## DXGErrors
ErrorCode | Identifier | Priority | IdOfTextToDisplay | TextToDisplay | ConfirmRequired | Description | ExtendedDescription | ServID | ServCode | ServiceMessage
1024 | DTC_PERFTX_RDSYNC | Error |  |  |  | Error while waiting for sync frame | Timeout or error occured while initiating the transmission of identification data from the PMI Board to the TagReader component at the Cassette Handling Board. Either the CAN bus is defective or at least the TagReader component at the Cassette Handling Board (or the complete Cassette Handling Board) does not respond anymore.
1025 | DTC_PERFTX_BKLWRITE | Error |  |  |  | Error during block transfer of ipdata | Either the CAN bus is defective or at least the TagReader component at the Cassette Handling Board (or the complete Cassette Handling Board) does not respond anymore.
1026 | DTC_INITX_FCHOPEN | Error |  |  |  | Cannot open CAN channel for data transmission | Unable to open a channel to transmit the identification data from the PMI Board to the TagReader component at the Cassettte Handling Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1027 | DTC_PERFRX_FCHOPEN | Error |  |  |  | Cannot open CAN channel for data retrieval | Unable to open a channel to transmit the identification data from the Cassette Handling Board to the PMI Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1028 | DTC_PERFRX_INITBLKRX | Error |  |  |  | Cannot initialize CAN channel for block retrieval | Timeout or error occured while initializing a block transfer over the CAN bus from the Cassette Handling Board to the PMI Board. This error is related to the PMI Board itself or to the software running on this PCB.
1029 | DTC_PERFRX_WRSYNC | Error |  |  |  | Error while writing sync frame | An error occured when trying to start the transmission of identification data from the Cassette Handling Board to the PMI Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1030 | DTC_PERFRX_BLKRX | Error |  |  |  | Error while performing the block read | An error occured when reading the identifiaction data that are sent from the Cassette Handling Board to the PMI Board. This error is related to one or more of the following reasons: - The CAN bus is defective. - The Cassette Handling Board aborted the transmission (e.g. software crash, hardware fault, ...) - The PMI Board itself reported an error condition in the CAN logic.
1031 | VHLDLL_INITUART | Error |  |  |  | DLL: Initialization of the UART failed | Initializing the serial line that connects the Cassette Handling Board with the Chip Reader failed. This is an internal error of the Cassette Handling Board and a string indication that this PCB is defective and should be replaced.
1032 | VHLDLL_RCF_RCT1 | Error |  |  |  | DLL: Timeout while reading a character from the TagReader component | The TagReader component at the Cassette Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the Cassette Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The Cassette Handling Board is defective.
1033 | VHLDLL_RCF_RCT2 | Error |  |  |  | DLL: Timeout while reading a character from the TagReader component | The TagReader component at the Cassette Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the Cassette Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The Cassette Handling Board is defective.
1034 | VHLDLL_SCF_RCT | Error |  |  |  | DLL: Timeout while waiting for Sync ACK from the TagReader component | The TagReader component at the Cassette Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the Cassette Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The Cassette Handling Board is defective.
1035 | VHLDLL_RA_RCT | Error |  |  |  | DLL: Timeout while waiting for Sync ACK from the TagReader component | The TagReader component at the Cassette Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the Cassette Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The Cassette Handling Board is defective.
1036 | VHLDLL_RX_UNKNOWN | Error |  |  |  | DLL: Unknown event in method receive() | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the Cassette Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The Cassette Handling Board is defective.
1037 | VHLDLL_TX_UNKNOWN | Error |  |  |  | DLL: Unknown event in method send() | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the Cassette Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The Cassette Handling Board is defective.
1038 | VHLDLL_STT_RXINITIAL_UNKNOWN | Error |  |  |  | DLL: Unknown event in initial transition in state OK::receiving | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the Cassette Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The Cassette Handling Board is defective.
1039 | VHLAL_POWERON_TX | Error |  |  |  | AL: Error sending command powerON() | Sending the command 'antenna on' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1040 | VHLAL_POWERON_RX | Error |  |  |  | AL: Error receiving answer for powerON() | Receiving the answer to the command 'antenna on' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1041 | VHLAL_POWERON_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to powerON() | The answer to the command 'antenna on' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1042 | VHLAL_POWERON_SEQID | Error |  |  |  | AL: Sequence error in answer for powerON() | The answer to the command 'antenna on' from the Chip Reader carries a wrong sequence ID.
1043 | VHLAL_SELECT_TX | Error |  |  |  | AL: Error sending command SELECT() | Sending the command 'select tag' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1044 | VHLAL_SELECT_RX | Error |  |  |  | AL: Error receiving answer for SELECT() | Receiving the answer to the command 'select tag' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1045 | VHLAL_SELECT_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to SELECT() | The answer to the command 'select tag' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1046 | VHLAL_SELECT_SEQID | Error |  |  |  | AL: Sequence error in answer for SELECT() | The answer to the command 'select tag' from the Chip Reader carries a wrong sequence ID.
1047 | VHLAL_GETINFO_TX | Error |  |  |  | AL: Error sending command GETINFO() | Sending the command 'get info' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1048 | VHLAL_GETINFO_RX | Error |  |  |  | AL: Error receiving answer for GETINFO() | Receiving the answer to the command 'get info' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1049 | VHLAL_GETINFO_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to GETINFO() | The answer to the command 'get info' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1050 | VHLAL_GETINFO_SEQID | Error |  |  |  | AL: Sequence error in answer for GETINFO() | The answer to the command 'get info' from the Chip Reader carries a wrong sequence ID.
1051 | VHLAL_READ_TX | Error |  |  |  | AL: Error sending command READ() | Sending the command 'read' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1052 | VHLAL_READ_RX | Error |  |  |  | AL: Error receiving answer for READ() | Receiving the answer to the command 'read' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1053 | VHLAL_READ_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to READ() | The answer to the command 'read' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1054 | VHLAL_READ_SEQID | Error |  |  |  | AL: Sequence error in answer for READ() | The answer to the command 'read' from the Chip Reader carries a wrong sequence ID.
1055 | VHLAL_WRITE_TX | Error |  |  |  | AL: Error sending command WRITE() | Sending the command 'write' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1056 | VHLAL_WRITE_RX | Error |  |  |  | AL: Error receiving answer for WRITE() | Receiving the answer to the command 'write' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1057 | VHLAL_WRITE_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to WRITE() | The answer to the command 'write' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1058 | VHLAL_WRITE_SEQID | Error |  |  |  | AL: Sequence error in answer for WRITE() | The answer to the command 'write' from the Chip Reader carries a wrong sequence ID.
1059 | VHLAL_POWEROFF_TX | Error |  |  |  | AL: Error sending command powerOFF() | Sending the command 'antenna off' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1060 | VHLAL_POWEROFF_RX | Error |  |  |  | AL: Error receiving answer for powerOFF() | Receiving the answer to the command 'antenna off' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1061 | VHLAL_POWEROFF_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to powerOFF() | The answer to the command 'antenna off' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1062 | VHLAL_POWEROFF_SEQID | Error |  |  |  | AL: Sequence error in answer for powerOFF() | The answer to the command 'antenna off' from the Chip Reader carries a wrong sequence ID.
1063 | VHLAL_SETMIF_TX | Error |  |  |  | AL: Error sending command setupMIFARE() | Sending the command 'select Mifare access mode' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1064 | VHLAL_SETMIF_RX | Error |  |  |  | AL: Error receiving answer for setupMIFARE() | Receiving the answer to the command 'select Mifare access mode' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1065 | VHLAL_SETMIF_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to setupMIFARE() | The answer to the command 'select Mifare access mode' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1066 | VHLAL_SETMIF_SEQID | Error |  |  |  | AL: Sequence error in answer for setupMIFARE() | The answer to the command 'select Mifare access mode' from the Chip Reader carries a wrong sequence ID.
1067 | VHLAL_GETSNR_TX | Error |  |  |  | AL: Error sending command getSNR() | Sending the command 'getSerialNumberOfTag' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the Cassette Handling Board and the Chip Reader is broken or not connected properly.
1068 | VHLAL_GETSNR_RX | Error |  |  |  | AL: Error receiving answer for getSNR() | Receiving the answer to the command 'getSerialNumberOfTag' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1069 | VHLAL_GETSNR_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to getSNR() | The answer to the command 'getSerialNumberOfTag' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the Cassette Handling Board.
1070 | VHLAL_GETSNR_SEQID | Error |  |  |  | AL: Sequence error in answer for getSNR() | The answer to the command 'getSerialNumberOfTag' from the Chip Reader carries a wrong sequence ID.
1071 | APP_TEST_ALINIT | Error |  |  |  | APP: Initializing the application layer failed | Cannot initialize the communication layer between the Cassette Handling Board and the Chip Reader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1072 | APP_TEST_GETINFO | Error |  |  |  | APP: Cannot retrieve firmware information | Cannot execute command 'getFirmwareVersion' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1073 | APP_GETINFO | Error |  |  |  | APP: Cannot retrieve firmware information | Cannot execute command 'getFirmwareVersion' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1074 | APP_TESTTAG_PWRON | Error |  |  |  | APP: Cannot switch on RF-power in testTAG() | Cannot execute command 'testForTag/AntennaOn' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1075 | APP_TESTTAG_PWROFF | Error |  |  |  | APP: Cannot switch off RF-power in testTAG() | Cannot execute command 'testForTag/AntennaOff' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1076 | APP_ENABLE_SETMIF | Error |  |  |  | APP: Cannot setupMifare access mode in method ENABLE | Cannot execute command 'Enable/setMifareAccMode' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1077 | APP_ENABLE_GETSNR | Error |  |  |  | APP: Cannot retrieve serial number of the tag in method ENABLE | Cannot execute command 'Enable/GetSerialNumberOfTag' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1078 | APP_ENABLE_SELECT | Error |  |  |  | APP: Cannot select tag in method ENABLE | Cannot execute command 'Enable/SelectTag' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1079 | APP_ENABLE_PWRON | Error | 117 | Error during writing on cassette chip after successful scan. Erase the image plate again. |  | APP: Cannot switch RF-power on in method ENABLE | Cannot execute command 'Enable/AntennaOn' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1080 | APP_DISABLE_PWROFF | Error |  |  |  | APP: Cannot switch RF-power off in method DISABLE | Cannot execute command 'Enable/AntennaOff' on the TagReader. Possible reasons are: - The cable between the Cassette Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The Cassette Handling Board might be defective.
1081 | APP_RD_GETPAR | Error |  |  |  | APP: Cannot retrieve signal parameters in method READ | Cannot execute command 'Read' on the TagReader. Unable to retrieve the parameters from the message. Internal software(?) error at the Cassette Handling Board.
1082 | APP_RD_SYNCFRAME | Error |  |  |  | APP: Cannot receive SYNC frame in method READ | Cannot execute command 'Read' on the TagReader. Unable to initiate data transfer to the PMI Board from the Cassette Handling Board. Error on CAN bus or PMI Board or Cassette Handling Board.
1083 | APP_RD_BLKWRITE | Error |  |  |  | APP: Cannot perform block write in method READ | Cannot execute command 'Read' on the TagReader. Unable to perform the data transfer from the Cassette Handling Board to the PMI Board. Error on CAN bus or PMI Board or Cassette Handling Board.
1084 | APP_RD_FCHOPEN | Error |  |  |  | APP: Cannot open CAN channel for data transfer in method READ | Cannot execute command 'Read' on the TagReader. Unable to open channel on CAN bus. Internal error of the Cassette Handling Board.
1085 | APP_WR_FCHOPEN | Error |  |  |  | APP: Cannot open CAN channel for data transfer in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to open channel on CAN bus. Internal error of the Cassette Handling Board.
1086 | APP_WR_GETPAR | Error |  |  |  | APP: Cannot retrieve signal parameters in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to retrieve the parameters from the message. Internal software(?) error at the Cassette Handling Board.
1087 | APP_WR_BLKRD | Error |  |  |  | APP: Cannot perform block read in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to perform the data transfer from the PMI Board. Error on CAN bus or PMI Board or Cassette Handling Board.
1088 | APP_WR_SYNCFRAME | Error |  |  |  | APP: Cannot send SYNC frame in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to initiate the data transfer from the PMI Board. Error on CAN bus or PMI Board or Cassette Handling Board.
1089 | APP_WR_BRXINIT | Error |  |  |  | APP: Cannot initialize block transfer in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to initialize the data transmission from the PMI Board. Error on Cassette Handling Board.
1090 | RD_FAILED | Error |  |  |  | command read failed. | Cure: Check cabling of the TagReader (broken?)
1536 | INFO_VERS_TERM | Error |  |  |  | IPInfo::Version: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::Version attribute read. | COM_SWCorruptIDData
1537 | INFO_VERS_CORRUPTED | Error |  |  |  | IPInfo::Version: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Version. | COM_SWCorruptIDData
1538 | INFO_IDNUM_TERM | Error |  |  |  | IPInfo::IDNumber: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IDNumber attribute read. | COM_SWCorruptIDData
1539 | INFO_IDNUM_CORRUPTED | Error |  |  |  | IPInfo::IDNumber: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IDNumber. | COM_SWCorruptIDData
1540 | INFO_INITDATE_TERM | Error |  |  |  | IPInfo::InitDate: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::InitDate attribute read. | COM_SWCorruptIDData
1541 | INFO_INITDATE_CORRUPTED | Error |  |  |  | IPInfo::InitDate: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::InitDate. | COM_SWCorruptIDData
1542 | INFO_IPSIZE_TERM | Error |  |  |  | IPInfo::IPSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSize attribute read. | COM_SWCorruptIDData
1543 | INFO_IPSIZE_CORRUPTED | Error |  |  |  | IPInfo::IPSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSize. | COM_SWCorruptIDData
1544 | INFO_SCANSIZE_TERM | Error |  |  |  | IPInfo::ScanSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::ScanSize attribute read. | COM_SWCorruptIDData
1545 | INFO_SCANSIZE_CORRUPTED | Error |  |  |  | IPInfo::ScanSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::ScanSize. | COM_SWCorruptIDData
1546 | INFO_IPTYPE_TERM | Error |  |  |  | IPInfo::IPType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPType attribute read. | COM_SWCorruptIDData
1547 | INFO_IPTYPE_CORRUPTED | Error |  |  |  | IPInfo::IPType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPType. | COM_SWCorruptIDData
1548 | INFO_IPSENS_TERM | Error |  |  |  | IPInfo::IPSensitivity: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSensitivity attribute read. | COM_SWCorruptIDData
1549 | INFO_IPSENS_CORRUPTED | Error |  |  |  | IPInfo::IPSensitivity: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSensitivity. | COM_SWCorruptIDData
1550 | INFO_IPERA_TERM | Error |  |  |  | IPInfo::IPErasibility: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPErasibility attribute read. | COM_SWCorruptIDData
1551 | INFO_IPERA_CORRUPTED | Error |  |  |  | IPInfo::IPErasibility: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPErasibility. | COM_SWCorruptIDData
1552 | INFO_IPTHICK_TERM | Error |  |  |  | IPInfo::IPThickness: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPThickness attribute read. | COM_SWCorruptIDData
1553 | INFO_IPTHICK_CORRUPTED | Error |  |  |  | IPInfo::IPThickness: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPThickness. | COM_SWCorruptIDData
1554 | INFO_CHECKSUM_TERM | Error |  |  |  | IPInfo::Checksum: unexpected attribute terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: No section terminator after IPInfo::Checksum attribute read. | COM_SWCorruptIDData
1555 | INFO_CHECKSUM_CORRUPTED | Error |  |  |  | IPInfo::Checksum: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Checksum. | COM_SWCorruptIDData
1556 | INFO_CHECKSUM_WRONG | Error |  |  |  | IPInfo::Checksum: checksum error | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Checksum error for section IPInfo. | COM_SWCorruptIDData
1557 | STATUS_VERS_TERM | Error |  |  |  | IPStatus::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Version attribute read. | COM_SWCorruptIDData
1558 | STATUS_VERS_NOTSUPPORTED | Error |  |  |  | IPStatus::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPStatus::Version not supported. | COM_SWCorruptIDData
1559 | STATUS_VERS_CORRUPTED | Error |  |  |  | IPStatus::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Version. | COM_SWCorruptIDData
1560 | STATUS_STATUS_TERM | Error |  |  |  | IPStatus::Status: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Status attribute read. | COM_SWCorruptIDData
1561 | STATUS_STATUS_CORRUPTED | Error |  |  |  | IPStatus::Status: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Status. | COM_SWCorruptIDData
1562 | STATUS_CYCLES_TERM | Error |  |  |  | IPStatus::Cycles: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Cycles attribute read. | COM_SWCorruptIDData
1563 | STATUS_CYCLES_CORRUPTED | Error |  |  |  | IPStatus::Cycles: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Cycles. | COM_SWCorruptIDData
1564 | STATUS_CHECKSUM_TERM | Error |  |  |  | IPStatus::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPStatus::Checksum attribute read. | COM_SWCorruptIDData
1565 | STATUS_CHECKSUM_CORRUPTED | Error |  |  |  | IPStatus::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Checksum. | COM_SWCorruptIDData
1566 | STATUS_CHECKSUM_WRONG | Error |  |  |  | IPStatus::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPStatus. | COM_SWCorruptIDData
1567 | CTRL_VERS_TERM | Error |  |  |  | IPCtrl::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Version attribute read. | COM_SWCorruptIDData
1568 | CTRL_VERS_NOTSUPPORTED | Error |  |  |  | IPCtrl::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPCtrl::Version not supported. | COM_SWCorruptIDData
1569 | CTRL_VERS_CORRUPTED | Error |  |  |  | IPCtrl::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Version. | COM_SWCorruptIDData
1570 | CTRL_SENSITIVITY_TERM | Error |  |  |  | IPCtrl::Sensitivity: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Sensitivity attribute read. | COM_SWCorruptIDData
1571 | CTRL_SENSITIVITY_CORRUPTED | Error |  |  |  | IPCtrl::Sensitivity: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Sensitivity. | COM_SWCorruptIDData
1572 | CTRL_EXPDOSE_TERM | Error |  |  |  | IPCtrl::ExposureDose: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::ExposureDose attribute read. | COM_SWCorruptIDData
1573 | CTRL_EXPDOSE_CORRUPTED | Error |  |  |  | IPCtrl::ExposureDose: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::ExposureDose. | COM_SWCorruptIDData
1574 | CTRL_OPCODE_TERM | Error |  |  |  | IPCtrl::OperationCode: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::OpCode attribute read. | COM_SWCorruptIDData
1575 | CTRL_OPCODE_CORRUPTED | Error |  |  |  | IPCtrl::CoperationCode: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::OpCode. | COM_SWCorruptIDData
1576 | CTRL_CHECKSUM_TERM | Error |  |  |  | IPCtrl::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPCtrl::Checksum attribute read. | COM_SWCorruptIDData
1577 | CTRL_CHECKSUM_CORRUPTED | Error |  |  |  | IPCtrl::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Checksum. | COM_SWCorruptIDData
1578 | CTRL_CHECKSUM_WRONG | Error |  |  |  | IPCtrl::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPCtrl. | COM_SWCorruptIDData
1579 | IPPSDATA_VERS_TERM | Error |  |  |  | IPPSData::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::Version attribute read. | COM_SWCorruptIDData
1580 | IPPSDATA_VERS_NOTSUPPORTED | Error |  |  |  | IPPSData::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPPSData::Version not supported. | COM_SWCorruptIDData
1581 | IPPSDATA_VERS_CORRUPTED | Error |  |  |  | IPPSData::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::Version. | COM_SWCorruptIDData
1582 | IPPSDATA_MPPSINSTUID_TERM | Error |  |  |  | IPPSData::MPPSInstanceUID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::MPPSInstanceUID attribute read. | COM_SWCorruptIDData
1583 | IPPSDATA_MPPSINSTUID_CORRUPTED | Error |  |  |  | IPPSData::MPPSInstanceUID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::MPPSInstanceUID. | COM_SWCorruptIDData
1584 | IPPSDATA_SOPINSTUID_TERM | Error |  |  |  | IPPSData::SOPInstanceUID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::SOPInstanceUID attribute read. | COM_SWCorruptIDData
1585 | IPPSDATA_SOPINSTUID_CORRUPTED | Error |  |  |  | IPPSData::SOPInstanceUID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::SOPInstanceUID. | COM_SWCorruptIDData
1586 | IPPSDATA_PROCSTAT_TERM | Error |  |  |  | IPPSData::ProcessingStation: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ProcessingStation attribute read. | COM_SWCorruptIDData
1587 | IPPSDATA_PROCSTAT_CORRUPTED | Error |  |  |  | IPPSData::ProcessingStation: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ProcessingStation. | COM_SWCorruptIDData
1588 | IPPSDATA_STIMIDENTID_TERM | Error |  |  |  | IPPSData::StimIdentID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::StimIdentID attribute read. | COM_SWCorruptIDData
1589 | IPPSDATA_STIMIDENTID_CORRUPTED | Error |  |  |  | IPPSData::StimIdentID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::StimIdentID. | COM_SWCorruptIDData
1590 | IPPSDATA_EXPROOMID_TERM | Error |  |  |  | IPPSData::ExposureRoomID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ExposureRoom attribute read. | COM_SWCorruptIDData
1591 | IPPSDATA_EXPROOMID_CORRUPTED | Error |  |  |  | IPPSData::ExposureRoomID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ExposureRoom. | COM_SWCorruptIDData
1592 | IPPSDATA_CHECKSUM_TERM | Error |  |  |  | IPPSData::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPPSData::Checksum attribute read. | COM_SWCorruptIDData
1593 | IPPSDATA_CHECKSUM_CORRUPTED | Error |  |  |  | IPPSData::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::Checksum. | COM_SWCorruptIDData
1594 | IPPSDATA_CHECKSUM_WRONG | Error |  |  |  | IPPSData::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPPSData. | COM_SWCorruptIDData
1595 | INFO_VERSION_NOTSUPPORTED | Error |  |  |  | IPInfo::Version: this version is not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPInfo::Version not supported. | COM_SWCorruptIDData
1596 | INFO_APPTYPE_TERM | Error |  |  |  | IPInfo::ApplicationType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::AppType attribute read. | COM_SWCorruptIDData
1597 | INFO_APPTYPE_CORRUPTED | Error |  |  |  | IPInfo::ApplicationType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::AppType. | COM_SWCorruptIDData
1598 | IPCTRL_PATNAME_TERM | Error |  |  |  | IPCtrl::PatientName: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::PatientName attribute read. | COM_SWCorruptIDData
1599 | IPCTRL_PATNAME_CORRUPTED | Error |  |  |  | IPCtrl::PatientName: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::PatientName. | COM_SWCorruptIDData
1600 | INFO_OPTION_TERM | Error |  |  |  | IPInfo::Option: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::Option attribute read. | COM_SWCorruptIDData
1601 | INFO_OPTION_CORRUPTED | Error |  |  |  | IPInfo::Option: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Option. | COM_SWCorruptIDData
7168 | CAS_LATENCY_NOT_SUPPORTED | Error |  |  |  | FPGA initialization failed because CAS_LATENCY is not supported | Meaning: FpgaAccess: FPGA initialization failed because CAS_LATENCY is not supported. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7169 | BURST_LENGTH_NOT_SUPPORTED | Error |  |  |  | FPGA initialization failed because BURST_LENGTH is not supported | Meaning: FpgaAccess: FPGA initialization failed because BURST_LENGTH is not supported. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7170 | SDRAM_CONTROLLER_START_FAILED | Error |  |  |  | FPGA initialization failed because SDRAM controller do not start | Meaning: FpgaAccess: FPGA initialization failed because SDRAM controller do not start. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7171 | TEST_FPGA_NOT_LOADED | Error |  |  |  | FPGA test failed FPGA is not loaded yet | Meaning: FpgaAccess: FPGA test failed FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_NotNeeded_ForServiceId
7172 | PATTGEN_FPGA_NOT_LOADED | Error |  |  |  | FPGA pattern generation failed because FPGA is not loaded yet | Meaning: FpgaAccess: FPGA pattern generation failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_NotNeeded_ForServiceId
7173 | PATTGEN_FLUSH_FIFO_UNDERRUN | Error |  |  |  | FPGA pattern generation failed because flush FIFO underrun occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO underrun occured. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7174 | PATTGEN_FLUSH_FIFO_OVERFLOW | Error |  |  |  | FPGA pattern generation failed because flush FIFO overflow occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO overflow occured. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7175 | PATTGEN_FLUSH_FIFO_DRAIN_NOK | Error |  |  |  | FPGA pattern generation failed because flush FIFO cannot be drained | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7176 | FLUSH_FIFO_DRAIN_FAILED | Error |  |  |  | flush FIFO cannot be drained; HW problem, fatal error | Meaning: FpgaAccess: flush FIFO cannot be drained; HW problem, fatal error. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7177 | FPGA_TEST_FAILED | Error |  |  |  | FPGA test failed (one of the RAMs was not correctly accessible), fatal error | Meaning: FpgaAccess: FPGA test failed (one of the RAMs was not correctly accessible). Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7178 | TIMEOUT_FILLING_SCAN_FIFO | Error |  |  |  | scan fifo write pump did not fill scan fifo with SDRAM data, fatal error | Meaning: FpgaAccess: scan fifo write pump did not fill scan fifo with SDRAM data. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7179 | UNKNOWN_FLUSH_FIFO_INPUT | Error |  |  |  | unknown input for flush fifo selected, fatal software error | Meaning: FpgaAccess: unknown input for flush fifo selected. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7180 | FILLSDRAM_FPGA_NOT_LOADED | Error |  |  |  | filling the SDRAM failed because FPGA is not loaded yet | Meaning: FpgaAccess: filling the SDRAM failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_NotNeeded_ForServiceId
7181 | FILLSDRAM_IMAGE_TOO_BIG | Error |  |  |  | filling the SDRAM failed because desired image is too big | Meaning: FpgaAccess: filling the SDRAM failed because desired image is too big. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7182 | PATTGEN_SDRAM_SIZE_UNKNOWN | Error |  |  |  | FPGA pattern generation failed because size of SDRAM is not determined yet (call initFpga first!) | Meaning: FpgaAccess: FPGA pattern generation failed because size of SDRAM is not determined yet. Reason: SW error (call initFpga first!). Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7183 | PATTGEN_EXCEEDS_SDRAM | Error |  |  |  | FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM | Meaning: FpgaAccess: FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7184 | LOAD_FPGA_CONFIGURATION | Error |  |  |  | Can't load FPGA configuration, fatal error | Meaning: FpgaAccess: Can't load FPGA configuration. Reason: Hardware defect or software mismatch. Cure: Upload PMI software and FPGA configuration; if still NOK, replace PMI board | COM_NotNeeded_ForServiceId
7185 | LOAD_FPGA_HARDWARE | Error |  |  |  | Load FPGA: can't set FPGA into configuration mode, fatal error | Meaning: FpgaAccess: Load FPGA can't set FPGA into configuration mode. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7186 | STORED_FPGA_CONFIGURATION | Error |  |  |  | No valid FPGA version found on PMI | Meaning: FPGA version doesn't fit to hard- or software version. Reason: SW or HW error. Cure: Update software including FPGA configuration; if still NOK, replace PMI board | COM_NotNeeded_ForServiceId
7187 | READ_PMI_CPU | Error |  |  |  | Can't read from FPGA CPU | Meaning: Can't read from FPGA CPU. Reason: Hardware error on PMI. Cure: Retry switch digitizer off/on; if still NOK, replace PMI board. | COM_NotNeeded_ForServiceId
7188 | CAL_ERROR_FLAG | Error |  |  |  | PMI CPU Error flag is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Polygon or Laser or PMI board defective. Cure: Retry switch digitizer off/on, replace lens assembly CR30X, replace PMI board. | COM_NotNeeded_ForServiceId
7189 | CAL_EOL | Error |  |  |  | EOL error counter is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Optical path is partly interrupted, calibration board is defective. Cure: Clean scan line with antistatic brush, check mounting of light guide, switch digitizer off/on, replace calibration plate | COM_GeoCalibrationErr
7190 | CAL_CAL_SENSORS | Error |  |  |  | Calibration sensor error counter is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Optical path is partly interrupted, calibration board is defective. Cure: Clean scan line with antistatic brush, switch digitizer off/on, replace calibration plate | COM_GeoCalibrationErr
7680 | IMAGE_SERVER_CREATION_FAILED | Error |  |  |  | ImageTransferControl cannot create image data server; fatal! | Meaning: ImageTransferControl cannot create image data server. Reason: SW error. Cure: Install new software. | COM_ImageTransferControlErr
7681 | INVALID_MAX_PACKET_SIZE | Error |  |  |  | ImageTransferControl received wrong FW packet size from PC; fatal! | Meaning: ImageTransferControl received wrong FW packet size from PC. Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7682 | INVALID_FW_SPEED_CODE | Error |  |  |  | ImageTransferControl determined invalid FW speed; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&D team in any case. | COM_ImageTransferControlErr
7683 | BLOCK_TRANSFER_TIMEOUT_EXCEEDED | Error |  |  |  | ImageTransferControl: block transfer takes too long; either FW cable was disconnected during the block transfer or a fatal HW error occured! | Meaning: ImageTransferControl: block transfer takes too long. Reason: The transmission of an image data block takes too long. Cure: Check if FW cable is connected correctly; if still NOK, replace FW cable; if still NOK, replace PMI board; if still NOK, replace FW-board on PC. | COM_ImageTransferControlErr
7684 | INVALID_PC_NODE_ID | Error |  |  |  | ImageTransferControl: block transfer failed because PC's node ID is invalid; fatal error! | Meaning: ImageTransferControl: block transfer failed because PC's node ID is invalid. Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7685 | SCAN_FIFO_OVERFLOW | Error |  |  |  | ImageTransferControl: scan fifo overflow during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo overflow during block transfer. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7686 | SCAN_FIFO_UNDERRUN | Error |  |  |  | ImageTransferControl: scan fifo underrun during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo underrun during block transfer. Reason: HW error. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI board and PMT/lens assembly; if OK, replace PMI board; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly | COM_ImageTransferControlErr
7687 | FPGA_BUS_ERROR | Error |  |  |  | ImageTransferControl: FPGA bus error occured; fatal! | Meaning: ImageTransferControl: FPGA bus error occured. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7688 | INVALID_NR_OF_REQUESTED_BYTES | Error |  |  |  | ImageTransferControl: invalid number of requested bytes ( <=0); fatal SW bug! | Meaning: ImageTransferControl: invalid number of requested bytes ( <=0). Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7689 | TIMEOUT_GET_PACKET_SIZE | Error |  |  |  | ImageTransferControl: timeout: cannot determine used FW packet size; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&D team in any case. | COM_ImageTransferControlErr
7690 | TIMEOUT_DATA_NOT_AVAILABLE | Error |  |  |  | ImageTransferControl: timeout: scanner didn't store enough data in SDRAM yet to start block transfer; wait, then try again! | Meaning: ImageTransferControl: timeout. Reason: Scanner didn't store enough data in SDRAM yet to start block transfer. Cure: No problem, PC software has to wait and try again. | COM_ImageTransferControlErr
7691 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump. Reason: HW error (DM (DataMover) pump or LLC (LinkLayerController) failed). Cure: Replace PMI board. | COM_ImageTransferControlErr
7692 | RECOVER_PACKET_FAILED | Error |  |  |  | ImageTransferControl: recovering the FW packet failed; fatal error! | Meaning: ImageTransferControl: recovering the FW packet failed. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7693 | TOO_MANY_PACKETS_PER_BLOCK | Error |  |  |  | ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block); request smaller block! | Meaning: ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block). Reason: SW error (PC must request smaller block!). Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7694 | FLUSH_FIFO_DRAIN_FAILED | Error |  |  |  | flush FIFO cannot be drained; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7695 | PIXELVIEWER_FLUSH_FIFO_OVERFLOW | Error |  |  |  | flush FIFO overflow occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO overflow occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7696 | PIXELVIEWER_FLUSH_FIFO_UNDERRUN | Error |  |  |  | flush FIFO underrun occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO underrun occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7697 | PIXELVIEWER_CRC1_AND_2_NOT_EQUAL | Error |  |  |  | after scanning in PixelViewer mode, the CRC1 and 2 are not equal; HW problem, fatal error! | Meaning: ImageTransferControl: after scanning in PixelViewer mode, the CRC1 and 2 are not equal. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7698 | PIXELVIEWER_START_NEXT_SCAN_FAILED | Error |  |  |  | starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running; HW failure (e.g. BOL missing) or SW synchronization problem, fatal error! | Meaning: ImageTransferControl: starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running. Reason: HW failure (e.g. BOL missing) or SW synchronization problem. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_ImageTransferControlErr
7699 | DM_ERROR_HANDLER_SETUP_FAILED | Error |  |  |  | registering the callback function for DM errors failed | Meaning: ImageTransferControl: registering the callback function for DM errors failed. Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7700 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_START | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7701 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_OVFL | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7702 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_URUN | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7703 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_RETRANS | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7704 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER | Error |  |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost. | COM_ImageTransferControlErr
7705 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_START | Error |  |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. Actual transfer was not started yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost. | COM_ImageTransferControlErr
7706 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_PUMP | Error |  |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. DM pump was not enabled yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image transmission automatically again, image is not lost. | COM_ImageTransferControlErr
7707 | IMAGE_BUFFER_CREATION_FAILED | Error |  |  |  | ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. | Meaning: ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. Reason: not enough memory resources (probably software bug) Cure: Install new software. | COM_ImageTransferControlErr
7708 | MORE_DATA_REQUESTED_THAN_SCANNED | Error |  |  |  | ImageTransferUDPControl: more data requested than scanned | Meaning: ImageTransferUDPControl: more data requested than scanned. Reason: Index out of range. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7709 | SERVER_DISCONNECTED_DURING_BLOCK_TRANSFER | Error |  |  |  | ImageTransferUDPControl: block transfer stopped because server was disconnected! | Meaning: ImageTransferUDPControl: block transfer stopped because server was disconnected! Reason: Cable was disconnected during transmission of image data to PC. Cure: Connect cable again; PC will request the image trasnmission automatically again, image is not lost. | COM_ImageTransferControlErr
7710 | NO_DATA_IN_DRAM_TIMEOUT | Error |  |  |  | ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. | Meaning: ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. Reason: Copying data into DRAM takes too long, or, no data anymore in FPGA. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7711 | DMA_NOT_PREPARED_TIMEOUT | Error |  |  |  | ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. | Meaning: ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. Reason: Timing problem. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7712 | PATGEN_ENABLE_PROBLEM | Error |  |  |  | ImageTransferUDPControl: The Pattern Generator cannot be enabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be enabled. Reason: Could be software or hardware problem. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7713 | PATGEN_DISABLE_PROBLEM | Error |  |  |  | ImageTransferUDPControl: The Pattern Generator cannot be disabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be disabled. Reason: Could be software or hardware problem. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7714 | IMAGE_TRANSFER_CANCELLED | Error |  |  |  | ImageTransferUDPControl: the image transfer was cancelled | Meaning: ImageTransferUDPControl: the image transfer was cancelled. | COM_ImageTransferControlErr
7715 | DM_ERROR | Error |  |  |  | ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error | Meaning: ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error. Reason: A DM error occurred during the transmission of an image block. Cure: The PC will automatically request the image block again (up to 10 times); if all repetitions fail an image transer error is displayed; in the latter case reboot the PC, the PC will start the image transmission automatically again, the image is not lost as long as the digitizer is not switched off. | COM_ImageTransferControlErr
8192 | LOST_MESSAGES | Error |  |  |  | Lost logmessages
8193 | SAVE_DIGITIZERSETTINGS | Error |  |  |  | Could not save digitizer settings
8704 | TEST_MODULE_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module tests | Meaning: CassetteHandlingControl:: Error occurred during module tests Reason: Cure:
8705 | TEST_CASS_EDGE_SWITCH_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest | Meaning: CassetteHandlingControl:: Error occurred during module tests Reason: Cassette Edge Sensor disconnected or defective Cure: 1) Check cables and connections, 2) check sensor. 3) Replace Cassette Edge Sensor, if error persists.
8706 | TEST_CASS_TRANSP_REF_LB_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest at CassTransportRefLB | Meaning: CassetteHandlingControl:: Error occurred during module tests Reason: Cassette Transport Reference Lightbarrier (LB) disconnected or defective. Cure: 1) Check cables and connections. 2) check LB. 3) Check the cassette transport belt and the driving cams on the belt for damage. In this case repair the belt and the lightbarrier with its connection cable.
8707 | TEST_LOCKER_CONTACT_LB_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest at Locker Contact LB | Meaning: CassetteHandlingControl: Error occurred during module tests Reason: LOCKING CASSETTE-CONTACT SENSOR disconnected or defect Cure: 1) Check cables and connections. 2) Check LOCKING CASSETTE-CONTACT SENSOR. 3) Replace LOCKING CASSETTE-CONTACT SENSOR, if error persists.
8708 | TEST_OPENER_CLOSE_FLAP_LB_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest at Opener Close Flap LB | Meaning: CassetteHandlingControl: Error occurred during module tests Reason: OPENER HOME-POSITION SENSOR disconnected or defect Cure: 1) Check cables and connections. 2) Check OPENER HOME-POSITION SENSOR. 3) Replace OPENER HOME-POSITION SENSOR, if error persists.
8709 | TEST_OPENER_REF_LB_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest at Opener Reference LB | Meaning: CassetteHandlingControl: Error occurred during module tests Reason: OPENER CONTROL SENSOR disconnected or defect Cure: 1) Check cables and connections. 2) Check OPENER CONTROL SENSOR. 3) Replace OPENER CONTROL SENSOR, if error persists.
8710 | TEST_PUSHER_REF_LB_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest at Pusher Reference LB | Meaning: CassetteHandlingControl: Error occurred during module tests Reason: PUSHER HOME-POSITION SENSOR disconnected or defect Cure: 1) Check cables and connections. 2) Check PUSHER HOME-POSITION SENSOR . 3) Replace PUSHER HOME-POSITION SENSOR, if error persists.
8711 | TEST_CASS_SLOT_LB_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest at Cassette Slot LB | Meaning: CassetteHandlingControl:: Error occurred during module tests Reason: Locking cassette contact sensor disconnected or defective Cure: Check cables and connections, check sensor (lightbarrier).
8712 | TEST_BUFFER_FULL_1_LB_FAILED | Error |  |  |  | CassetteHandlingControl:: Error occurred during module selftest at Buffer Full 1 LB | Meaning: CassetteHandlingControl:: Error occurred during module tests Reason: Cassette output buffer sensor (lightbarrier) disconnected or defective Cure: Check cables and connections, check lightbarrier.
8713 | INIT_MOTOR_TIMEOUT | Error |  |  |  | CassetteHandlingControl: Timeout while initializing the motors. | Meaning: CassetteHandlingControl: Timeout in initialization of motors. Reason: A motor did not complete its movement in time. Cure: See logging (file: log.xml) for the precise error code and the motor that gave the problem, then check mechanics, belts, ... of the faulty drive.
8714 | SM_LOCKER_GENERAL_FAILURE | Error |  |  |  | CassetteHandlingControl: Stepper Motor Locker terminated an action with failure condition. | Meaning: SM LOCKER did not complete its movement as expected. Reason: Failure due to mechanical, sensoric or electronic influence. Cure: 1) Check LOCKER module mechanics, Cassette slider mechanics. 2) Check Locking Cassette Contact Sensor with cabling. 3) Check cassette handling board with cabling to SM LOCKER. 4) Replace LOCKER module. 5) Replace cassette handling board.
8715 | INIT_MOTOR_FIXER_TIMEOUT | Error |  |  |  | CassetteHandlingControl: Timeout while initializing the motor Fixer. | Meaning: CassetteHandlingControl: Timeout in initialization of SM CASSETTE FIXATION. Reason: The motor did not complete its movement in time. Cure: 1) Redo stall calibration 3 times. Check values for consistency (+/- 2 digits). 2) Check mechanics of SM CASSETTE FIXATION drive. 3) Check cabling to CASSETTE HANDLING BOARD. 4) Replace SM CASSETTE FIXATION if error persists. 5) Replace CASSETTE HANDLING BOARD if error persists.
8716 | INIT_MOTOR_LOCKER_TIMEOUT | Error |  |  |  | CassetteHandlingControl: Timeout while initializing the motor Locker. | Meaning: CassetteHandlingControl: Timeout in initialization of SM LOCKING. Reason: The motor did not complete its movement in time. Cure: 1) Redo stall calibration 3 times. Check values for consistency (+/- 2 digits). 2) Check mechanics of SM LOCKING drive. 3) Check cabling to CASSETTE HANDLING BOARD. 4) Replace SM SM LOCKING if error persists. 5) Replace CASSETTE HANDLING BOARD if error persists.
8717 | INIT_MOTOR_OPENER_TIMEOUT | Error |  |  |  | CassetteHandlingControl: Timeout while initializing the motor Opener. | Meaning: CassetteHandlingControl: Timeout in initialization of SM OPENER. Reason: The motor did not complete its movement in time. Cure: 1) Check mechanics of SM OPENER drive. 2) Check light barriers and motor cabling to CASSETTE HANDLING BOARD. 3) Check the light barriers OPENER HOME-POSITION SENSOR and OPENER CONTROL SENSOR 4) Replace OPENER HOME-POSITION SENSOR, OPENER CONTROL SENSOR if error persists. 5) Replace SM OPENER if error persists. 6) Replace CASSETTE HANDLING BOARD if error persists.
8718 | INIT_MOTOR_PUSHER_TIMEOUT | Error |  |  |  | CassetteHandlingControl: Timeout while initializing the motor Pusher. | Meaning: CassetteHandlingControl: Timeout in initialization of SM PUSHER. Reason: The motor did not complete its movement in time. Cure: 1) Redo stall calibration 3 times. Check values for consistency (+/- 2 digits). 2) Check mechanics of SM PUSHER drive. 3) Check cabling to CASSETTE HANDLING BOARD. 4) Replace SM SM PUSHER if error persists. 5) Replace CASSETTE HANDLING BOARD if error persists.
8719 | INIT_MOTOR_CASS_TRANSPORT_TIMEOUT | Error |  |  |  | CassetteHandlingControl: Timeout while initializing the motor Cassette Transport. | Meaning: CassetteHandlingControl: Timeout in initialization of SM CASSETTE TRANSPORT. Reason: The motor did not complete its movement in time. Cure: 1) Redo stall calibration 3 times. Check values for consistency (+/- 2 digits). 2) Check mechanics of SM CASSETTE TRANSPORT drive. 3) Check cabling to CASSETTE HANDLING BOARD. 4) Replace SM SM CASSETTE TRANSPORT if error persists. 5) Replace CASSETTE HANDLING BOARD if error persists.
8720 | CASS_TRANSPORT_RESPONSE_TIMEOUT | Error |  |  |  | CassetteHandlingControl: Response Timeout while moving the motor Cassette Transport. | Meaning: CassetteHandlingControl: Timeout at expected response of SM CASSETTE TRANSPORT. Reason: The motor did not complete its movement in time. Expected response message is missing. Cure: 1) Restart machine 2) Upgrade Software
8804 | WARN_CLAMP_CASS_INVALID_WIDTH | Warning |  |  |  | Warning: Clamp Cassette Movement detected invalid Cassette Width | Reason: The clamped Cassette has a invalid width. Meaning: Retry possible, function disturbed. Cure: Re-clamp Cassette. If error persists please calibrate mechanical Cassette Unit.
9216 | TEST_MODULE_FAILED | Error |  |  |  | BufferControl:: Error occurred during module tests | Meaning: BufferControl:: Error occurred during module tests Reason: Cure:
10240 | MOVEMENT_ENDED_WITH_TIMEOUT | Error |  |  |  | IPHandlingControl:: A mechanical movement ended with a timeout. | Meaning: A mechanical movement ended with a timeout. Reason: Cure: see 'cure' of corresponding movement
10241 | CABLE_ERASURE_UNIT_NOT_CONNECTED | Error |  |  |  | The data cable of the erasure unit is not connected or erasure unit power supply defective. | Meaning: The data cable of the erasure unit is not connected or the power supply for the erasure unit is defective. Reason: -- Cure: - Check the cable connection. Replace the IP handling erasure cable and/or the erasure unit if necessary.- Check the erasure unit power supply and replace if defective.
10242 | LED_ERASURE_UNIT_NOT_OPERATIONAL | Error |  |  |  | At least one of the LED's of the erasure unit is not operational. | Meaning: At least one of the LED's of the erasure unit is not operational. Reason: 1) Erasure unit defective. 2) Cabeling (power and/or data) to erasure unit defective. 3) Power supply of erasure unit defective. 4) IP-Handling board defective. Cure: 1) Replace erasure unit. 2) Replace appropiate cable. 3) Replace power supply of erasure unit. 4) Replace IP-Handling board.
10243 | ERASURE_UNIT_NOT_OPERATIONAL | Error |  |  |  | The erasure unit is not operational. | Meaning: The erasure unit is not operational. Reason: 1) Erasure unit defective. 2) Data cable to erasure unit defective. 3) IP-Handling board defective. 4) Power unit defective. Cure:1) Replace the erasure unit. 2) Replace the data cable to erasure unit. 3) Replace IP-Handling board. 4) Replace power unit
10244 | CLEAN_UP_PREPARE_INIT_TEST_FOR_IP_FAILED | Error |  |  |  | The clean-up procedure 'PrepareInitTestForIP' failed. | Meaning: The clean-up procedure failed. This error code is obsolete and has been replaced by specific error codes. Reason: -- Cure: Check Vacuum Valve drive and Swivel drive.
10245 | CLEAN_UP_INIT_TEST_FOR_Z_FREE_IP_FAILED | Error |  |  |  | The clean-up procedure 'InitTestForIPZFree' failed. | Meaning: The clean-up procedure failed. This error code is obsolete and has been replaced by specific error codes. Reason: -- Cure: Check the following drives: IP Carrier, Connector, IP Transport, Slowscan, Swivel, Vacuum Valve, Snaphooks.
10246 | CLEAN_UP_INIT_TEST_FOR_IP_Z_INTERRUPTED_FAILED | Error |  |  |  | The clean-up procedure 'InitTestForIPZInterrupted' failed. | Meaning: The clean-up procedure failed. This error code is obsolete. Reason: -- Cure:
10247 | CLEAN_UP_FINISH_INIT_TEST_FOR_IP_FAILED | Error |  |  |  | The clean-up procedure 'FinishInitTestForIP' failed. | Meaning: The clean-up procedure failed. This error code is obsolete and has been replaced by specific error codes. Reason: -- Cure: Check the following drives: IP Transport, Slowscan, Swivel, Vacuum Valve, Snaphooks.
10248 | CLEAN_UP_RETURN_IP_FAILED | Error |  |  |  | The clean-up procedure 'ReturnIP' failed. | Meaning: The clean-up procedure failed. Reason: There is a mechanical problem with the IP Transport drive or with the Z-sensor (not interrupted as expected). Cure:Remove cassette and IP manually and restart the digitizer.
10249 | CLEAN_UP_RELEASE_IP_FAILED | Error |  |  |  | The clean-up procedure 'ReleaseIP' failed. | Meaning: The clean-up procedure failed. Reason: There is a mechanical problem with the Snaphook drive. Cure: Remove cassette and IP manually and restart the digitizer.
10250 | MOVE_FEEDER_TO_IDLE_POS_FAILED | Error |  |  |  | The feeder could not be moved correctly to the idle position. | Meaning: The linear transport could not be moved correctly to the idle position. Reason: The movement was not controlled correctly or there are mechanical problems Cure: 1) Repeat the stall calibration of the SM Linear Transport 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the linear transport. 3) Replace SM Linear Transport, if error persists.
10251 | MOVE_FEEDER_TO_PRE_SNAP_POS_FAILED | Error |  |  |  | The feeder could not be moved correctly to the pre-snap position. | Meaning: The linear transport could not be moved correctly to the pre-snap position. Reason: The movement was not controlled correctly or there are mechanical problems Cure: 1) Repeat the stall calibration of the SM Linear Transport 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the linear transport. 3) Check Z-sensor actuation and cabling. 4) Replace SM Linear Transport, if error persists.
10252 | MOVE_FEEDER_TO_SNAP_POS_FAILED | Error |  |  |  | The feeder could not be moved correctly to the snap position. | Meaning: The linear transport could not be moved correctly to the snap position. Reason: The movement was not controlled correctly or there are mechanical problems Cure: 1) Repeat the stall calibration of the SM Linear Transport 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the linear transport. 3) Check Z-sensor actuation and cabling. 4) Replace Z-Sensor, if error perists. 5) Replace SM Linear Transport, if error persists.
10253 | MOVE_FEEDER_OR_WAGON_TO_STAGE_FAILED | Error |  |  |  | The feeder / wagon could not be moved correctly to the stage. | Meaning: The feeder could not be moved correctly to the stage. Reason: The movement was not controlled correctly or there are mechanical problems Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the paths of the IP transport and the wagon.
10254 | MOVE_FEEDER_TO_ERASURE_START_POS_FAILED | Error |  |  |  | The feeder could not be moved correctly to the erasure start pos. | Meaning: The linear transport could not be moved correctly to the erasure start pos. Reason: The movement was not controlled correctly or there are mechanical problems Cure: 1) Repeat the stall calibration of the SM Linear Transport 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the linear transport. 3) Replace SM Linear Transport, if error persists.
10255 | MOVE_FEEDER_TO_ERASURE_END_POS_FAILED | Error |  |  |  | The feeder could not be moved correctly to the erasure end pos. | Meaning: The linear transport could not be moved correctly to the erasure end pos. Reason: The movement was not controlled correctly or there are mechanical problems Cure: 1) Repeat the stall calibration of the SM Linear Transport 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the linear transport. 3) Replace SM Linear Transport, if error persists.
10256 | MOVE_WAGON_TO_HOME_POS_FAILED | Error |  |  |  | The wagon could not be moved correctly to the home pos. | Meaning: The linear transport could not be moved correctly to the home position. Reason: The movement was not controlled correctly or there are mechanical problems Cure: 1) Repeat the stall calibration of the SM Linear Transport 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the Linear Transport (including belts). 3) Check Linear Transport Home Position Sensor. 4) Replace Linear Transport Home Position Sensor if error persists. 5) Replace SM Linear Transport if error persists.
10257 | MOVE_FEEDER_BLIND_DOWN_FAILED | Error |  |  |  | The movement 'blind down' could not be accomplished correctly. | Meaning: The movement 'blind down' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the paths of the IP transport and the wagon, (iii) the LB 'Linear Transport Home' .
10258 | MOVE_FEEDER_UP_FOR_RETOUCH_FAILED | Error |  |  |  | The movement 'move up for retouch' could not be accomplished correctly. | Meaning: The movement 'move up for retouch' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the IP transport.
10259 | MOVE_FEEDER_RELEASE_SNAP_HOOKS_FROM_Z_FAILED | Error |  |  |  | The movement 'release snap hooks from IP' could not be accomplished correctly. | Meaning: The movement 'release snap hooks from IP' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the IP transport.
10260 | MOVE_FEEDER_UP_TEST_IP_FAILED | Error |  |  |  | The movement 'move feeder up for testing if IP at snaphooks' could not be accomplished correctly. | Meaning: The movement 'move feeder up for testing if IP at snaphooks' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the IP transport.
10261 | MOVE_SNAP_HOOKS_TO_HOME_POS_FAILED | Error |  |  |  | The snap hooks could not be moved to the home position correctly. | Meaning: The SM Snap hooks could not move the snap hooks correctly to the home position. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM Snap hooks 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the Snap Hooks. 3) Replace SM Snap Hooks if error persists.
10262 | MOVE_SNAP_HOOKS_TO_STOPPER2_FAILED | Error |  |  |  | The snap hooks could not be moved correctly to the 'stopper2' position. | Meaning: The snap hooks could not be moved correctly to the 'stopper2' position. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the snap hooks.
10263 | UNGRIP_SNAP_HOOKS_FAILED | Error |  |  |  | The snap hooks could not be ungripped correctly. | Meaning: The snap hooks could not be ungripped correctly from the image plate. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Check SNAP HOOK GRIP lightbarriers (left and right). 2) Check for mechanic problems blocking the movement of the Snaphooks. 3) Replace SNAP HOOK GRIP lightbarriers if error persists. 4) Replace traverse with snaphooks if error persists.
10264 | GRIP_SNAP_HOOKS_WITH_LB_FAILED | Error |  |  |  | The snap hooks could not be gripped with LB correctly. | Meaning: Light barrier indicates that IP is not gripped correctly. Reason: The image plate could not be gripped with the snaphooks correctly. Cure: Check 1) the mechanics of the snaphooks, 2)the light barriers of the snaphooks.
10265 | POWER_GRIP_SNAP_HOOKS_FAILED | Error |  |  |  | The movement 'power grip snap hooks' could not be accomplished correctly. | Meaning: The movement to engage the snap hooks with the image plate could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM_Snaphooks 3 times. Check values for consistency (+/- 2 digits). 2) Check Snaphook grip sensors (right and left) 3) check for damages to the snaphook mechanics 4) check for damages on the image plate connectors 5) Replace sensors, if error persists.
10266 | PREPARE_SNAP_HOOKS_FOR_RETURN_IP_FAILED | Error |  |  |  | The snap hooks could not be prepared for returning the IP. | Meaning: The snap hooks could not be prepared for returning the IP. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the snap hooks.
10267 | FIND_REFERENCE_SNAP_HOOKS_FAILED | Error |  |  |  | The reference position of the snap hooks could not be found correctly. | Meaning: The reference position of the snap hooks could not be found correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the snap hooks.
10268 | UNFIX_SNAP_HOOKS_ON_LB_FAILED | Error |  |  |  | The movement 'unfix snap hooks on LB' could not be accomplished correctly. | Meaning: The movement 'unfix snap hooks on LB' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the snap hooks.
10269 | UNFIX_SNAP_HOOKS_COMPLETELY_FAILED | Error |  |  |  | The movement 'unfix snap hooks completely' could not be accomplished correctly. | Meaning: The movement 'unfix snap hookscompletely' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the snap hooks.
10270 | FIND_REFERENCE_CARRIER_FAILED | Error |  |  |  | The reference position of the carrier could not be found correctly. | Meaning: The reference position of the carrier could not be found correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the carrier, (iii) the reference positon light barriers of the carrier.
10271 | MOVE_CARRIER_TO_COUPLE_POSITION_FAILED | Error |  |  |  | The carrier could not be moved correctly to the couple position. | Meaning: The carrier could not be moved correctly to the couple position (to the traverse). Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM Carrier Drive 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanical problems blocking the movement of the Carrier Drive.
10272 | MOVE_CARRIER_TO_UNCOUPLE_POSITION_FAILED | Error |  |  |  | The carrier could not be moved correctly to the uncouple position. | Meaning: The carrier could not be moved correctly to the uncouple position (to the traverse). Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM Carrier Drive 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanical problems blocking the movement of the Carrier Drive.
10273 | MOVE_CARRIER_TO_FOCUS_LEVEL_FAILED | Error |  |  |  | The carrier could not be moved correctly to the focus level. | Meaning: The carrier could not be moved correctly to the focus level. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the value of the focus level, (iii) the path of the carrier.
10274 | MOVE_CARRIER_FREE_IP_TRANSPORT_FAILED | Error |  |  |  | The movement 'free IP transport' could not be accomplished correctly. | Meaning: The movement of the IP Carrier to 'free IP transport' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM Carrier Drive 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanical problems blocking the movement of the Carrier Drive.
10275 | MOVE_CARRIER_OUT_ON_STOPPER_FAILED | Error |  |  |  | The movement 'move out on stopper' could not be accomplished correctly. | Meaning: The movement 'move out on stopper' could not be accomplished correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the carrier.
10276 | FIND_REFERENCE_CONNECTOR_FAILED | Error |  |  |  | The reference position of the connector could not be found correctly. | Meaning: The reference position of the connector could not be found correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the connector.
10277 | OPEN_LINK_CONNECTOR_FAILED | Error |  |  |  | The link (connector) could not be opened correctly. | Meaning: The connector could not be opened correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM Connector 3 times. Check values for consistency (+/- 2 digits). 2) Check for mechanic problems blocking the movement of the Connector (including belts). 3) Replace SM Connector if error persists.
10278 | CLOSE_LINK_CONNECTOR_FAILED | Error |  |  |  | The link (connector) could not be closed correctly. | Meaning: The link (connector) could not be closed correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the connector.
10279 | FIND_REFERENCE_SWIVEL_FAILED | Error |  |  |  | The reference position of the swivel could not be found correctly. | Meaning: The reference position of the swivel could not be found correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the swivel.
10280 | LIFT_OPTICS_FAILED | Error |  |  |  | The optics could not be lifted correctly. | Meaning: Stall control indicates that the movement of the stepper motor SWIVEL DRIVE may be insufficient. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check 1) the swivel mechanics for smooth movement, 2) repeat the stall calibration, 3) replace stepper motor SWIVEL DRIVE, if error persists.
10281 | LOWER_OPTICS_FAILED | Error |  |  |  | The optics could not be lowered correctly. | Meaning: The light collector could not be lowered correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the swivel.
10282 | FIND_REFERENCE_VALVE_FAILED | Error |  |  |  | The reference position of the valve could not be found correctly. | Meaning: The reference position of the vacuum valve could not be found correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the valve .
10283 | RELEASE_VACUUM_FAILED | Error |  |  |  | The vacuum could not be released correctly. | Meaning: The vacuum could not be released correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the valve.
10284 | OPEN_SMALL_FORMATS_FAILED | Error |  |  |  | The valve could not be opened for small formats correctly. | Meaning: The vacuum valve could not be opened for lsmall formats correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM_VACUUM_VALVE 3 times. Check values for consistency (+/- 2 digits). 2) Replace the vacuum valve if error persists.
10285 | OPEN_LARGE_FORMATS_FAILED | Error |  |  |  | The valve could not be opened for large formats correctly. | Meaning: The vacuum valve could not be opened for large formats correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: 1) Repeat the stall calibration of the SM_VACUUM_VALVE 3 times. Check values for consistency (+/- 2 digits). 2) Replace the vacuum valve if error persists.
10286 | RELEASE_VACCUM_FAILED | Error |  |  |  | The vacuum could not be released correctly. | Meaning: The vacuum could not be released correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the valve.
10287 | CLOSE_ALL_FAILED | Error |  |  |  | The valves could not be closed correctly. | Meaning: The vacuum valve could not be closed correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the valve.
10288 | MEASURE_VACUUM_FAILED | Error |  |  |  | The vacuum could not be measured correctly. | Meaning: The vacuum could not be measured correctly. Reason: The measurement was not performed correctly or there are electronic or mechanical problems. Cure: Check (i) the corresponding measurement procedures (SW), (ii) the pressure sensors and the valves.
10289 | MOVE_SLOW_SCAN_TO_HOME_POSITION_FAILED | Error |  |  |  | The slow scan could not be moved to the home position correctly. | Meaning: The slow scan could not be moved to the home positon correctly after the scan. Reason: The movement was not controlled correctly or there are mechanical or electronic problems. Cure: 1) Check for mechanic problems blocking the movement of the Slow Scan. 2) Check Slow Scan Reference sensor motor and cabling. 3) Replace defective part, reference sensor, motor, cabeling or IPHandling Board, if error persists.
10290 | START_SLOW_SCAN_FAILED | Error |  |  |  | The slow scan could not perform the movement for scanning the IP correctly. | Meaning: The slow scan could not perform the movement for scanning the IP correctly. Reason: Probably any electronic problems. Cure: Restart the system. Replace IP handling board if error persists.
10291 | MOVE_SLOW_SCAN_HOME_AFTER_SCAN_FAILED | Error |  |  |  | The slow scan could not be moved to the home positon correctly after the scan. | Meaning: The slow scan could not be moved to the home positon correctly after the scan. Reason: The movement was not controlled correctly or there are mechanical or electronic problems. Cure: 1) Check for mechanic problems blocking the movement of the Slow Scan. 2) Check Slow Scan Reference sensor motor and cabling. 3) Replace defective part, reference sensor, motor, cabeling or IPHandling Board, if error persists.
10292 | MOVE_SLOW_SCAN_TO_CALIBRATION_POSITION_FAILED | Error |  |  |  | The slow scan could not be moved to the calibration positon correctly (for calibration of the optics). | Meaning:The slow scan could not be moved to the calibration positon correctly (for calibration of the optics). Reason: The movement was not controlled correctly or there are mechanical or electronic problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the slow scan, (iii) the mechanics and electronics of the slow scan.
10293 | TEST_SLOWSCAN_UPPER_END_CTL_LB_FAILED | Error |  |  |  | The selftest of Slow Scan Home LB failed. | Meaning:The selftest of the light barrier SLOWSCAN REFERENCE failed. Reason: The light barrier is disconnected or defective. Cure: 1) Check corresponding cables and connections. 2) Replace light barrier if error persists.
10294 | TEST_LINEAR_TRANSPORT_HOME_LB_FAILED | Error |  |  |  | The selftest of Linear Transport Home LB failed. | Meaning:The selftest of the light barrier LINEAR TRANSPORT HOME POSITION failed. Reason: The light barrier is disconnected or defective. Cure: 1) Check corresponding cables and connections. 2) Replace light barrier if error persists.
10295 | TEST_LINEAR_TRANSPORT_Z_LB_FAILED | Error |  |  |  | The selftest of Linear Transport Z LB failed. | Meaning:The selftest of the light barrier Z-SENSOR failed. Reason: The light barrier is disconnected or defective. Cure: 1) Check corresponding cables and connections. 2) Replace light barrier if error persists.
10296 | TEST_LINEAR_TRANSPORT_TRAVERSE_ON_STAGE_LB_FAILED | Error |  |  |  | The selftest of Linear Transport 'Traverse on Stage' LB failed. | Meaning:The selftest of the TRAVERSE ON IP-CARRIER failed. Reason: The light barrier is disconnected or defective. Cure: 1) Check corresponding cables and connections. 2) Replace light barrier if error persists.
10297 | TEST_LINEAR_TRANSPORT_TRAVERSE_LOCKED_LB_FAILED | Error |  |  |  | The selftest of Linear Transport 'Traverse Locked' LB failed. | Meaning:The selftest lightbarrier TRAVERSE LOCKED ON CARRIAGE failed. Reason: The lightbarrier is disconnected or defective. Cure: 1) Check corresponding cables and connections. 2) Replace light barrier if error persists.
10298 | TEST_CARRIER_HOME_LB_FAILED | Error |  |  |  | The selftest of Carrier Home LB failed. | Meaning:The selftest of the lightbarrier CARRIER_DRIVE HOME POSITION COARSE failed. Reason: The light barrier is disconnected or defective. Cure: 1) Check corresponding cables and connections. 2) Replace light barrier if error persists.
10299 | TEST_CARRIER_REF_POS_FINE_LB_FAILED | Error |  |  |  | The selftest of Carrier 'Ref Pos Fine' LB failed. | Meaning:The selftest of the light barrier CARRIER_DRIVE HOME POSITION FINE failed. Reason: The light barrier is disconnected or defective. Cure: 1) Check corresponding cables and connections. 2) Replace light barrier if error persists.
10300 | VACUUM_MEASUREMENT_FAILED | Error |  |  |  | The measurement of the vacuum failed. | Meaning:The measurement of the vacuum failed. Reason: (i) timeout of the AD Converter during measurement, (ii) pressure value out of the required range (value returned by the AD Converter is too high). Cure: Check the sensors and corresponding electronics.
10301 | VACUUM_BELOW_THRESHOLD | Error |  |  |  | The vacuum is below the defined threshold. | Meaning:The vacuum is below the defined threshold. Reason: (i) IP is not fixed correctly, (ii) the corresponding vacuum sensor or the vacuum unit do not work correctly. Cure: Check (i) the IP, (ii) the sensors and the vacuum unit.
10302 | TEST_VACUUM_UNIT_FAILED | Error |  |  |  | The selftest of the vacuum unit failed. | Meaning: The selftest of the vacuum unit failed; the pumping capacity is not high enough. Reason: (i) The vacuum unit is not conected correctly or does not work correctly, (ii) The pressure sensors do not work correctly. Cure: Check (i) the corresponding cables and connections, (ii) the pressure sensors, (iii) the vacuum unit.
10303 | STATUS_SENSOR_ROPE_NOK | Error |  |  |  | At least one of the ropes of the Linear Transport is mechanically not operable. | Meaning: At least one of the ropes of the Linear Transport is mechanically not operable. Therefore the device is stopped immediately Reason: (i) At least one of the ropes of the Linear Transport is mechanically not operable, (ii) the 'rope' sensor is not connected, (iii) the 'rope' sensor defective. Cure: Check (i) the ropes of the LinearTransport, (ii) the 'rope' sensor on the IP Handling Board.
10304 | LINEAR_TRANSPORT_Z_LB_WRONG_STATUS_BEFORE_SCAN | Error |  |  |  | The Linear Transport Z LB has the wrong status before the scan. | Meaning: The Linear Transport Z LB has the wrong status before the scan ('interrupted' instead of 'free'). Reason: The Linear Transport Z LB is disconnected or defect Cure: Check corresponding cables and connections, check LB.
10305 | LINEAR_TRANSPORT_TRAVERSE_ON_STAGE_LB_WRONG_STATUS_BEFORE_SCAN | Error |  |  |  | The Linear Transport 'Traverse on Stage' LB has the wrong status before the scan. | Meaning: The Linear Transport 'Traverse on Stage' LB has the wrong status before the scan ('interrupted' instead of 'free'). Reason: The Linear Transport 'Traverse on Stage' LB is disconnected or defect Cure: Check corresponding cables and connections, check LB.
10306 | LINEAR_TRANSPORT_TRAVERSE_LOCKED_LB_WRONG_STATUS_BEFORE_SCAN | Error |  |  |  | The Linear Transport 'Traverse Locked' LB has the wrong status before the scan. | Meaning: The Linear Transport 'Traverse Locked' LB has the wrong status before the scan ('interrupted' instead of 'free'). Reason: The Linear Transport 'Traverse Locked' LB is disconnected or defect Cure: Check corresponding cables and connections, check LB.
10307 | LINEAR_TRANSPORT_HOME_LB_WRONG_STATUS_BEFORE_SCAN | Error |  |  |  | The Linear Transport home LB has the wrong status before the scan. | Meaning: The Linear Transport home LB has the wrong status before the scan ('interrupted' instead of 'free'). Reason: The Linear Transport home LB is disconnected or defect Cure: Check corresponding cables and connections, check LB.
10308 | SNAP_HOOKS_LBS_WRONG_STATUS_BEFORE_SCAN | Error |  |  |  | The Snap Hooks LBs have the wrong status before the scan. | Meaning: The Snap Hooks LBs have the wrong status before the scan ('interrupted' instead of 'free'). Reason: The Snap Hooks LBs are disconnected or defect Cure: Check corresponding cables and connections, check LBs.
10309 | LINEAR_TRANSPORT_Z_LB_WRONG_STATUS_BEFORE_ERASURE_OF_IP | Error |  |  |  | The Linear Transport Z LB has the wrong status before the erasure of the IP. | Meaning: The Linear Transport Z LB has the wrong status before thethe erasure of the IP ('interrupted' instead of 'free'). Reason: The Linear Transport Z LB is disconnected or defect Cure: Check corresponding cables and connections, check LB.
10310 | PREPARE_SNAP_HOOKS_FOR_GRIP_FAILED | Error |  |  |  | The snap hooks could not be prepared for gripping the IP in cassette. | Meaning: The snap hooks could not be prepared for gripping the IP in the cassette. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the snap hooks.
10311 | AUTO_ADJUST_CARRIER_COUPLE_POS_FAILED | Error |  |  |  | The automatic adjustment of carrier couple correction position failed. | Meaning: The automatic adjustment of carrier couple correction position failed. The range of carrier positions where Couple/Uncouple procedure of waggon works, is too small. Reason: Any mechanical disadjustment or fault at components IP-stage, adapter plate, traverse or waggon. Cure: Check and repair mechanics.
10312 | WRITE_ELABEL_ENTRY_IPHDL_FAILED | Error |  |  |  | A value could not be writte correctly to the Elabel of the IP Handling Board. | Meaning: Error during writing a value in the Elabel of the IP Handling Board. Reason: The IP Handling Unit writes the Elabel value with an error condition. Cure: Restart device, if the error persists, replace the IP Handling board.
10313 | SNAPHOOKS_LBCTRL_LEFT_WRONG_STATUS_INTERRUPTED | Error |  |  |  | The Snap Hooks left control LB has the wrong status Interrupted.. | Meaning: The Snap Hooks left control long distance LB has the wrong status 'interrupted' instead of 'free'. Reason: The Snap Hooks LB is disconnected or defective. Cure: Check corresponding cables and connections, check LB.
10314 | SNAPHOOKS_LBCTRL_RIGHT_WRONG_STATUS_INTERRUPTED | Error |  |  |  | The Snap Hooks right control LB has the wrong status Interrupted.. | Meaning: The Snap Hooks right control long distance LB has the wrong status 'interrupted' instead of 'free'. Reason: The Snap Hooks LB is disconnected or defective. Cure: Check corresponding cables and connections, check LB.
10315 | CLEANUP_LB_TRAVERSE_ON_STAGE_NOT_REACHED | Error |  |  |  | LB travers on stage not interrupted by traverse in Cleanup cycle. | Meaning: The LB travers on stage becomes not interrupted in Cleanup cycle, when moving traverse down to stage. Reason: (i) The traverse is blocked mechanically. (ii) The traverse reaches end position at stage but LB is not interrupted anyhow. (iii) LB is interrupted mechanically but remains in state FREE. Cure: (i) Check linear transport + traverse mechanics. (ii) Check positions of LB-flag at traverse and of LB itself. (iii) Check LB plus cabeling on electric functionality and replace defect part(s).
10316 | CLEANUP_LB_LINTRA_REFERENCE_NOT_REACHED | Error |  |  |  | Reference LB of Linear transport not interrupted by waggon in Cleanup cycle. | Meaning: The Reference LB of Linear transport not interrupted by waggon in Cleanup cycle. Reason: (i) Uncouple waggon from traverse not possible, carrier uncouple position not OK. (ii) Couple waggon to traverse not possible, snaphooks relief lever at wrong position. (iii) The waggon moves down but LB is not interrupted anyhow. (iiii) LB is interrupted mechanically but remains in state FREE. Cure: (i) Do Autoadjustment of carrier couple position. (ii) Check if snaphooks relief lever is versatile and always returns to initial position by spring force, and repair if not. (iii) Check positions of LB-flag at traverse and of LB itself. (iiii) Check LB plus cabeling on electric functionality and replace defect part(s).
10317 | CLEANUP_LB_TRAVERSE_LOCKED_NOT_REACHED | Error |  |  |  | LB traverse locked not interrupted by traverse/waggon in Cleanup cycle. | Meaning: The LB travers locked becomes not interrupted in Cleanup cycle, when moving waggon up to stage. Reason: (i) Couple waggon to traverse not possible, carrier couple position not OK. (ii) Couple waggon to traverse not possible, snaphooks relief lever at wrong position. (iii) Couple waggon to traverse is OK, but LB is not interrupted anyhow. Cure: (i) Do Autoadjustment of carrier couple position. (ii) Check if snaphooks relief lever is versatile and always returns to initial position by spring force, and repair if not. (iii) Check positions of LB-flag at waggon and of locked LB itself. Check LB plus cabeling on electric functionality and replace defect part(s).
10318 | CLEANUP_TRAVERSE_LOST | Error |  |  |  | Status of traverse is undefined, probably lost during CleanUp. Sensors show traverse on stage, unlocked from waggon and carrier not in unlock position to fix traverse with magnets. | Meaning: Control Sensors show traverse on stage, unlocked from waggon and carrier not in unlock position to fix traverse with magnets. Therefore status of traverse is undefined, this is a deadlock situation. Reason: (i) Coarse Home LB of carrier shows wrong status 'interrupted' and stage is in uncouple position. (ii) Reference LB of linear transport shows wrong status 'interrupted' by waggon. (iii) Status of traverse is really undefined, uncoupled and not fixed at traverse with magnets Cure: (i) Check/Replace Coarse Home-LB of carrier drive with cables. (ii) Check/Replace Reference-LB of linear transport drive with cables. (iii) Return digitizer manually to well defined status: move carrier to uncouple position, couple traverse with waggon, move carrier backwards to reference position.
10319 | CLEANUP_LB_TRAVERSE_AND_ZSENSOR_WRONG_STATE | Error |  |  |  | LB traverse on stage and LB Z-Sensor are both interrupted in Cleanup cycle. this is an invalid status. | Meaning: LB traverse on stage and LB Z-Sensor are both interrupted in Cleanup cycle. this is an invalid status. Reason: (i) LB traverse on stage defective if not interrupted. (ii) LB Z-Sensor defective if not interrupted. Cure: (i) Check LB traverse on stage plus cabeling on electric functionality and replace defect part(s). (ii) Check LB Z-sensor plus cabeling on electric functionality and replace defect part(s).
10320 | MOVE_FEEDER_DOWN_Z_FREE_FAILED | Error |  |  |  | The movement 'feeder down to free Z-Sensor' could not be accomplished correctly. | Meaning: The movement 'feeder down to free Z-Sensor' could not be accomplished correctly. Reason: (i) Traverse is mechanically blocked and cannot move. (ii) Z-Sensor is already free but status remains interrupted. (iii) Linear transport mechanics or electronic defective. Cure: (i) Check cause of blockade and remove. (ii) Check functionality of Z-Sensor and cabeling and remove if necessary. (iii) Check mechanics and electronics(motor, cabeling,IP-handlingBoard) of linear transport and replace if defective.
10321 | RELEASE_SNAP_HOOKS_RETURN_FAILED | Error |  |  |  | The snap hooks could not move to release position correctly, after IP returned into cassette . | Meaning: The snap hooks could not move to release position correctly, after IP returned into cassette . Reason: (i) The movement was not controlled correctly or (ii) there are electrical problems. Cure: (i) Fix bug in Control Software.(ii) Check snaphooks motor, cabeling and IPHandling Board and replace if necessary.
10322 | RETURN_IP_FAILED | Error |  |  |  | The procedure 'ReturnIP' failed. | Meaning: The procedure 'ReturnIP' failed. Reason: mechanics or mechanical movements of IP transfer between IP handling and cassette handling units do not work correctly. Cure: Check movements of SnapHooks, LinearTransport or mechanics of IP transfer between IP handling and cassette handling units
10323 | RELEASE_HOOKS_RETURN_FAILED | Error |  |  |  | The snap hooks could not move to release position correctly, after IP returned into cassette . | Meaning: The snap hooks could not move to release position correctly, after IP returned into cassette . Reason: (i) The movement was not controlled correctly or (ii) there are electrical problems. Cure: (i) Fix bug in Control Software.(ii) Check snaphooks motor, cabeling and IPHandling Board and replace if necessary.
10324 | FIND_REFERENCE_CONNECTOR_FAILED_MECHANICS | Error |  |  |  | The reference position of the connector could not be found correctly. | Meaning: The reference position of the connector could not be found correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW and stall parameters), (ii) the path of the connector, (iii) the belt of the connector motor
10325 | CLOSE_LINK_CONNECTOR_FAILED_MECHANICS | Error |  |  |  | The link (connector) could not be closed correctly. | Meaning: The link (connector) could not be closed correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the connector, (iii) the belt of the connector motor.
10326 | CONNECTOR_LOW_LEVEL_MOVEMENT_FAILED | Error |  |  |  | The link (connector) could not be moved correctly. | Meaning: The link (connector) could not be moved correctly. Reason: The movement was not controlled correctly or there are mechanical problems. Cure: Check (i) the corresponding movement control (SW), (ii) the path of the connector, (iii) the belt of the connector motor.
10327 | IP_NOT_DETECTED_WITH_LEFT_LB | Error |  |  |  | IP not detected with left snaphooks control LB during transport. | Meaning: Image Plate not detected with left snaphooks control lightbarrier during transport. Reason: (i) IP lost because not gripped correctly by snaphooks. (ii) light barrier hardware defective Cure: (i) Remove IP manually from digitizer and make restart. (ii) Check hardware (left snaphooks control LB,Cable,IPHandling-Board) and replace if defective.
10328 | IP_NOT_DETECTED_WITH_RIGHT_LB | Error |  |  |  | IP not detected with right snaphooks control LB during transport. | Meaning: Image Plate not detected with right snaphooks control lightbarrier during transport. Reason: (i) IP lost because not gripped correctly by snaphooks. (ii) light barrier hardware defective Cure: (i) Remove IP manually from digitizer and make restart. (ii) Check hardware right snaphooks control LB,Cable,IPHandling-Board) and replace if defective.
10329 | TEST_TRANSPORT_LOCK_ATTACHED | Error | 282 | Transport locks still attached. Remove transport locks. |  | Transport lock attached at start up. | Meaning: The mobile kit is mounted and the transport lock is attached. Reason: User forgot to remove the transport lock. Cure: Remove transport lock and restart the digitizer.
10330 | TEST_MOBILE_CODER_FAILED | Error |  |  |  | Inconsistent sensor configuration | Meaning: The coder of the mobile kit is not connected, while the transport lock light barrier is. Reason: Defective hardware. Cure: Check Check connector at P31 of the IP handling board and the cable.
10331 | TEST_TRANSPORT_LOCK_LB_FAILED | Error |  |  |  | Inconsistent sensor configuration | Meaning: The lightbarrier of the transport lock is not connected, while the mobile coder is. Reason: Defective hardware or unplugged light barrier. Cure: Check connector at P31 of the IP handling board, the cable and the connection to the lightbarrier.
10332 | INTERNAL_ELABEL_INIT_FAILED | Error |  |  |  | Elabel initialisation failed | Meaning: Elabel data not accessible or inconsistent Reason: Defective hardware or software failure. Cure: 1) Restart the machine. 2) Reload the software. 3) Replace the IP handling board.
10333 | INTERNAL_ELABEL_INIT_MISSING_ENTRY_DETECTED | Info |  |  |  | Elabel initialisation detected a missing entry | Meaning: The data in the Elabel is not complete Reason: Software upgrade, Elabel Data not up to date. Cure: 1) Perform Stall Calibration of motor SM Connector.
10334 | VALVE_LOW_LEVEL_MOVEMENT_FAILED | Error |  |  |  | Any Valve drive movement was not peformed correctly. | Meaning: Valve movement failed. Reason: Any electrical or mechanical problem of valve drive. Cure: Check the valve drive if all mechanical parts are moving and mounted correctly, and if the motor is moving at all. Replace defective part(s).
10335 | SWIVEL_LOW_LEVEL_MOVEMENT_FAILED | Error |  |  |  | Any Swivel drive movement was not peformed correctly. | Meaning: Swivel drive movement failed. Reason: Any electrical or mechanical problem of swivel drive. Cure: Check the swivel drive if all mechanical parts are moving and mounted correctly, and if the motor is moving at all. Replace defective part(s).
10336 | LINTRA_LOW_LEVEL_MOVEMENT_FAILED | Error |  |  |  | Any LinearTransport drive movement was not peformed correctly. | Meaning: LinearTransport drive movement failed. Reason: Any electrical or mechanical problem of LinearTransport drive. Cure: Check the LinearTransport drive if all mechanical parts are moving and mounted correctly, and if the motor is moving at all. Replace defective part(s).
10337 | SNAPHOOKS_LOW_LEVEL_MOVEMENT_FAILED | Error |  |  |  | Any Snaphooks drive movement was not peformed correctly. | Meaning: Snaphooks drive movement failed. Reason: Any electrical or mechanical problem of Snaphooks drive. Cure: Check the Snaphooks drive if all mechanical parts are moving and mounted correctly, and if the motor is moving at all. Replace defective part(s).
10338 | CARRIER_LOW_LEVEL_MOVEMENT_FAILED | Error |  |  |  | Any Carrier drive movement was not peformed correctly. | Meaning: Carrier drive movement failed. Reason: Any electrical or mechanical problem of Carrier drive. Cure: Check the Carrier drive if all mechanical parts are moving and mounted correctly, and if the motor is moving at all. Replace defective part(s).
10339 | VALVE_IRUNREF_OUTRANGE | Warning |  |  |  | Valve adjustable run current for reference positioning out of range. Default value is used. | Meaning: Valve adjustable run current for reference positioning out of range. Default value is used. Reason: Cure:
10340 | SWIVEL_IRUNLIFT_OUTRANGE | Warning |  |  |  | Swivel adjustable run current for lift optics out of range. Default value is used. | Meaning: Swivel adjustable run current for lift optics out of range. Default value is used. Reason: Cure:
10341 | LINTRA_IRUNDOWN_OUTRANGE | Warning |  |  |  | Linear transport adjustable run current for moving down stage out of range. Default value is used. | Meaning: Linear transport adjustable run current for moving down stage out of range. Default value is used. Reason: Cure:
10342 | LINTRA_IRUNUP_OUTRANGE | Warning |  |  |  | Linear transport adjustable run current for moving waggon up to stage out of range. Default value is used. | Meaning: Linear transport adjustable run current for moving waggon up to stage out of range. Default value is used. Reason: Cure:
10343 | CONN_IRUNOPEN_OUTRANGE | Warning |  |  |  | Connector adjustable run current for open link out of range. Default value is used. | Meaning: Connector adjustable run current for open link out of range. Default value is used. Reason: Cure:
10344 | CONN_IRUNCLOSE_OUTRANGE | Warning |  |  |  | Connector adjustable run current for close link out of range. Default value is used. | Meaning: Connector adjustable run current for close link out of range. Default value is used. Reason: Cure:
10345 | CARRIER_IRUNOUTSTOPPER_OUTRANGE | Warning |  |  |  | Carrier adjustable run current for out on stopper movement out of range. Default value is used. | Meaning: Carrier adjustable run current for out on stopper movement out of range. Default value is used. Reason: Cure:
10346 | CARRIER_POSCOUPLECORR_OUTRANGE | Warning |  |  |  | Carrier couple correction position out of range. Default value is used. | Meaning: Carrier couple correction position out of range. Default value is used. Reason: Cure:
10347 | CARRIER_POSCOUPLE_OUTRANGE | Warning |  |  |  | Carrier couple position out of range. Default value is used. | Meaning: Carrier couple position out of range. Default value is used. Reason: Cure:
10348 | CARRIER_POSFOCUSBASE_OUTRANGE | Warning |  |  |  | Carrier focus base position out of range. Default value is used. | Meaning: Carrier focus base position out of range. Default value is used. Reason: Cure:
10349 | CARRIER_POSUNCOUPLE_OUTRANGE | Warning |  |  |  | Carrier uncouple position out of range. Default value is used. | Meaning: Carrier uncouple position out of range. Default value is used. Reason: Cure:
10350 | CARRIER_POSFREEIPTRANSPORT_OUTRANGE | Warning |  |  |  | Carrier free IP transport position out of range. Default value is used. | Meaning: Carrier free IP transport position out of range. Default value is used. Reason: Cure:
10351 | CARRIER_POSOUTSTOPPER_OUTRANGE | Warning |  |  |  | Carrier out on stopper position out of range. Default value is used. | Meaning: Carrier out on stopper position out of range. Default value is used. Reason: Cure:
10752 | INCIDENCE_OF_LIGHT | Error |  |  |  | Incidence of light --> adj.PMM/PMT failed ; fatal! | Meaning: It is too bright to adjust the Photo Multiplier. Reason: Housing of digitizer is open. Cure: Close the digitizer housing.
10753 | PMM_ADJUST | Error |  |  |  | PMM/PMT adjust failed | Meaning: Photo multiplier malfunction. Reason: - Cure: Switch digitizer off/on and try a scan, if failed again replace PMT + LIGHT COLLECTOR.
10754 | GEO_CALIBRATION | Error |  |  |  | Geometric and jitter calibration failed | Meaning: Geometric and jitter calibration failed. Reason: Optic module, calibration board, sensor path misadjusted. Cure: Check proper seating of optic module, clean scan path with dust brush.
10755 | START_POLYGON | Error |  |  |  | Start polygon failed | Meaning: Polygon locked signal is not set. Reason: - Cable broken - Polygon(Optic Mudule) defective - PMI board defective - PMI POWER DISTRIBUTION board defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check the cabling of the following boards: - PMI BOARD - Optic Module - DISTRIBUTOR BOARD - PMI BOARD 3. Exchange DISTRIBUTOR BOARD, refer to the Service Manual for order and replacement 4. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement 5. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement
10756 | ADC_ADJUST | Error |  |  |  | ADC adjustment out of spec | Meaning: ADC offset adjust out of range. Reason: ADC defective or PMM/PMT offset voltage out of range. Cure: Switch digitizer off/on, replace PMI board, replace PMT + LIGHT COLLECTOR.
10757 | START_SCAN | Error |  |  |  | Start scan failed | Meaning: Start scan failed, because scan is already running Reason: Software error Cure: Update software or inform RDE
10758 | POLYGON_LOCKED | Error |  |  |  | Polygon get not locked | Meaning: Polygon can't reach the requested speed. Reason: Polygon defective. Cure: Replace optic unit.
10759 | PREPARE_SCAN | Error |  |  |  | Prepare Scan failed
10760 | TOO_LESS_DATA | Error |  |  |  | Too less data stored in SDRAM --> SDRAM filled with RAMP | Meaning: Too less data stored in SDRAM --> SDRAM filled with stripes. Reason: Electrostatic discharge, error in begin or end of line signal (BOL,EOL). Cure: Check grounding of calibration board and grounding of front cover.
10761 | TEST_SCAN | Error |  |  |  | Test scan failed --> BOL or EOL out of range | Meaning: One or both of the scanline control sensors in the Optic Module, for detecting begin of line (BOL) and end of line (EOL), does not detect the laser beam at all or get to less laser power. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at data cable between Optic Module and connector P2 at PMI-Board. 2) Contact problem(connector P2) or defect of the PMI-Board. 3) Misalignment of the Optic Module. 4) Contact problem or defect on the Optic Module. Cure: 1) Restart the Digitizer and check, if the problem persists. 2) Check the data cable between Optic Module and connector P2 at PMI-Board. Exchange cable if required, refer to the Service Manual for order and replacement. 3) Exchange the Optic Module, refer to the Service Manual for order and replacement. 4) Check and exchange PMI-Board if required, refer to the Service Manual for order and replacement.
10762 | PMM_REFCAL | Error |  |  |  | Calculation of REFCAL from IPF out of range | Meaning: The IP depended calculation of sensitivity factor results in a value out of range. Reason: Wrong PMM/PMT adjustment values stored in PMT eLabel. Cure: Replace PMT + LIGHT COLLECTOR.
10763 | PMM_AGING | Error |  |  |  | PMT sensitivity become critical --> reaches 94% = 750V | Meaning: The PMTs end is coming soon. Cure: Replace PMT + LIGHT COLLECTOR
10764 | PMM_FINE_ADJUST | Error |  |  |  | PMM/PMT fine adjust failed --> possibly bad image quality | Meaning: PMT sensitivity adjustment failed Reason: - PMI POWER DISTRIBUTION board defective - PMT defective or end of life Cure: 1. Replace PMI POWER DISTRIBUTION board 2. If error still persists replace PMT + LIGHT COLLECTOR
10765 | START_GEO_CALIBRATION | Error |  |  |  | Start geometric and jitter calibration failed
10766 | IPF_OUT_OF_RANGE | Error |  |  |  | IP sensitivity factor out of range | Meaning: IP is not compatible with digitizer Reason: IP does nit fit to this digitizer Cure: Use an IP, which is compatible to digitizer
10767 | TEST_SCAN_CAL | Error |  |  |  | Test scan failed: Not all CAL sensors detected --> defect or IP in digitizer | Meaning: Can't detect the expected amount of calibration signals. Reason: Optic path interrupted, IP in digitizer, broken light collector, cable PMI to calibration board defective, CAL-board defective Cure: Clean the scan line, check optic path, check cable PMI to calibration board, replace calibration board.
10768 | READ_PMM_ELABEL | Error |  |  |  | Can't read PMM settings from PMM eLabel |  | COM_PMMConnectionErr
10769 | READY_TO_SCAN_CAL_SENSORS | Error |  |  |  | Calibration sensor defective or two IP's in scanner | Meaning: Can't detect the expected amount of calibration signals. Reason: CAL-board defective or optic path interrupted (IP or dust). Cure: Clean the scan line, replace calibration board.
10770 | PMM_BLUE_LED | Error |  |  |  | Photo Multiplier Module (PMM/PMT), blue LED or power supply board defective | Meaning: Can't detect PMT blue LED. Reason: - PMT blue LED defective or not connected - PMI POWER DISTRIBUTION board defective - PMT defective Cure: 1. Check connection and cable from blue LED to PMT 2. Replace PMI POWER DISTRIBUTION board 3. If error still persists replace PMT + LIGHT COLLECTOR
10771 | FPGA_BUS | Error |  |  |  | FPGA BUS error detected | Meaning: FPGA BUS error occurred. Reason: unknown Cure: Restart digitizer, if error occurred again replace PMI board.
10772 | AUTOEOL_ADDED | Error |  |  |  | FPGA has added an EOL at end of EOL valid | Meaning: No EOL detected, FPGA adds one at end of EOL valid window. Reason: polygon speed out of range. Cure: Restart digitizer, replace optic module.
10773 | PMM_OFFSET_VOLTAGE | Error |  |  |  | PMM/PMT offset voltage out of range | Meaning: PMT offset voltage adjustment failed Reason: - Incidence of light - PMI POWER DISTRIBUTION board defective - PMT defective (end of life) Cure: 1. Make sure all covers are closed 2. Replace PMI POWER DISTRIBUTION board 3. If error still persists replace PMT + LIGHT COLLECTOR
10774 | IMAGE_DMA | Error |  |  |  | Image transfer error | Meaning: Image transfer failed Reason: - Hard- or Software error Cure: 1. Save Logfile and send it to RDE 2. Restart the Digitizer and check, if the problem persists
10775 | TOO_LESS_DATA_WARNING | Warning | 129 | Problem during scan. The image may not have the required quality. |  | Too less data stored in SDRAM --> SDRAM filled with RAMP | Meaning: Too less data stored in SDRAM --> SDRAM filled with stripes. Reason: Electrostatic discharge, error in begin or end of line signal (BOL,EOL). Cure: Check grounding of calibration board and grounding of front cover.
10776 | PMM_SPI_ACCESS | Error |  |  |  | Access to PMM SPI-BUS failed | Meaning: Can not get PMM SPI-BUS in time. Reason: -. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check the connector PMI - PMT 3. If error still persists replace PMT + LIGHT COLLECTOR | COM_PMMConnectionErr
10777 | PMM_REFLUM | Error |  |  |  | Adjust RefLum failed | Meaning: Can not adjust blue LED for the needed brightness. Reason: - blue LED defective - Phototransistor defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. If error still persists replace PMT + LIGHT COLLECTOR
10778 | IMAGE_DMA_OPEN | Error |  |  |  | Image transfer error | Meaning: Initializing image transfer failed Reason: - Hard- or Software error Cure: 1. Save Logfile and send it to RDE 2. Restart the Digitizer and check, if the problem persists
10779 | IMAGE_DMA_PEND | Error |  |  |  | Image transfer error | Meaning: sc_pend at beginning of image transfer failed Reason: - Hard- or Software error Cure: 1. Save Logfile and send it to RDE 2. Restart the Digitizer and check, if the problem persists
10780 | TEST_SCAN_CIR | Error |  |  |  | Test scan failed --> BOL or EOL out of range | Meaning: Test of the EOL / BOL sensors, which are located on the Calibration Board failed. The Digitizer is out of operation and requires service support. Reason: 1) The optical path of the laser beam is obstructed e.g. fibres in the dust cover, contamination on the pane or the sensors 2) Laser diode of the optic module is defective or the laser power is not sufficient 3) The Optic Module and/or the Calibration Board are not proper mounted 4) Contact problem or defect at data cables between CPMI Board (B410) and connector P3 of the CMAM Adapter Board or connector P1 of the CMAM Adapter Board and connector P1 of the Calibration Board (B412) or data cable between CPMI Board (B410) and connector P1 of the Calibration Board (B412) 5) Contact problem (connector P1) or defect of the Calibration Board (B412) 6) Contact problem (connector P1 and P3) or defect of the CMAM Adapter Board on the Optic Module (B317) 7) Contact problem (connector P2#22) or defect of the CPMI Board (B410) Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the presence of the laser beam on the pane of the scanner. If required replace the Optic Module, refer to the Service Manual for order and replacement 3) Check if the optical path of the laser beam is obstructed e.g. contamination on the EOL / BOL sensors, fibres in the dust cover 4) Check the Optic Module and the Calibration Board for proper fixation 5) Check the data cable and connectors between CPMI Board (B410, connector P2#22) and CMAM Adapter Board (connector P3) and the data cable and connectors between CMAM Adapter Board (connector P1) and the Calibration Board (B412, connector P1). If required replace the data cable, contact your Regional Support Network for further assistance 6) Exchange the Calibration Board (B412), refer to the Service Manual for order and replacement 7) Exchange the CMAM Adapter Board, contact your Regional Support Network for further assistance 8) Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement | TEST_SCAN_BOL_EOL_ERROR
10781 | PMM_VOLTAGE | Error |  |  |  | PMM/PMT set high voltage failed | Meaning: Photo multiplier malfunction. High voltage out of range Reason: PMM control unit defective Cure: Switch digitizer off/on, if failed again replace PMT + LIGHT COLLECTOR.
11264 | PC_TO_DIC_ADAPTER_RECEIVE_DEFAULT_NOT_POSSIBLE | Error |  |  |  | Adapter tried to receive a message using the 'receiveDefault' method, but it contains unsupported number of bytes |  | COM_NotNeeded_ForServiceId
11265 | BASE_APP_SEND_FAILED_NO_CONNECTION | Error |  |  |  | RTPort_send failed because connection to receiver is not established |  | COM_SWError
11266 | BASE_APP_SEND_FAILED_NO_MEMORY | Error |  |  |  | RTPort_send failed because no memory available (list of unused messageobjects exhausted) |  | COM_SWError
11267 | BASE_APP_SEND_FAILED_INVALID_PRIO | Error |  |  |  | RTPort_send failed because message priority is invalid |  | COM_SWError
11268 | BASE_APP_SEND_FAILED_OTHER_REASON | Error |  |  |  | RTPort_send failed for any other reason |  | COM_SWError
11776 | SM_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of SMx terminated with failure. | Reason: The Handling Motor SM terminated its selftest with an errror condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
11777 | SM_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of SMx terminated with UNPLUGGED failure. | Reason: The Handling Motor SM terminated its selftest with an errror condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
11778 | SM_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of SMx terminated with SIN COIL failure. | Reason: The Handling Motor SM terminated its selftest with an errror condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
11779 | SM_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of SMx terminated with COS COIL failure. | Reason: The Handling Motor SM terminated its selftest with an errror condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
11780 | SM_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of SMx terminated with DIVER failure. | Reason: The Handling Motor SM terminated its selftest with an errror condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
11781 | SM_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of SMx terminated with overcurrent to Ground detected failure. | Reason: The Handling Motor SM terminated its selftest with an errror condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
11782 | ADC_MEAS_TOT_FAILURE | Error |  |  |  | Selftest of SMx ended due to ADC measurement timeout failure. | Reason: The Handling Motor SM terminated its selftest with an errror condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
11876 | WARN_INVALID_PAR | Error |  |  |  | Warning: Test detected invalid par | Reason: The SM Test has a invalid par. Meaning: Retry possible, function disturbed. Cure: Restart, Call Service
11877 | WARN_INVALID_ACCESS_FUNC | Error |  |  |  | Warning: Test used invalid access par | Reason: The SM Test has a invalid par. Meaning: Retry possible, function disturbed. Cure: Restart, Call Service
11878 | WARN_INVALID_POINTER | Error |  |  |  | Warning: Test used invalid pointer | Reason: The SM Test has a invalid par. Meaning: Retry possible, function disturbed. Cure: Restart, Call Service
11879 | WARN_LOW_VOLT_MEAS_POS | Error |  |  |  | Warning: Test pos current, measured voltage value has low reserve to limit | Reason: The SM Test at driver voltage of an active pos current has a low reserve to the error limit. Meaning: Retry possible, function normal, risk for malfunction increased. Cure: Restart, check electronic and motor.
11880 | WARN_LOW_VOLT_MEAS_NEG | Error |  |  |  | Warning: Test neg current, measured voltage value has low reserve to limit | Reason: The SM Test at driver voltage of an active neg current has a low reserve to the error limit. Meaning: Retry possible, function normal, risk for malfunction increased. Cure: Restart, check electronic and motor.
11881 | WARN_LOW_SIN_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement sin value has low reserve to limit | Reason: The SM Test at sin current has a low reserve to the error limit. Meaning: Retry possible, function normal, risk for malfunction increased. Cure: Restart, check electronic and motor.
11882 | WARN_LOW_COS_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement cos value has low reserve to limit | Reason: The SM Test at cos current has a low reserve to the error limit. Meaning: Retry possible, function normal, risk for malfunction increased. Cure: Restart, check electronic and motor.
11883 | WARN_LOW_SIN_NEG_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement sin neg value has low reserve to limit | Reason: The SM Test at sin negative current has a low reserve to the error limit. Meaning: Retry possible, function normal, risk for malfunction increased. Cure: Restart, check electronic and motor.
11884 | WARN_LOW_COS_NEG_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement cos neg value has low reserve to limit | Reason: The SM Test at cos negative current has a low reserve to the error limit. Meaning: Retry possible, function normal, risk for malfunction increased. Cure: Restart, check electronic and motor.
12288 | FIXER_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of FIXER terminated with failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSETTEFIXATION terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12289 | FIXER_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of FIXER terminated with UNPLUGGED failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSETTEFIXATION terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12290 | FIXER_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of FIXER terminated with SIN COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSETTEFIXATION terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12291 | FIXER_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of FIXER terminated with COS COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSETTEFIXATION terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12292 | FIXER_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of FIXER terminated with DRIVER failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSETTEFIXATION terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12293 | FIXER_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of FIXER terminated with overcurrent to Ground detected failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSETTEFIXATION terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12294 | LOCKER_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of LOCKER terminated with failure. | Meaning: Failure in selftest. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12295 | LOCKER_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of LOCKER terminated with UNPLUGGED failure. | Meaning: Failure in selftest. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12296 | LOCKER_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of LOCKER terminated with SIN COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12297 | LOCKER_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of LOCKER terminated with COS COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12298 | LOCKER_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of LOCKER terminated with DRIVER failure. | Meaning: Failure in selftest. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12299 | LOCKER_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of LOCKER terminated with overcurrent to Ground detected failure. | Meaning: Failure in selftest. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12300 | OPENER_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of OPENER terminated with failure. | Reason: The Handling Motor OPENER terminated its selftest with an error condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, check cabling and motor. Call Service.
12301 | OPENER_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of OPENER terminated with UNPLUGGED failure. | Reason: The Handling Motor OPENER terminated its selftest with an error condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
12302 | OPENER_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of OPENER terminated with SIN COIL failure. | Meaning: The SM OPENER terminated its selftest with an error condition Reason: Cabling of motor or motor itself defective. Cure: 1) Check for mechanic problems blocking the movement of the opener drive. 2) Check cassette handling board. 3) Replace SM OPENER, if error persists. 4) Replace CASSETTE HANDLING BOARD, if error persists.
12303 | OPENER_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of OPENER terminated with COS COIL failure. | Meaning: The SM OPENER terminated its selftest with an error condition Reason: Cabling of motor or motor itself defective. Cure: 1) Check for mechanic problems blocking the movement of the opener drive. 2) Check cassette handling board. 3) Replace SM OPENER, if error persists. 4) Replace CASSETTE HANDLING BOARD, if error persists.
12304 | OPENER_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of OPENER terminated with DRIVER failure. | Reason: The Handling Motor OPENER terminated its selftest with an error condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
12305 | OPENER_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of OPENER terminated with overcurrent to Ground detected failure. | Meaning: The SM OPENER terminated its selftest with an error condition. Reason: Cabling of SM OPENER or IP handling board defective. Cure: 1) Check cabling of SM OPENER. 2) Check cassette handling board. 3) Replace cassette handling board, if error persists.
12306 | PUSHER_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of PUSHER terminated with failure. | Meaning: Failure in selftest. Reason: The Handling Motor PUSHER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12307 | PUSHER_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of PUSHER terminated with UNPLUGGED failure. | Meaning: Failure in selftest. Reason: The Handling Motor PUSHER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12308 | PUSHER_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of PUSHER terminated with SIN COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor PUSHER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12309 | PUSHER_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of PUSHER terminated with COS COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor PUSHER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12310 | PUSHER_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of PUSHER terminated with DRIVER failure. | Meaning: Failure in selftest. Reason: The Handling Motor PUSHER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12311 | PUSHER_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of PUSHER terminated with overcurrent to Ground detected failure. | Meaning: Failure in selftest. Reason: The Handling Motor PUSHER terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12312 | CASSTRANSPORT_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of CASSTRANSPORT terminated with failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSTRANSPORT terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12313 | CASSTRANSPORT_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of CASSTRANSPORT terminated with UNPLUGGED failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSTRANSPORT terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12314 | CASSTRANSPORT_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of CASSTRANSPORT terminated with SIN COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSTRANSPORT terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12315 | CASSTRANSPORT_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of CASSTRANSPORT terminated with COS COIL failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSTRANSPORT terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12316 | CASSTRANSPORT_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of CASSTRANSPORT terminated with DRIVER failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSTRANSPORT terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12317 | CASSTRANSPORT_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of CASSTRANSPORT terminated with overcurrent to Ground detected failure. | Meaning: Failure in selftest. Reason: The Handling Motor CASSTRANSPORT terminated its selftest with an error condition Cure: Restart device, check cabling of motor, replace motor if error persists.
12318 | OPENER_CLOSE_COVER_CAL_POS_FAILED | Error |  |  |  | Procedure closeCoverCal for calibration of OPENER terminated with detected failure. | Reason: The Handling Motor OPENER terminated its movement with an error condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
12319 | OPENER_OPEN_COVER_CAL_POS_FAILED | Error |  |  |  | Procedure openCoverCal for calibration of OPENER terminated with detected failure. | Reason: The Handling Motor OPENER terminated its movement with an error condition Meaning: Retry possible, function disturbed. Cure: Retry, Restart, Call Service.
12320 | OPENER_OPEN_POS_STEP_DEVIATION | Error |  |  |  | Opener Open Pos Movement terminated with failure. | Reason: Moved steps have deviation above the accepted limit or mechanical stop not found. Meaning: Function disturbed. Open Cover Pos not found. Cure: Restart Digitizer. If error persists please call Service.
12321 | OPENER_CLOSE_POS_STEP_DEVIATION | Error |  |  |  | Opener Close Pos Movement terminated with failure. | Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Meaning: Function disturbed. Close Cover position not found. Cure: Restart Digitizer. If error persists please call Service.
12322 | OPENER_CLOSE_POS_FAILED | Error |  |  |  | Opener Close Pos Movement terminated with failure. | Reason: Moving the motor to the Open-Cover-Position a unexpected maximum step limit is reached while waiting for sensoric input; Mechanical position can be blocked or Light Barrier malfunction. Mechanics, Cabling of motor, cabling of the light barrier, the light barrier part itself, the motor itself or the control board can be defective. Meaning: The movement to the motor position for an Open-Cover-Position is not reached in the expected way. No end signal was received from the Light-Barrier. Cure: 1) Check for mechanic problems blocking the movement of the opener drive. 2) Check the cabling to the Light-Barrier at the Open-Cover-Position, check the Light-Barrier at the Open-Cover-Position, check the cassette handling board, check the motor SM OPENER. 3) Replace the cabling to the Light-Barrier at the Open-Cover-Position, if error persists. 4) Replace the Light-Barrier at the Open-Cover-Position, if error persists. 5) Replace the SM OPENER, if error persists. 6) Replace CASSETTE HANDLING BOARD, if error persists.
12323 | OPENER_OPEN_POS_MAX_STEP_LIMIT | Error |  |  |  | Opener Open Pos Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop not found. Meaning: Function disturbed. Open Cover Pos not found. Cure: Restart Digitizer. If error persists please call Service.
12324 | OPENER_OPEN_POS_FAILED | Error |  |  |  | Opener Open Pos Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical position blocked or Light Barrier malfunction. Meaning: Function disturbed. Open Cover Pos not found at Light Barrier Sensor. Cure: Restart Digitizer. If error persists please call Service.
12325 | FIXER_RE_CLAMP_POS_FAILED | Error |  |  |  | Fixer Re-Clamp Pos Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop not found. Meaning: Function disturbed. Fixer Clamp Pos not found. Cure: Restart Digitizer. If error persists please call Service.
12326 | FIXER_RE_UNCLAMP_POS_FAILED | Error |  |  |  | Fixer Re-UnClamp Pos Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop not found. Meaning: Function disturbed. Fixer UnClamp Pos not found. Cure: Restart Digitizer. If error persists please call Service.
12327 | FIXER_UNCLAMP_POS_STEP_DEVIATION | Error |  |  |  | Fixer UnClampPos Movement terminated with failure. | Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Meaning: Function disturbed. UnClamp position not found. Cure: Restart Digitizer. If error persists please call Service.
12328 | FIXER_UNCLAMP_POS_MAX_STEP_LIMIT | Error |  |  |  | Fixer Home UnClamp Movement stopped due to maximum step limit | Reason: Maximum step limit reached. Meaning: Function disturbed. Unclamp-Pos not found. Cure: Restart Digitizer. If error persists please call Service.
12329 | FIXER_CLAMP_CASS_FAILED | Error |  |  |  | Fixer Clamp Cassette Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop not found. Meaning: Function disturbed. Fixer Clamp Pos not found. Cure: Restart Digitizer. If error persists please call Service.
12330 | LOCKER_LOCK_IP_MOVE_STOP_UNEXPECTED | Error |  |  |  | Locker home movement stopped too early. | Reason: The Locker home movement stopped early as stall reacted in position range before stopper. Meaning: Retry possible, function may be disturbed. Cure: Restart Digitizer. If error persists please call Service.
12331 | LOCKER_LOCK_IP_POS_STEP_DEVIATION | Error |  |  |  | Locker Lock Ip Pos Movement terminated with failure. | Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Meaning: Function disturbed. LockIp position not found. Cure: Restart Digitizer. If error persists please call Service.
12332 | LOCKER_LOCK_IP_POS_FAILED | Error |  |  |  | Locker Lock Ip Pos Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop not found. Meaning: Function disturbed. Lock Ip pos not found. Cure: Restart Digitizer. If error persists please call Service.
12333 | LOCKER_UNLOCK_IP_POS_STEP_DEVIATION | Error |  |  |  | Locker UnLock Ip Pos Movement terminated with failure. | Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Meaning: Function disturbed. UnLockIp position not found. Cure: Restart Digitizer. If error persists please call Service.
12334 | LOCKER_UNLOCK_IP_POS_FAILED | Error |  |  |  | Locker UnLock Ip Pos Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop not found. Meaning: Function disturbed. UnLock Ip pos not found. Cure: Restart Digitizer. If error persists please call Service.
12335 | LOCKER_UNLOCK_IP_MOVE_STOP_UNEXPECTED | Error |  |  |  | Locker UnLock Ip Pos Movement terminated with failure. | Reason: Stop before breakpoint step limit reached, or mechanical way blocked. Meaning: Function disturbed. UnLock Ip pos not found. Cure: Restart Digitizer. If error persists please call Service.
12336 | LOCKER_ABS_POS_FAILED | Error |  |  |  | Locker Absolute Pos Movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop not found. Meaning: Function disturbed. Homepos/Absolutepos not found. Cure: Restart Digitizer. If error persists please call Service.
12337 | LOCKER_CAL_LOCK_IP_MOVE_STOP_UNEXPECTED | Error |  |  |  | Locker calibration Lock Ip Pos movement stopped too early. | Reason: The Locker calibration Lock Ip Pos movement stopped early as stall reacted in position range before target position reached. Meaning: Retry possible, function may be disturbed. Cure: Restart Digitizer. If error persists please call Service.
12338 | LOCKER_CAL_LOCK_IP_MOVE_FAILED | Error |  |  |  | Locker calibration Lock Ip Pos movement terminated with failure. | Reason: Maximum step limit reached, or mechanical stop at wrong position. Meaning: Retry possible, function may be disturbed. Cure: Restart Digitizer. If error persists please call Service.
12339 | LOCKER_CAL_UNLOCK_IP_MOVE_FAILED | Error |  |  |  | Locker calibration UnLock Ip Pos movement terminated with failure. | Reason: Maximum step limit reached, or mechanical problem Meaning: Retry possible, function may be disturbed. Cure: Restart Digitizer. If error persists please call Service.
12340 | FIXER_CLAMP_CASS_CAL_MAX_STEP_LIMIT | Error |  |  |  | Clamp Cassette Movement stopped due to maximum step limit | Reason: Maximum step limit reached. Meaning: Function disturbed. Cure: Re-Start Calibration Process Cassette. Restart Digitizer. If error persists please call Service.
12341 | PREPARE_SM_SELFTEST_FAILED | Error |  |  |  | Preparation of the Selftest of the stepper motors failed. | Reason: SW-Function to set all motor currents failed Meaning: Function disturbed. Cure: Load new SW. Restart Digitizer. If error persists please call Service.
12342 | LOCKER_LOCK_IP_CONTACT_NO_CASSETTE_DETECTED | Error |  |  |  | Locker lockIp movement terminated without contact to Cassette Slider . | Reason: The Locker lockIp movement terminated, but the Locker Contact Sensor has no contact at expected position range. Meaning: Retry possible, function may be disturbed. Cassette Slider not reacting or blocked, no Cassette present. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12343 | LOCKER_CAPSULE_LOCK_IP_CONTACT_NO_CASSETTE_DETECTED | Error |  |  |  | Locker Capsule lockIp movement terminated without contact to Cassette Slider. | Reason: The Locker Capsule lockIp movement terminated, but the Locker Contact Sensor has no contact at expected position range. Meaning: Retry possible, function may be disturbed. Cassette Slider not reacting or blocked, no Cassette present. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12344 | LOCKER_CAPSULE_LOCK_IP_CONTACT_FAILED | Error |  |  |  | Locker Capsule lockIp movement terminated with failure | Reason: The Locker Capsule lockIp movement terminated, but the result indicates failure. Meaning: Retry possible, function may be disturbed. Cassette Slider not reacting or blocked, no Cassette present. Mechanics defect or blocked. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12345 | LOCKER_LOCK_IP_CONTACT_STALL_STOP_UNEXPECTED | Error |  |  |  | Locker lockIp movement terminated with unexpected Stall Stop | Reason: The Locker lockIp movement terminated, but the Locker has reached stall stop on the way. Meaning: Retry possible, function may be disturbed. Cassette Slider blocked, or mechanic blocking. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12346 | LOCKER_UNLOCK_IP_CONTACT_NO_CASSETTE_DETECTED | Error |  |  |  | Locker unlockIp movement terminated without contact to Cassette Slider . | Reason: The Locker unlockIp movement terminated, but the Locker Contact Sensor has no more contact at expected position range. Meaning: Retry possible, function may be disturbed. Cassette Slider not reacting or blocked. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12347 | LOCKER_UNLOCK_IP_CONTACT_POS_LIMIT_REACHED_UNEXPECTED | Error |  |  |  | Locker unlock movement stopped not with expected stall reaction. | Reason: The Locker unlock movement stopped not as stall reacted at stopper. Meaning: Retry possible, function may be disturbed. Movement terminated unexpectedly by position limit. Cure: Restart Digitizer. If error persists please call Service.
12348 | LOCKER_CAPSULE_UNLOCK_IP_CONTACT_FAILED | Error |  |  |  | Locker Capsule unlockIp movement terminated with failure | Reason: The Locker Capsule unlockIp movement terminated, but the result indicates failure. Meaning: Retry possible, function may be disturbed. Cassette Slider not reacting or blocked, no Cassette present. Mechanics defect or blocked. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12349 | LOCKER_CAPSULE_UNLOCK_IP_CONTACT_NO_CASSETTE_DETECTED | Error |  |  |  | Locker Capsule unlockIp movement terminated without contact to Cassette Slider . | Reason: The Locker Capsule unlockIp movement terminated, but the Locker Contact Sensor has no more contact at expected position range. Meaning: Retry possible, function may be disturbed. Cassette Slider not reacting or blocked. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12350 | LOCKER_CAPSULE_UNLOCK_IP_CONTACT_FAILED_NO_CASSETTE_DETECTED | Error |  |  |  | Locker Capsule unlockIp movement terminated with failure as started without contact to Cassette Slider . | Reason: The Locker Capsule unlockIp movement terminated with failure, as the Locker Contact Sensor has no contact at expected position range. Start of the movement therefore blocked. Precondition violated. Meaning: Retry possible, function may be disturbed. Cassette Slider not reacting or blocked. Cure: Restart Digitizer. Change Cassette. If error persists please call Service.
12351 | CASSTRANSPORT_IRUN_SHIFT_DEFAULT_AT_ELABEL_FAILURE | Error |  |  |  | CassTransport iRun from Elabel set to default value at Elabel Failure. | Meaning: Failure in reading Elabel value CTiRunShift results in default value. Reason: The Handling Motor CassetteTransport reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12352 | CASSTRANSPORT_IRUN_SHIFT_ELABEL_ENTRY_FAILURE | Error |  |  |  | CassTransport iRun entry not found at Elabel Failure. | Meaning: Failure in reading Elabel value CTiRunShift results in not found entry. At startup: try to make new entry. Reason: The Handling Motor CassetteTransport reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12353 | CASSTRANSPORT_IRUN_SHIFT_ELABEL_READ_ENTRY_FAILURE | Error |  |  |  | CassTransport iRun read entry at Elabel Failure. | Meaning: Failure in reading Elabel value CTiRunShift. Reason: The Handling Motor CassetteTransport reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12354 | CASSTRANSPORT_IRUN_SHIFT_ELABEL_WRITE_ENTRY_FAILURE | Error |  |  |  | CassTransport iRun write entry at Elabel Failure. | Meaning: Failure in writing Elabel value CTiRunShift. Reason: The Handling Motor CassetteTransport writes the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12355 | CASSTRANSPORT_IRUN_SHIFT_ELABEL_WRITE_RANGE_FAILURE | Error |  |  |  | CassTransport iRun write at Elabel range Failure. | Meaning: Failure at value range in writing Elabel entry CTiRunShift. Reason: The Handling Motor CassetteTransport Elabel entry CTiRunShift is written with a value causing an error condition at the range check. Cure: Restart device, check ELabel on Cassette Handling Board.
12356 | CASS_HDL_BOARD_INTERNAL_ELABEL_INIT_FAILED | Error |  |  |  | Cassette Handling Board init Lib at Elabel Failure. | Meaning: Failure at initialisation of the Elabel library. Reason: The Cassette Handling Board Elabel initialisation of the SW-Lib terminated with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12357 | CASS_HDL_BOARD_INTERNAL_ELABEL_CHECK_ENTRIES_FAILED | Error |  |  |  | Cassette Handling Board check entries at Elabel Failure. | Meaning: Failure at initialisation of the Elabel at checking the entries. Reason: The Cassette Handling Board Elabel initialisation of the Elabel entries terminated with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12358 | CASSTRANSPORT_SHIFT_CASS_OUT_FAILURE | Error |  |  |  | CassTransport Shift Cassette Out movement terminated with Failure. | Meaning: Failure at movement Shift Cassette Out to Buffer. Cassette Jam possible. Reason: The Handling Motor CassetteTransport has an error condition. Retry possible, function may be disturbed. Cassette Belt not reacting or blocked. Cure: Restart device, Remove Cassettes. If error persists please call Service.
12359 | CASSTRANSPORT_SHIFT_CASS_TO_SLUICE_FAILURE | Error |  |  |  | CassTransport Shift Cassette to Sluice movement terminated with Failure. | Meaning: Failure at movement Shift Cassette to Sluice. Cassette Jam possible. Reason: The Handling Motor CassetteTransport has an error condition. Retry possible, function may be disturbed. Cassette Belt not reacting or blocked. Cure: Restart device, Remove Cassettes. If error persists please call Service.
12360 | CASSTRANSPORT_SHIFT_CASS_TO_EDGE_FAILURE | Error |  |  |  | CassTransport Shift Cassette to Edge movement terminated with Failure. | Meaning: Failure at movement Shift Cassette to Edge. Cassette Jam possible. Reason: The Handling Motor CassetteTransport has an error condition. Retry possible, function may be disturbed. Cassette Belt not reacting or blocked. Cure: Restart device, Remove Cassettes. If error persists please call Service.
12361 | CASSTRANSPORT_PREPARE_TO_EJECT_FAILURE | Error |  |  |  | CassTransport Prepare to Eject movement terminated with Failure. | Meaning: Failure at movement Prepare to Eject. Cassette Jam possible. Reason: The Handling Motor CassetteTransport has an error condition. Retry possible, function may be disturbed. Cassette Belt not reacting or blocked. Cure: Restart device, Remove Cassettes. If error persists please call Service.
12362 | CASSTRANSPORT_FIND_REF_FAILURE | Error |  |  |  | CassTransport Find Ref movement terminated with Failure. | Meaning: Failure at movement Find Ref. Cassette Jam possible. Reason: The Handling Motor CassetteTransport has an error condition. Retry possible, function may be disturbed. Cassette Belt not reacting or blocked. Cure: Restart device, Remove Cassettes. If error persists please call Service.
12363 | CASSTRANSPORT_FIND_REF_FWD_FAILURE | Error |  |  |  | CassTransport Find Ref Fwd movement terminated with Failure. | Meaning: Failure at movement Find Ref Fwd. Cassette Jam possible. Reason: The Handling Motor CassetteTransport has an error condition. Retry possible, function may be disturbed. Cassette Belt not reacting or blocked. Cure: Restart device, Remove Cassettes. If error persists please call Service.
12364 | CASSTRANSPORT_REPUSH_CASS_FAILURE | Error |  |  |  | CassTransport Repush Cass movement terminated with Failure. | Meaning: Failure at movement Repush Cass. Cassette Jam possible. Reason: The Handling Motor CassetteTransport has an error condition. Retry possible, function may be disturbed. Cassette Belt not reacting or blocked. Cure: Restart device, Remove Cassettes. If error persists please call Service.
12365 | CASSTRANSPORT_IRUN_SHIFT_OUT_ELABEL_ENTRY_FAILURE | Error |  |  |  | CassTransport iRun ShiftOut entry not found at Elabel Failure. | Meaning: Failure in reading Elabel value CTiRunShiftOut results in not found entry. At startup: try to make new entry. Reason: The Handling Motor CassetteTransport reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12366 | CASSTRANSPORT_IRUN_SHIFT_OUT_DEFAULT_AT_ELABEL_FAILURE | Error |  |  |  | CassTransport iRun ShiftOut from Elabel set to default value at Elabel Failure. | Meaning: Failure in reading Elabel value CTiRunShiftOut results in default value. Reason: The Handling Motor CassetteTransport reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12367 | CASSTRANSPORT_IRUN_SHIFT_OUT_ELABEL_READ_ENTRY_FAILURE | Error |  |  |  | CassTransport iRun ShiftOut read entry at Elabel Failure. | Meaning: Failure in reading Elabel value CTiRunShiftOut. Reason: The Handling Motor CassetteTransport reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12368 | CASSTRANSPORT_IRUN_SHIFT_OUT_ELABEL_WRITE_ENTRY_FAILURE | Error |  |  |  | CassTransport iRun ShiftOut write entry at Elabel Failure. | Meaning: Failure in writing Elabel value CTiRunShiftOut. Reason: The Handling Motor CassetteTransport writes the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12369 | CASSTRANSPORT_IRUN_SHIFT_OUT_ELABEL_WRITE_RANGE_FAILURE | Error |  |  |  | CassTransport iRun ShiftOut write at Elabel range Failure. | Meaning: Failure at value range in writing Elabel entry CTiRunShiftOut. Reason: The Handling Motor CassetteTransport Elabel entry CTiRunShiftOut is written with a value causing an error condition at the range check. Cure: Restart device, check ELabel on Cassette Handling Board.
12370 | CHC_CASS_SLOT_WIDTH_ELABEL_ENTRY_FAILURE | Error |  |  |  | CHC CassSlotWidth entry not found at Elabel Failure. | Meaning: Failure in reading Elabel value CassSlotWidth results in not found entry. At startup: try to make new entry. Reason: The Cassette Handling reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12371 | CHC_CASS_SLOT_WIDTH_DEFAULT_AT_ELABEL_FAILURE | Error |  |  |  | CHC CassSlotWidth from Elabel set to default value at Elabel Failure. | Meaning: Failure in reading Elabel value CassSlotWidth results in default value. Reason: The Cassette Handling reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12372 | CHC_CASS_SLOT_WIDTH_ELABEL_READ_ENTRY_FAILURE | Error |  |  |  | CHC CassSlotWidth read entry at Elabel Failure. | Meaning: Failure in reading Elabel value CassSlotWidth. Reason: The Cassette Handling reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12373 | CHC_CASS_SLOT_WIDTH_ELABEL_WRITE_ENTRY_FAILURE | Error |  |  |  | CHC CassSlotWidth write entry at Elabel Failure. | Meaning: Failure in writing Elabel value CassSlotWidth. Reason: The Cassette Handling writes the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12374 | CHC_CASS_SLOT_WIDTH_ELABEL_WRITE_RANGE_FAILURE | Error |  |  |  | CHC CassSlotWidth write at Elabel range Failure. | Meaning: Failure at value range in writing Elabel entry CassSlotWidth. Reason: The Cassette Handling Elabel entry CassSlotWidth is written with a value causing an error condition at the range check. Cure: Restart device, check ELabel on Cassette Handling Board.
12375 | FIXER_IRUN_CLAMP_ELABEL_ENTRY_FAILURE | Error |  |  |  | Fixer iRunClamp entry not found at Elabel Failure. | Meaning: Failure in reading Elabel value iRunClamp results in not found entry. At startup: try to make new entry. Reason: The Handling Motor Fixer reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12376 | FIXER_IRUN_CLAMP_DEFAULT_AT_ELABEL_FAILURE | Error |  |  |  | Fixer iRunClamp from Elabel set to default value at Elabel Failure. | Meaning: Failure in reading Elabel value iRunClamp results in default value. Reason: The Handling Motor Fixer reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12377 | FIXER_IRUN_CLAMP_ELABEL_READ_ENTRY_FAILURE | Error |  |  |  | Fixer iRunClamp read entry at Elabel Failure. | Meaning: Failure in reading Elabel value iRunClamp. Reason: The Handling Motor Fixer reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12378 | FIXER_IRUN_CLAMP_ELABEL_WRITE_ENTRY_FAILURE | Error |  |  |  | Fixer iRunClamp write entry at Elabel Failure. | Meaning: Failure in writing Elabel value iRunClamp. Reason: The Handling Motor Fixer writes the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12379 | FIXER_IRUN_CLAMP_ELABEL_WRITE_RANGE_FAILURE | Error |  |  |  | Fixer iRunClamp write at Elabel range Failure. | Meaning: Failure at value range in writing Elabel entry iRunClamp. Reason: The Handling Motor Fixer Elabel entry iRunClamp is written with a value causing an error condition at the range check. Cure: Restart device, check ELabel on Cassette Handling Board.
12380 | LOCKER_IRUN_UNLOCK_ELABEL_ENTRY_FAILURE | Error |  |  |  | Locker iRunUnlock entry not found at Elabel Failure. | Meaning: Failure in reading Elabel value iRunUnlock results in not found entry. At startup: try to make new entry. Reason: The Handling Motor Locker reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12381 | LOCKER_IRUN_UNLOCK_DEFAULT_AT_ELABEL_FAILURE | Error |  |  |  | Locker iRunUnlock from Elabel set to default value at Elabel Failure. | Meaning: Failure in reading Elabel value iRunUnlock results in default value. Reason: The Handling Motor Locker reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12382 | LOCKER_IRUN_UNLOCK_ELABEL_READ_ENTRY_FAILURE | Error |  |  |  | Locker iRunUnlock read entry at Elabel Failure. | Meaning: Failure in reading Elabel value iRunUnlock. Reason: The Handling Motor Locker reads the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12383 | LOCKER_IRUN_UNLOCK_ELABEL_WRITE_ENTRY_FAILURE | Error |  |  |  | Locker iRunUnlock write entry at Elabel Failure. | Meaning: Failure in writing Elabel value iRunUnlock. Reason: The Handling Motor Locker writes the Elabel value with an error condition. Cure: Restart device, check ELabel on Cassette Handling Board.
12384 | LOCKER_IRUN_UNLOCK_ELABEL_WRITE_RANGE_FAILURE | Error |  |  |  | Locker iRunUnlock write at Elabel range Failure. | Meaning: Failure at value range in writing Elabel entry iRunUnlock. Reason: The Handling Motor Locker Elabel entry iRunUnlock is written with a value causing an error condition at the range check. Cure: Restart device, check ELabel on Cassette Handling Board.
12385 | LOCKER_FIND_REFERENCE_FAILURE | Error |  |  |  | Stepper Motor Locker terminated find reference action with failure condition. | Meaning: CassetteHandlingControl: Failure in initialization of SM LOCKING at find reference. Reason: Failure due to mechanical, sensoric or electronic influence. Cure: 1) Redo stall calibration 3 times. Check values for consistency (+/- 2 digits). 2) Check mechanics of SM LOCKING drive. 3) Check cabling to CASSETTE HANDLING BOARD. 4) Replace SM SM LOCKING if error persists. 5) Replace CASSETTE HANDLING BOARD if error persists.
12386 | FIXER_FIND_REFERENCE_FAILURE | Error |  |  |  | Stepper Motor Fixer terminated find reference action with failure condition. | Meaning: CassetteHandlingControl: Failure in initialization of motor CASSETTEFIXATION at find reference. Reason: Failure due to mechanical, or electronic influence. Cure: 1) Redo stall calibration 3 times. Check values for consistency (+/- 2 digits). 2) Check mechanics of CASSETTEFIXATION drive. 3) Check cabling to CASSETTE HANDLING BOARD. 4) Replace SM CASSETTEFIXATION if error persists. 5) Replace CASSETTE HANDLING BOARD if error persists.
12800 | CARRIER_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of CARRIER terminated with failure. | Meaning: The SM CARRIER DRIVE terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM CARRIER DRIVE. 2) Check IP handling board. 3) Replace IP handling board if error persists.
12801 | CARRIER_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of CARRIER terminated with UNPLUGGED failure. | Meaning: The motor SM CARRIER DRIVE terminated its selftest with an error condition Reason: Cable or connector unplugged. Cure: 1) check connectors of SM CARRIER DRIVE. 2) check cabling of SM CARRIER DRIVE.
12802 | CARRIER_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of CARRIER terminated with SIN COIL failure. | Meaning: The SM CARRIER DRIVE terminated its selftest with an error condition Reason: Cabling of motor or motor itself defective. Cure: 1) Repeat the stall calibration of the SM_CARRIER DRIVE 3 times. Check values for consistency (+/- 2 digits). 2) Check lightbarriers CARRIER_DRIVE HOME POSITION COARSE SENSOR and CARRIER_DRIVE HOME POSITION FINE SENSOR. 3) Check for mechanic problems blocking the movement of the carrier drive. 4) Check IP handling board.
12803 | CARRIER_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of CARRIER terminated with COS COIL failure. | Meaning: The SM CARRIER DRIVE terminated its selftest with an error condition Reason: Cabling of motor or motor itself defective. Cure: 1) Repeat the stall calibration of the SM_CARRIER DRIVE 3 times. Check values for consistency (+/- 2 digits). 2) Check lightbarriers CARRIER_DRIVE HOME POSITION COARSE SENSOR and CARRIER_DRIVE HOME POSITION FINE SENSOR. 3) Check for mechanic problems blocking the movement of the carrier drive. 4) Check IP handling board.
12804 | CARRIER_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of CARRIER terminated with DRIVER failure. | Meaning: The SM CARRIER DRIVE terminated its selftest with an error condition Reason: Driver failure of SM CARRIER DRIVE. Cure: 1) Check IP handling board. 2) Replace IP handling board if error persists.
12805 | CARRIER_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of CARRIER terminated with overcurrent to Ground detected failure. | Meaning: The motor SM CARRIER DRIVE terminated its selftest with an error condition. Reason: Cabling of SM CARRIER DRIVE or IP handling board defective. Cure: 1) Check cabling of SM CARRIER DRIVE. 2) Check IP handling board. 3) Replace IP handling board, if error persists.
12806 | CONNECTOR_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of CONNECTOR terminated with failure. | Meaning: The SM CONNECTOR terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM CONNECTOR. 2) Check IP handling board. 3) Replace IP handling board if error persists.
12807 | CONNECTOR_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of CONNECTOR terminated with UNPLUGGED failure. | Meaning: The motor CONNECTOR terminated its selftest with an error condition Reason: No cabling connection to SM CONNECTOR. Cure: 1) Check cable connectors for proper seat. 2) Check for cable damages. 3) Replace SM Connector with cabling in case error persists.
12808 | CONNECTOR_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of CONNECTOR terminated with SIN COIL failure. | Meaning: The SM CONNECTOR terminated its selftest with an error condition Reason: Cabling or motor itself defective. Cure: 1) Check cabling of SM CONNECTOR. 2) Check IP handling board. 3) Replace SM CONNECTOR if error persists.
12809 | CONNECTOR_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of CONNECTOR terminated with COS COIL failure. | Meaning: The SM CONNECTOR terminated its selftest with an error condition Reason: Cabling or motor itself defective. Cure: 1) Check cabling of SM CONNECTOR. 2) Check IP handling board. 3) Replace SM CONNECTOR if error persists.
12810 | CONNECTOR_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of CONNECTOR terminated with DRIVER failure. | Meaning: The SM CONNECTOR terminated its selftest with an error condition Reason: Driver failure of SM CONNECTOR. Cure: 1) Check IP handling board. 2) Replace IP handling board if error persists.
12811 | CONNECTOR_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of CONNECTOR terminated with overcurrent to Ground detected failure. | Meaning: The motor SM CONNECTOR terminated its selftest with an error condition. Reason: Cabling of SM CONNECTOR or IP handling board defective. Cure: 1) Check cabling of SM CONNECTOR. 2) Check IP handling board. 3) Replace IP handling board, if error persists.
12812 | LINEARTRANSPORT_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of LINEARTRANSPORT terminated with failure. | Meaning: The SM LINEARTRANSPORT terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM LINEARTRANSPORT. 2) Check IP handling board. 3) Replace IP handling board if error persists.
12813 | LINEARTRANSPORT_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of LINEARTRANSPORT terminated with UNPLUGGED failure. | Meaning: The motor LINEARTRANSPORT terminated its selftest with an error condition Reason: No cabling connection to SM LINEARTRANSPORT. Cure: 1) Check cable connectors for proper seat. 2) Check for cable damages. 3) Replace SM LINEARTRANSPORT with cabling in case error persists.
12814 | LINEARTRANSPORT_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of LINEARTRANSPORT terminated with SIN COIL failure. | Meaning: The SM LINEARTRANSPORT terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM LINEARTRANSPORT. 2) Replace SM LINEARTRANSPORT if error persists.
12815 | LINEARTRANSPORT_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of LINEARTRANSPORT terminated with COS COIL failure. | Meaning: The SM LINEARTRANSPORT terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM LINEARTRANSPORT. 2) Replace SM LINEARTRANSPORT if error persists.
12816 | LINEARTRANSPORT_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of LINEARTRANSPORT terminated with DRIVER failure. | Meaning: The SM LINEARTRANSPORT terminated its selftest with an errror condition. Reason: Driver problem during selftest detected. Cure: 1) Check cabling of SM LINEARTRANSPORT. 2) Replace IP Handling board if error persists.
12817 | LINEARTRANSPORT_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of LINEARTRANSPORT terminated with overcurrent to Ground detected failure. | Meaning: The SM LINEARTRANSPORT terminated its selftest with an errror condition. Reason: Overcurrent to ground during selftest detected. Cure: 1) Check cabling of SM LINEARTRANSPORT. 2) Replace IP Handling board if error persists.
12818 | SLOWSCAN_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of SLOWSCAN terminated with failure. | Meaning: The SM SLOWSCAN terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SLOWSCAN 2) Check IP handling board. 3) Replace IP handling board if error persists.
12819 | SLOWSCAN_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of SLOWSCAN terminated with UNPLUGGED failure. | Meaning: The motor SLOWSCAN terminated its selftest with an error condition Reason: No cabling connection to SM SLOWSCAN. Cure: 1) Check cable connectors for proper seat. 2) Check for cable damages. 3) Replace SM SLOWSCAN with cabling in case error persists.
12820 | SLOWSCAN_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of SLOWSCAN terminated with SIN COIL failure. | Meaning: The SM SLOWSCAN terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SLOWSCAN. 2) Replace SM SLOWSCAN if error persists.
12821 | SLOWSCAN_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of SLOWSCAN terminated with COS COIL failure. | Meaning: The SM SLOWSCAN terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SLOWSCAN. 2) Replace SM SLOWSCAN if error persists.
12822 | SLOWSCAN_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of SLOWSCAN terminated with DRIVER failure. | Meaning: The SM SLOWSCAN terminated its selftest with an errror condition. Reason: Driver problem during selftest detected. Cure: 1) Check cabling of SM SLOWSCAN. 2) Replace IP Handling board if error persists.
12823 | SLOWSCAN_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of SLOWSCAN terminated with overcurrent to Ground detected failure. | Meaning: The SM SLOWSCAN terminated its selftest with an errror condition. Reason: Overcurrent to ground during selftest detected. Cure: 1) Check cabling of SM SLOWSCAN. 2) Replace IP Handling board if error persists.
12824 | SNAPHOOKS_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of SNAPHOOKS terminated with failure. | Meaning: The SM SNAPHOOKS terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SNAPHOOKS 2) Check IP handling board. 3) Replace IP handling board if error persists.
12825 | SNAPHOOKS_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of SNAPHOOKS terminated with UNPLUGGED failure. | Meaning: The motor SNAPHOOKS terminated its selftest with an error condition Reason: No cabling connection to SM SNAPHOOKS. Cure: 1) Check cable connectors for proper seat. 2) Check for cable damages. 3) Replace SM SNAPHOOKS with cabling in case error persists.
12826 | SNAPHOOKS_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of SNAPHOOKS terminated with SIN COIL failure. | Meaning: The SM SNAPHOOKS terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SNAPHOOKS. 2) Replace SM SNAPHOOKS if error persists.
12827 | SNAPHOOKS_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of SNAPHOOKS terminated with COS COIL failure. | Meaning: The SM SNAPHOOKS terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SNAPHOOKS. 2) Replace SM SNAPHOOKS if error persists.
12828 | SNAPHOOKS_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of SNAPHOOKS terminated with DRIVER failure. | Meaning: The SM SNAPHOOKS terminated its selftest with an errror condition. Reason: Driver problem during selftest detected. Cure: 1) Check cabling of SM SNAPHOOKS. 2) Replace IP Handling board if error persists.
12829 | SNAPHOOKS_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of SNAPHOOKS terminated with overcurrent to Ground detected failure. | Meaning: The SM SNAPHOOKS terminated its selftest with an errror condition. Reason: Overcurrent to ground during selftest detected. Cure: 1) Check cabling of SM SNAPHOOKS. 2) Replace IP Handling board if error persists.
12830 | SWIVEL_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of SWIVEL terminated with failure. | Meaning: The SM SWIVELDRIVE terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SWIVELDRIVE 2) Check IP handling board. 3) Replace IP handling board if error persists.
12831 | SWIVEL_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of SWIVEL terminated with UNPLUGGED failure. | Meaning: The motor SWIVELDRIVE terminated its selftest with an error condition Reason: No cabling connection to SM SWIVELDRIVE Cure: 1) Check cable connectors for proper seat. 2) Check for cable damages. 3) Replace SM SWIVELDRIVE with cabling in case error persists.
12832 | SWIVEL_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of SWIVEL terminated with SIN COIL failure. | Meaning: The SM SWIVELDRIVE terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SWIVELDRIVE. 2) Replace SM SWIVELDRIVE if error persists.
12833 | SWIVEL_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of SWIVEL terminated with COS COIL failure. | Meaning: The SM SWIVELDRIVE terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM SWIVELDRIVE. 2) Replace SM SWIVELDRIVE if error persists.
12834 | SWIVEL_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of SWIVEL terminated with DRIVER failure. | Meaning: The SM SWIVELDRIVE terminated its selftest with an errror condition. Reason: Driver problem during selftest detected. Cure: 1) Check cabling of SM SWIVELDRIVE. 2) Replace IP Handling board if error persists.
12835 | SWIVEL_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of SWIVEL terminated with overcurrent to Ground detected failure. | Meaning: The SM SWIVELDRIVE terminated its selftest with an errror condition. Reason: Overcurrent to ground during selftest detected. Cure: 1) Check cabling of SM SWIVELDRIVE. 2) Replace IP Handling board if error persists.
12836 | VALVE_SELFTEST_FAILED | Error |  |  |  | Procedure selftest of VALVE terminated with failure. | Meaning: The SM VACUUMVALVE terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM VACUUMVALVE. 2) Check IP handling board. 3) Replace IP handling board if error persists.
12837 | VALVE_UNPLUGGED_FAILURE | Error |  |  |  | Procedure selftest of VALVE terminated with UNPLUGGED failure. | Meaning: The motor VACUUMVALVE terminated its selftest with an error condition Reason: No cabling connection to SM VACUUMVALVE Cure: 1) Check cable connectors for proper seat. 2) Check for cable damages. 3) Replace SM VACUUMVALVE with cabling in case error persists.
12838 | VALVE_SIN_COIL_FAILURE | Error |  |  |  | Procedure selftest of VALVE terminated with SIN COIL failure. | Meaning: The SM VACUUMVALVE terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM VACUUMVALVE. 2) Replace SM VACUUMVALVE if error persists.
12839 | VALVE_COS_COIL_FAILURE | Error |  |  |  | Procedure selftest of VALVE terminated with COS COIL failure. | Meaning: The SM VACUUMVALVE terminated its selftest with an errror condition. Reason: Problem during selftest detected. Cure: 1) Check cabling of SM VACUUMVALVE. 2) Replace SM VACUUMVALVE if error persists.
12840 | VALVE_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of VALVE terminated with DRIVER failure. | Meaning: The SM VACUUMVALVE terminated its selftest with an errror condition. Reason: Driver problem during selftest detected. Cure: 1) Check cabling of SM VACUUMVALVE. 2) Replace IP Handling board if error persists.
12841 | VALVE_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of VALVE terminated with overcurrent to Ground detected failure. | Meaning: The SM VACUUMVALVE terminated its selftest with an errror condition. Reason: Overcurrent to ground during selftest detected. Cure: 1) Check cabling of SM VACUUMVALVE. 2) Replace IP Handling board if error persists.
12842 | WRONG_VALUE_STORED | Error |  |  |  | The value stored in the eLabel is wrong. | Meaning: The run current value stored in the eLabel is wrong. The required and the set values are logged. Reason: -- Cure: Check the corresponding SW routines and eventually the hardware.
12843 | WRONG_STATE | Error |  |  |  | The event arrived in the wrong state. | Meaning: The event arrived in the wrong state. Reason: -- Cure: Check the corresponding SW routines.
12844 | EXTERNAL_PORT_POINTER_ZERO | Error |  |  |  | The pointer of an external port was zero. | Meaning: The pointer of an external port for sending a low level response to a motor class was zero (therefore the response could not be sent). Reason: -- Cure: Check the corresponding SW routine
12845 | LOW_LEVEL_RESPONSE_TO_MOTOR_CLASS_NOT_SENT | Error |  |  |  | A low level response to a motor class was not sent. | Meaning: A low level response to a motor class was not sent. The function 'RTPort_raiseExternal' did not work correctly. Reason: -- Cure: Check the corresponding SW routine
12846 | RACK_FAN_SELFTEST_FAILURE | Error |  |  |  | Procedure selftest of the RackFan terminated with failure. | Reason: The Rack Fan terminated its selftest with an error condition. No rotational pulses were detected sufficiently. Meaning: Retry possible, function disturbed. Rack Fan disconnected, blocked or defect. Cure: Retry, Restart, check cables, check Rack Fan, check IpHandling electronic board, Call Service.
13924 | FIXER_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Fixer failed. | Meaning: The calibration of the Fixer failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
13925 | FIXER_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Fixer was aborted. | Meaning: The calibration procedure of the Fixer was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
13926 | FIXER_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Fixer was invalid. | Meaning: The determined optimal run current value of the Fixer was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
13927 | LOCKER_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Locker failed. | Meaning: The calibration of the Locker failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
13928 | LOCKER_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Locker was aborted. | Meaning: The calibration procedure of the Locker was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
13929 | LOCKER_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Locker was invalid. | Meaning: The determined optimal run current value of the Locker was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
13930 | CASS_TRANSPORT_CALIBRATION_SHIFT_CASS_FAILED | Error |  |  |  | The calibration of the Cassette Transport failed. | Meaning: The calibration of the Cassette Transport failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
13931 | CASS_TRANSPORT_CALIBRATION_SHIFT_CASS_ABORTED | Error |  |  |  | The calibration procedure of the Cassette Transport was aborted. | Meaning: The calibration procedure of the Cassette Transport was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
13932 | CASS_TRANSPORT_CALIBRATION_SHIFT_CASS_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Cassette Transport was invalid. | Meaning: The determined optimal run current value of the Cassette Transport was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
13933 | CASS_TRANSPORT_CALIBRATION_SHIFT_CASS_OUT_FAILED | Error |  |  |  | The calibration of the Cassette Transport failed. | Meaning: The calibration of the Cassette Transport failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
13934 | CASS_TRANSPORT_CALIBRATION_SHIFT_CASS_OUT_ABORTED | Error |  |  |  | The calibration procedure of the Cassette Transport was aborted. | Meaning: The calibration procedure of the Cassette Transport was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
13935 | CASS_TRANSPORT_CALIBRATION_SHIFT_CASS_OUT_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Cassette Transport was invalid. | Meaning: The determined optimal run current value of the Cassette Transport was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14436 | SNAP_HOOKS_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Snaphooks failed. | Meaning: The calibration of the Snaphooks failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14437 | SNAP_HOOKS_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Snaphooks was aborted. | Meaning: The calibration procedure of the Snaphooks was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14438 | SNAP_HOOKS_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Snaphooks was invalid. | Meaning: The determined optimal run current value of the Snaphooks was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14439 | SWIVEL_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Swivel failed. | Meaning: The calibration of the Swivel failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14440 | SWIVEL_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Swivel was aborted. | Meaning: The calibration procedure of the Swivel was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14441 | SWIVEL_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Swivel was invalid. | Meaning: The determined optimal run current value of the Swivel was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14442 | CONNECTOR_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Connector failed. | Meaning: The calibration of the Connector failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14443 | CONNECTOR_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Connector was aborted. | Meaning: The calibration procedure of the Connector was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14444 | CONNECTOR_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Connector was invalid. | Meaning: The determined optimal run current value of the Connector was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14445 | VALVE_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Valve failed. | Meaning: The calibration of the Valve failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14446 | VALVE_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Valve was aborted. | Meaning: The calibration procedure of the Valve was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14447 | VALVE_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Valve was invalid. | Meaning: The determined optimal run current value of the Valve was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14448 | LINEAR_TRANSPORT_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Linear Transport failed. | Meaning: The calibration of the Linear Transport failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14449 | LINEAR_TRANSPORT_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Linear Transport was aborted. | Meaning: The calibration procedure of the Linear Transport was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14450 | LINEAR_TRANSPORT_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Linear Transport was invalid. | Meaning: The determined optimal run current value of the Linear Transport was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14451 | WAGON_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Wagon failed. | Meaning: The calibration of the Wagon failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14452 | WAGON_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Wagon was aborted. | Meaning: The calibration procedure of the Wagon was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14453 | WAGON_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Wagon was invalid. | Meaning: The determined optimal run current value of the Wagon was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14454 | CARRIER_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Carrier failed. | Meaning: The calibration of the Carrier failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14455 | CARRIER_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Carrier was aborted. | Meaning: The calibration procedure of theCarrier was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14456 | CARRIER_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Carrier was invalid. | Meaning: The determined optimal run current value of the Carrier was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14457 | CONNECTOR_CALIBRATION2_FAILED | Error |  |  |  | The calibration of the Connector failed. | Meaning: The calibration of the Connector failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14458 | CONNECTOR_CALIBRATION2_ABORTED | Error |  |  |  | The calibration procedure of the Connector was aborted. | Meaning: The calibration procedure of the Connector was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14459 | CONNECTOR_CALIBRATION2_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Connector was invalid. | Meaning: The determined optimal run current value of the Connector was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14848 | READ_OPTICS_TYPE | Error |  |  |  | Can't read optics type from optics eLabel -> using default type | Meaning: Default value for optic type are used. Reason: Unreadable entry in 'optic type' of optic module elabel. Cure: Replace optic module. | COM_NotNeeded_ForServiceId
14849 | READ_POLYGON_TYPE | Error |  |  |  | Can't read polygon type from optics eLabel -> using default type | Meaning: Default values for polygon type are used. Reason: Unreadable entry in polygon type of optic module elabel. Cure: Replace optic module. | COM_NotNeeded_ForServiceId
14850 | READ_OPERATION_MODE | Error |  |  |  | Can't read operation mode from PMI eLabel -> setting to normal mode | Meaning: 'Normal mode' is used as operation mode. Reason: Unreadable entry in 'operation mode' of PMI elabel. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
14851 | READ_POLYGON_SPEED_OFFSET | Error |  |  |  | Can't read polygon speed offset from optics eLabel -> using offset 0 |  | COM_NotNeeded_ForServiceId
14852 | INIT_PMM | Error |  |  |  | Initializing I2C hardware for PMM failed | Meaning: Communication error PMI to PMM. Reason: PMI and PMM connection not OK or PMI or PMM defective. Cure: 1) Check PMM cabling. 2) Replace PMM cabling. 3) Replace PMI board. 4) Replace PMM. | COM_PMMConnectionErr
14853 | CAL_BOARD_CONNECTOR | Error |  |  |  | Calibration board not detected | Meaning: Calibration board not detected. Reason: Cable broken or calibration board defective. Cure: 1) Check calibration board cabling. 2) Replace calibration board, if error persists. | COM_CalBoardConnectionErr
14854 | OPTICS_BOARD_CONNECTOR | Error |  |  |  | Optics board not detected | Meaning: Optic board not detected. Reason: Cable broken/not connected or optic board defective. Cure: 1) Check optic module cabling. 2) Replace optic module, if error persists. | COM_OpticsConnectionErr
14855 | LASER_DIODE | Error |  |  |  | Check laser diode failed | Meaning: Check laser diode failed. Reason: Laser diode defective. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement | COM_NotNeeded_ForServiceId
14856 | OPTICS_TYPE | Error |  |  |  | Unknown optics type detected | Meaning: Unknown optics type detected Reason: Unknown optics type in elabel. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement | COM_NotNeeded_ForServiceId
14857 | INIT_ELABEL | Error |  |  |  | Error initializing optics eLabel | Meaning: Initialization of OPTIC MODULE e-label (module specific parameter storage) failed. The Digitizer is out of operation and requires service support. Reason: OPTIC MODULE e-label can't be initialized due to: - OPTIC MODULE not connected - Defective of the PMI-BOARD - OPTIC MODULE cable or connectors - Corrupted data structure on the e-label - Defective OPTIC MODULE Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between PMI-BOARD and OPTIC MODULE 3. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement 4. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement | COM_NotNeeded_ForServiceId
14858 | PMM_ELABEL | Error |  |  |  | Error getting data from PMM eLabel | Meaning: Can't read data from PMM eLabel. Reason: PMM eLabel defective or PMM not connected. Cure: 1) Check PMM cabling. 2) Replace PMM cabling. 3) Replace PMI board. 4) Replace PMM. | COM_NotNeeded_ForServiceId
14859 | PMM_BOARD_CONNECTOR | Error |  |  |  | PMM board not detected | Meaning: PMM board not detected Reason: PMM is not connected to PMI board. Cure: 1) Check PMM cabling. 2) Replace PMM cabling, if error persists. 3) Replace PMM, if error persists. | COM_PMMConnectionErr
14860 | START_POLYGON | Error |  |  |  | Start or Lock polygon failed | Meaning: Polygon does not reach its speed. Reason: Power supply or polygon defective. Cure: Check power supply, replace optics module. | COM_NotNeeded_ForServiceId
14861 | UPDATE_PMM | Error |  |  |  | Error: PMM firmware update failed - Switch digitizer Off/On for retry | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists. | COM_NotNeeded_ForServiceId
14862 | UPDATE_PMM_FAILED | Error |  |  |  | Error: PMM firmware update failed | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists. | COM_NotNeeded_ForServiceId
14863 | PMMADJ_LPV_TOO_HIGH | Error |  |  |  | Prepare PMM adjust: LPV is too high |  | COM_NotNeeded_ForServiceId
14864 | PMMADJ_LED_BRIGHTNESS | Error |  |  |  | Prepare PMM adjust: Error LED not bright enough |  | COM_NotNeeded_ForServiceId
14865 | PMMADJ_SEL_SENS | Error |  |  |  | Prepare PMM adjust: Error selecting sensitivity |  | COM_NotNeeded_ForServiceId
14866 | PMMADJ_INCIDENCE_OF_LIGHT | Error |  |  |  | Prepare PMM adjust: Error because incidence of light |  | COM_NotNeeded_ForServiceId
14867 | PMMADJ_WRITE_ROV | Error |  |  |  | Analyse PMM adjust: Error writing new REFCAL to PMM |  | COM_NotNeeded_ForServiceId
14868 | PMMADJ_PARAMETER | Error |  |  |  | Analyse PMM adjust: Error parameter real dose or target index out of range |  | COM_NotNeeded_ForServiceId
14869 | BOL_TIMEOUT | Error |  |  |  | Lost BOL signal | Meaning: ESD or optics problem. Reason: electric discharge or optics module defective. Cure: if permanent check optics module. | COM_NotNeeded_ForServiceId
14870 | PMI_ELABEL | Error |  |  |  | Error getting data from PMI eLabel | Meaning: Error getting data from PMI eLabel. Reason: PMI eLabel defective. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
14871 | DIST_BOL_EOL_OUT_OF_RANGE | Error |  |  |  | Distance BOL to EOL out of range | Meaning: Distance BOL to EOL out of range. Reason: Polygon speed out of range, Optics module out of range. Cure: Check polygon speed (image width), replace optics module. | COM_NotNeeded_ForServiceId
14872 | PMI_HW_VER | Error |  |  |  | Software needs PMI-Board hardware version 1 or later | Meaning: PMI software does not fit to PMI hardware. Reason: Software update to an incompatible PMI board. Cure: Replace PMI board to index 1 or higher. | COM_NotNeeded_ForServiceId
14873 | PMI_FPGA_VER | Error |  |  |  | FPGA version does not fit to software version | Meaning: PMI software does not fit to PMI FPGA hardware. Reason: Software update failed or not completed. Cure: Update PMI software. | COM_NotNeeded_ForServiceId
14874 | OPTICS_ELABEL_CORRUPTED | Error |  |  |  | Optics eLabel data corrupted | Meaning: Optics eLabel can't be read, need to be restored from PC. Reason: Cable from PMI to optics module too long or unknown. Cure: Restore optics eLabel from PC. | COM_NotNeeded_ForServiceId
14875 | LD_AGING | Error |  |  |  | Laser diode is aging | Meaning: Laser diode is supplied with additional electrical power (1.5 of the standard value) to compensate aging for a while. Reason: Laser diode is short before end of life. | COM_NotNeeded_ForServiceId
14876 | INIDOSEDET_IPTYPEORDOSE | Error |  |  |  | IP type or dose out of range | Meaning: IP type or dose out of range. Reason: Software error. Cure: Install correct software. | COM_SWError
14877 | INIDOSEDET_WRONGHARDWARE | Error |  |  |  | initOverExpComparatorsFromData called on PMI1/2 | Meaning: The function initOverExpComparatorsFromData was called on the wrong hardware PMI1 or 2. Reason: Software error. Cure: Install correct software. | COM_NotNeeded_ForServiceId
14878 | PMM_BLUE_LED | Error |  |  |  | Photo Multiplier Module (PMM/PMT), blue LED defective or not connected | Meaning: Can't detect PMT blue LED. Reason: - PMT blue LED defective or not connected - PMT defective Cure: 1. Check connection and cable from blue LED to PMT 2. If error still persists replace PMT + LIGHT COLLECTOR | COM_NotNeeded_ForServiceId
14879 | PMM_TUBE_BOARD | Error |  |  |  | Photo Multiplier Module (PMM/PMT)defective or tube board not connected | Meaning: Can't read set HV. Reason: - PMM defective or tube board not connected Cure: 1. Check connection and cable from PMM main- to PMM tube board 2. If error still persists replace PMT + LIGHT COLLECTOR | COM_PMMConnectionErr
14880 | OPTICS_ELABEL | Error |  |  |  | Error getting data from optics eLabel | Meaning: Can't read data from optics eLabel. Reason: Optic module eLabel defective or optic module not connected. Cure: 1) Check optic module cabling. 2) Replace optic module cabling. 3) Replace PMI board. 4) Replace optic module. | COM_NotNeeded_ForServiceId
14881 | COULD_NOT_READ_PMM_FLASH_CNT | Warning |  |  |  | Error reading counters from PMM flash | Meaning: Cannot read counters from PMM flash memory. Cycle counter and operating hours may not correspond to real values of this PMM module. | COM_NotNeeded_ForServiceId
14882 | COULD_NOT_WRITE_PMM_FLASH_CNT | Warning |  |  |  | Error writing counters into PMM flash | Meaning: Cannot write counters into PMM flash memory. Cycle counter and operating hours may not correspond to real values of this PMM module. | COM_NotNeeded_ForServiceId
14883 | COULD_NOT_WRITE_OPTICS_ELABEL_CNT | Warning |  |  |  | Error writing counters into optics eLabel | Meaning: Cannot write counters into optics eLabel. Cycle counter and operating hours may not correspond to real values of this optics module. | COM_NotNeeded_ForServiceId
14884 | COULD_NOT_READ_OPTICS_ELABEL_CNT | Warning |  |  |  | Error reading counters from optics eLabel | Meaning: Cannot read counters from optics eLabel. Cycle counter and operating hours may not correspond to real values of this optics module. | COM_NotNeeded_ForServiceId
14885 | COULD_NOT_READ_OPTICS_ELABEL_INFO | Error |  |  |  | Error reading serial number from optics eLabel | Meaning: Cannot read serial number from optics eLabel . Reason: The optics eLabel is empty, was formatted, contains invalid data or is disconnected. Cure: 1. Check cabling connection between PMI board and optics module. 2. Format the optics eLabel via ESF, then write valid data to the optics eLabel via ESF. | COM_OpticModuleDataAccessErr
14886 | COULD_NOT_READ_PMM_INFO | Error |  |  |  | Error reading serial number from PMM flash memory | Meaning: Cannot read serial number from PMM flash memory . Reason: The flash memory is empty, was formatted, contains invalid data or is disconnected. Cure: 1. Check cabling connection between PMI board and PMM module. 2. Write valid data to the PMM flash memory via ESF. | COM_PMMDataAccessErr
14887 | COULD_NOT_READ_SENS_DATA | Error |  |  |  | Could not read sensitivity data | Meaning: Cannot read PMM sensisitivity data. Reason: PMM data not initialized (at production only) or PMM defective Cure: Adjust PMM (at production only) 1. Restart the Digitizer and check, if the problem persists 2. Check connection and cable from PMT to PMI 3. If error still persists replace PMT + LIGHT COLLECTOR | COM_PMMDataAccessErr
15972 | MAXIMUM_NUMBER_OF_PLATEAUS_EXCEEDED | Error |  |  |  | The maximum number of plateaus was exceeded. | Meaning: The maximum number of plateaus was exceeded. Reason: The system (mechanics/stepper motor) is not stable enough for a calibration. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
15973 | NO_PLATEAU_FOUND_FOR_EVALUATION | Error |  |  |  | No plateau was found for determining the optimal run current level. | Meaning: No plateau was found for determining the optimal run current level. Reason: The system (mechanics/stepper motor) is not stable enough for a calibration. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
15974 | WRONG_RUN_CURRENT_VALUE_SET | Error |  |  |  | The run current value set for the calibration cycle is wrong. | Meaning: The run current value set for the calibration cycle is wrong. The required and the set values are logged. Reason: -- Cure: Check the corresponding SW routine
15975 | FUNCTION_POINTER_ZERO | Error |  |  |  | A function pointer was zero. | Meaning: A function pointer was zero. Reason: -- Cure: Check the corresponding SW routine
15976 | MOTOR_SIM_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Motor Simulation failed. | Meaning: The calibration of the Motor Simulation failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding Software or Settings.
15977 | MOTOR_SIM_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Motor Simulation was aborted. | Meaning: The calibration procedure of the Motor Simulation was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
15978 | MOTOR_SIM_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Motor Simulation was invalid. | Meaning: The determined optimal run current value of the Motor Simulation was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding Software or Settings. Cure: Check the corresponding Software and Settings and retry the calibration procedure.
15979 | CB_FUNCTION_POINTER_ZERO | Error |  |  |  | A cb function pointer was zero. | Meaning: A cb function pointer was zero. Reason: -- Cure: Check the corresponding SW routine
15980 | WRONG_RUN_CURRENT_VALUE_STORED | Error |  |  |  | The run current value stored in the eLabel is wrong. | Meaning: The run current value stored in the eLabel is wrong. The required and the set values are logged. Reason: -- Cure: Check the corresponding SW routines and eventually the hardware.
15981 | WRONG_SM_ID | Error |  |  |  | The ID for selecting a motor is wrong (i.e. not in the required range of values). | Meaning: The ID for selecting a motor is wrong (i.e. not in the required range of values). Reason: -- Cure: Check the corresponding SW routines.
16896 | CAN_VCC_FAILURE | Error |  |  |  | CAN Vcc failure | Meaning: CAN voltage error. Reason: PMI power distribution board defective, cabling to PMI board defective or PMI board defective. Cure: 1) Check PMI power distribution board cabling. 2) Replace PMI power distribution board. 3) Replace PMI board.
16897 | PMM_PLUS20V_FAILURE | Error |  |  |  | PMM +20V failure | Meaning: PMM voltage error; Reason: PMI power supply fuse F18 or F19 broken, cable power supply to PMI broken or PMI power supply defect; Cure: Replace F18 or F19 [1A T], replace cable power supply to PMI, replace PMI power supply
16898 | PMM_MINUS10V_FAILURE | Error |  |  |  | PMM -10V failure | Meaning: PMM voltage error; Reason: PMI power supply fuse F17 broken, cable power supply to PMI broken or PMI power supply defect; Cure: Replace F17 [1A T], replace cable power supply to PMI, replace PMI power supply
16899 | ANALOG_VCC_FAILURE | Error |  |  |  | Analog Vcc failure | Meaning: Analog (ADC) voltage error; Reason: 24V missing to PMM. Cure: 1) Check PMM cabling. 2) Check PMI power distribution board cabling. 3) Replace PMI power distribution board. 4) Replace PMI board.
16900 | OPTICS_PLUS12V_FAILURE | Error |  |  |  | optics +12V failure | Meaning: Optics voltage error; Reason: PMI power supply fuse F1#1 broken, cable power supply to PMI broken or PMI power supply defect; Cure: Replace F1#1 [2.5A T], replace cable power supply to PMI, replace PMI power supply
16901 | OPTICS_MINUS12V_FAILURE | Error |  |  |  | optics -12V failure | Meaning: Optics voltage error; Reason: PMI power supply fuse F1 broken, cable power supply to PMI broken or PMI power supply defect; Cure: Replace F1 [2.5A T], replace cable power supply to PMI, replace PMI power supply
16902 | PMM_5V | Error |  |  |  | Check PMM 5 failed | Meaning: Failure of 5V for PMM. Reason: PMI power distribution board defective, cabling to PMI board defective or PMI board defective. Cure: 1) Check PMM cabling. 2) Check PMI power distribution board cabling. 3) Replace PMI power distribution board. 4) Replace PMI board.
16903 | INIT_BOARD_FAILED | Error |  |  |  | initBoard() failed | Meaning: Initialization of PMI board failed. Reason: Check underlying error. Cure: Check underlying error.
16904 | SET_IPDATA_WRONG_STATE | Error |  |  |  | setIPData received in wrong state | Meaning: the command setIPData has been received in a state in which it cannot be handled. Reason: Either a scanner preparation is already pending or the device is scanning or the image transmission to the client (PC) has not been completed. Cure: No cure, this is a violation of the protocol and therefore regarded as a software bug.
16905 | CASS_TO_EDGE_BLOCKED | Error | 287 | Unknown cassette format detected (cassette inserted correctly?). Remove cassette and press the confirm button. | Confirm | Cassette transport to edge sensor blocked | Meaning: Cassette slot edge sensor does not detect cassette while clamping with belt. Reason: 1) Cassette wrongly inserted (top down). 2) Cassette not completely down in slot. 3) Cassette blocked by any object in slot. 4) Slot edge sensor defective. Cure: 1) Check if cassette is inserted correctly. 2) Push cassette down. 3) remove object. 4) Adjust edge sensor, replace if needed.
16906 | UNKNOWN_CASS_FORMAT | Error | 287 | Unknown cassette format detected (cassette inserted correctly?). Remove cassette and press the confirm button. | Confirm | Cassette format determined by fixer drive is not valid. | Meaning: With Fixer drive detected cassette width does not match with any valid cassette format. Reason: 1) Cassette not correctly in slot and clamped with fixer. 2) Any unknown object is in slot. Cure: 1)Insert cassette correctly. 2) Remove object. 3) Check cassette initialization.
16907 | TAG_NOT_FOUND | Error | 270 | Error during read-out of cassette chip data. Remove and re-identify the cassette. | Confirm | IDData tag not found by tagreader | Meaning: IDData tag was not found by tagreader. Reason:1) No IP in cassette. 2) Wrong cassette in slot. 3) RF Tag defective Cure: 1) Put IP in cassette. 2) Exchange cassette. 3) Exchange RF tag on IP.
16908 | TAG_CORRUPTED | Error | 270 | Error during read-out of cassette chip data. Remove and re-identify the cassette. | Confirm | IDData tag found, but not readable with tagreader | Meaning:IDData tag not readable with tagreader. Reason: RF tag defective. Cure: Exchange RF tag on IP.
16909 | IDDATA_CORRUPTED | Error | 271 | Error during read-out of the ID data. Remove and re-identify the cassette. | Confirm | IDData read from tag are not valid. | Meaning:IDData read from tag are not valid. Reason: Wrong Identification. Cure: Reidentify tag. Replace RF tag if problem persists.
16910 | NOT_IDENTIFIED_DECPRECATED | Error |  |  |  | DEPRECATED | DEPRECATED
16911 | WRONG_CASS_DEPRECATED | Error |  |  |  | DEPRECATED | DEPRECATED
16912 | REFUSED_BY_PS_DEPRECATED | Error |  |  |  | DEPRECATED | DEPRECATED
16913 | OPEN_COVER_FAILED | Error | 273 | Cassette cannot be opened. The cassette may be damaged. Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. | Confirm | Open cassette cover failed. | Meaning: Open cassette cover failed. Reason: 1)Any damage at cassette. 2) Any defect at opener mechanics. Cure: 1) Check / Exchange cassette. 2) Check more specific error codes in infocounters; check mechanics with diagnose cycle of the service client,
16914 | GRIP_IP_FAILED | Error | 265 | Image plate is not scanned and erased! Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. | Confirm | Grip IP in cassette failed. | Meaning: Grip IP in cassette with snaphooks failed. Reason: 1) Any damage at cassette/IP. 2) Any defect at locking, snaphooks or IP-Transport mechanics. Cure: 1) Replace cassette/IP. 2) Try again; if fails permanently; check more specific error codes in infocounters; check mechanics with diagnose cycle of the service client
16915 | FEED_IP_FAILED | Error | 265 | Image plate is not scanned and erased! Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. | Confirm | Move IP down to stage failed. | Meaning: Moving IP down to stage failed. Reason: 1) Any damage at IP. 2) Any problem with IP transport down to stage. Cure: 1) Replace IP. 2) Try again; if fails permanently, check more specific error codes in infocounters; check mechanics with diagnose cycle of the service client
16916 | WRONG_HC_PROTVER | Error |  |  |  | Protocol version of HC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of HC (HandlingControl) is wrong. Cure: Install new software.
16917 | WRONG_BC_PROTVER | Error |  |  |  | Protocol version of BC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of BC (BufferControl) is wrong. Cure: Install new software.
16918 | WRONG_EC_PROTVER | Error |  |  |  | Protocol version of EC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of EC (ErasureControl) is wrong. Cure: Install new software.
16919 | WRONG_TRC_PROTVER | Error |  |  |  | Protocol version of TRC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of TRC (TagReadereControl) is wrong. Cure: Install new software.
16920 | WRONG_ICC_PROTVER | Error |  |  |  | Protocol version of ICC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ICC (ImageCaptureControl) is wrong. Cure: Install new software.
16921 | WRONG_UID_PROTVER | Error |  |  |  | Protocol version of UID is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of UID (UserInterfaceControl) is wrong. Cure: Install new software.
16922 | UNCHECKED_PROTVER | Error |  |  |  | Protocol version not checked | Meaning: Protocol version of a SW communication partner was not checked. Reason: SW error, protocol version check is missing in SW. Cure: Install new software.
16923 | CAN_ADAPTER_ERROR | Error |  |  |  | fatal error in CAN adapter layer | Meaning: Fatal error in CAN adapter layer Reason: CAN adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help, replace PMI board.
16924 | DELAY_TIMER_ALREADY_IN_USE_WAIT_EU_ON_TEST | Error |  |  |  | fatal SW error | Meaning: The DelayTimer is already running when the timer should be used for waiting a certain time to check if all LEDs of the erasure unit are ON. Reason: SW error. Cure: Install new software.
16925 | DELAY_TIMER_ALREADY_IN_USE_WAIT_EU_OFF_TEST_AGAIN | Error |  |  |  | fatal SW error | Meaning: The DelayTimer is already running when the timer should be used for waiting a certain time to check if all LEDs of the erasure unit are OFF in case of several erasure cycles. Reason: SW error. Cure: Install new software.
16926 | DELAY_TIMER_ALREADY_IN_USE_WAIT_EU_OFF_TEST | Error |  |  |  | fatal SW error | Meaning: The DelayTimer is already running when the timer should be used for waiting a certain time to check if all LEDs of the erasure unit are OFF. Reason: SW error. Cure: Install new software.
16927 | LOGGING_ADAPTER_ERROR | Error |  |  |  | fatal error in logging adapter layer | Meaning: Fatal error in logging adapter layer Reason: Logging adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help, replace PMI board.
16928 | FW_UDP_ADAPTER_ERROR | Error |  |  |  | fatal error in FW /UDP adapter layer | Meaning: Fatal error in FireWire (FW) resp. Ethernet (UDP) adapter layer Reason: FW/UDP adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and Ethernet connection; update SW; if this doesn't help, replace PMI board.
16929 | OUTPUT_BUFFER_FULL | Error | 266 | Output buffer full. Remove cassettes from the output buffer. |  | Output buffer is full, eject not possible. | Meaning: Ejecting cassette to output buffer is not possible. Reason: Output buffer is full. Cure: Remove cassettes from output buffer
16930 | CLEANUP_CASS_AND_PROCESSIBLE_IP_DETECTED_DEPRECATED | Error |  |  |  | DEPRECATED | DEPRECATED
16931 | CLEANUP_IP_WITHOUT_CASS | Error |  |  |  | CleanUp fatal Error: An IP was found in the digitizer, but no cassette! | Meaning: The cleanUp cycle during start-up failed with a fatal error. Reason: An IP was found in the digitizer, but no cassette. Cure: Remove IP from digitizer by hand and restart digitizer to start the self test.
16932 | CYCLE_FAILED_TO_FATAL_NO_ERROR_CODE | Error |  |  |  | DigitizerControl (DIC) received an error message during the cycle with status 'FAILED' which leads to a fatal error. No specific error code was passed over. | Meaning: DIC received an error message which leads to a fatal digitizer error. No specific error code was passed over. Reason: The error code is missing resp. the error should be handled specifically. This is a SW error. Cure: Remove cassettes and IP from digitizer and restart digitizer to start the self test. If error persists, update SW.
16933 | STARTUP_FAILED_TO_FATAL_NO_ERROR_CODE | Error |  |  |  | DigitizerControl (DIC) received an error message during the cycle with status 'FAILED' which leads to a fatal error. No specific error code was passed over. | Meaning: Digitizer Control (DIC) received an error message which leads to a fatal digitizer error. No specific error code was passed over. Reason: The error code is missing resp. the error should be handled specifically. This is a SW error. Cure: Remove cassettes and IP from digitizer and restart digitizer to start the self test. If error persists, update SW.
16934 | CLEANUP_EMPTY_CASS_DETECTED | Error | 106 | Empty cassette found at start-up. Remove cassette. | Confirm | CleanUp detected an empty cassette. | Meaning: The cleanUp cycle found an empty cassette without IP in the digitizer. Reason: Digitizer was switched off while clamping the empty cassette or empty cassette was inserted while digitizer was switched off. Thus the empty cassette was found at start-up. Cure: Confirm message, remove cassette and insert IP into the cassette.
16935 | NOBODY_WAITED_ON_ERROR_ACKNOWLEDGE | Error |  |  |  | Nobody waited on the confirmation or cancellation of an error, just remove the displayed error on PC | Meaning: An error message was confirmed or cancelled, but no SW module waited on the confirmation or cancellation of the specific error code. Reason: Digitizer was disconnected while an error message was displayed on the PC. Cure: None. This is only an internal warning for the R&D team.
16936 | CONFIRM_WRONG_DATAFLOW_SPEC | Error | 157 | This cassette is not supported by this digitizer. | Confirm | cassette is initialized with wrong data flow spec; spec version must be >= 1.3 | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong data flow specification (less than version 1.3) Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
16937 | CONFIRM_WRONG_IP_TYPE | Error | 157 | This cassette is not supported by this digitizer. | Confirm | wrong image plate type; must be an inflexible image plate on a carrier | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong setting of the image plate type; the image plate type must be set to Powder IP or Needle IP on carrier. Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
16938 | CONFIRM_WRONG_APP_TYPE | Error | 157 | This cassette is not supported by this digitizer. | Confirm | wrong application type; must be GenRad SR/HR | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong setting of the application type; the application type must be set to 'GenRad HR' or 'GenRad SR'. Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
16939 | CONFIRM_WRONG_MAMMO_SIZE | Error | 270 | Error during read-out of cassette chip data. Remove and re-identify the cassette. | Confirm | IP size is wrong for Mammo | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong setting of the application type or the IP size; a Mammo cassette must have the size 18x24 or 24x30. Cure: Check initialization of the cassette.
16940 | CONFIRM_WRONG_IP_SIZE | Error | 286 | Error occurred. Cassette is not correctly initialized. Remove the cassette and press the confirm button. Please contact service to fix the problem. Do not use the cassette again before an inspection. | Confirm | cassette is initialized with wrong IP size | Meaning: Cassette is rejected because IP size is wrong or cassette is initialized with wrong size. Reason: Cassette is rejected due to a wrong setting of the IP size. Cure: Check initialization of the cassette.
16941 | CONFIRM_UNSUPPORTED_IP_SIZE | Error | 157 | This cassette is not supported by this digitizer. | Confirm | IP size is not supported | Meaning: Cassette is rejected because IP size is not supported. Reason: IP size stored on RF-tag is not supported by this digitizer. Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
16942 | CONFIRM_CANCEL_ERASE_IDENTIFIED_CASSETTE | Error | 278 | Cassette already identified and possibly exposed - Press confirm to erase or cancel to scan the image plate and leave the erase mode. | Confirm | an identified cassette was detected in erase mode | Meaning: An identified cassette was detected in erase mode. The user is asked if the image plate should be erased or scanned. Reason: An identified cassette was detected in erase mode. Cure: Press 'confirm' to erase the image plate, press 'cancel' to scan the image plate.
16943 | CONFIRM_IP_OUT_OF_CASSETTE_FOUND_AT_STARTUP | Error | 280 | Image plate detected during self-test in the digitizer. The image plate was moved back into the cassette. Please check whether the image is present at the workstation and re-scan the cassette, if necessary. If the image is present, the cassette was possibly not erased. Erase the cassette again in any case. Remove the cassette and press the confirm button to continue. | Confirm | an image plate was found outside of the cassette in the digitizer during start-up | Meaning: An image plate was found outside of the cassette in the digitizer during start-up. It is not known if the IP was scanned and/or erased. Reason: User switched off the digitizer during a running cycle or after an error occurred. Cure: Remove the cassette, press 'confirm' to finish start-up. Note: if you don't remove the cassette and press 'confirm' , the cassette will be moved to the output buffer.
16944 | FATAL_ERROR_ASAP_DISPLAYS_ERROR | Error |  |  |  | fatal error, going into error state; ASAP will display error | Meaning: A fatal digitizer error occured which will be displayed by PC software component ASAP and not by the digitizer itself. Reason: A fatal digitizer error occured while a PC command was executed; therefore the fatal error code is sent in the response for the PC command and the PC (and not the digitizer) must display the corresponding error message. Cure: Check displayed error message.
16945 | FATAL_ERROR | Error |  |  |  | fatal error, going into error state | Meaning: Fatal error occurred, digitizer is in the error state and terminated current cycle. Reason: A SW component detected a fatal error. Cure: Check underlying error description.
16946 | FATAL_CMD_EJECT | Error |  |  |  | PC command 'eject' stopped with FATAL response | Meaning: PC command 'eject' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16947 | FATAL_CMD_ERASE | Error |  |  |  | PC command 'erase' stopped with FATAL response | Meaning: PC command 'erase' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16948 | FATAL_CMD_FEED_IP | Error |  |  |  | PC command 'feedIP' stopped with FATAL response | Meaning: PC command 'feedIP' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16949 | FATAL_CMD_RETURN_IP | Error |  |  |  | PC command 'returnIP' stopped with FATAL response | Meaning: PC command 'returnIP' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16950 | FATAL_CMD_SCAN | Error |  |  |  | PC command 'scan' stopped with FATAL response | Meaning: PC command 'scan failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16951 | FATAL_CMD_ENABLE_FEEDER | Error |  |  |  | enableFeeder stopped with FATAL response | Meaning: PC command 'enableFeeder' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16952 | FATAL_CMD_DISABLE_FEEDER | Error |  |  |  | PC command 'disableFeeder' stopped with FATAL response | Meaning: PC command 'disableFeeder' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16953 | FATAL_CMD_START_SELFTEST | Error |  |  |  | PC command 'startSelftest' stopped with FATAL response | Meaning: PC command 'startSelftest' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16954 | FAILED_CMD_SET_IPDATA | Error |  |  |  | PC command 'setIPData' stopped with FAILED response | Meaning: PC command 'setIPData' failed due to a fatal digitizer error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16955 | FAILED_CMD_STORE_IPCALDATA | Error |  |  |  | PC command 'storeIPCalData' stopped with FAILED response | Meaning: PC command 'storeIPCalData' failed due to a fatal digitizer error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16956 | FAILED_CMD_READ_IPCALDATA | Error |  |  |  | PC command 'readIPCalData' stopped with FAILED response | Meaning: PC command 'readIPCalData' failed due to a fatal digitizer error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16957 | FAILED_CMD_READ_IPCALDATA_LENGTH | Error |  |  |  | PC command 'readIPCalLength' stopped with FAILED response | Meaning: PC command 'readIPCalLength' failed due to a fatal digitizer error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
16958 | ERASING_IP_FAILED | Error | 274 | Error occurred. Please contact service to fix the problem. Make sure that cassette is erased before next use ! |  | Erasing the image plate failed. | Meaning: Erasing the image plate failed. Reason: 1) Erasure unit is not connected or not operating correctly. 2) Power supply for erasure unit defective. (See logging (file: log.xml) for details.) Cure: 1) Restart the digitizer to execute the selftest of the erasure unit and to get a more specific error code. 2) Exchange power supply for erasure unit.
16959 | ERASE_MODE_GRIP_IP_FAILED | Error | 277 | Image plate is not erased! Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. | Confirm | Grip IP in cassette failed in erase mode. | Meaning: Grip IP in cassette with snaphooks failed. Reason: 1) Any damage at cassette/IP. 2) Any defect at locking, snaphooks or IP-Transport mechanics. Cure: 1) Replace cassette/IP. 2) Try again; if fails permanently, check more specific error codes in infocounters; check mechanics with diagnose cycle of the service client
16960 | ERASE_MODE_FEED_FAILED | Error | 277 | Image plate is not erased! Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. | Confirm | Move IP down to begin of erasure failed. | Meaning: Move IP down to begin of erasure failed. Reason: 1) Any damage at IP. 2) Any problem with IP transport to erasure unit. Cure: 1) Replace IP. 2) Try again; if fails permanently, check more specific error codes in infocounters; check mechanics with diagnose cycle of the service client
16961 | PMM_ANALOG_POWER_FAILURE | Error |  |  |  | PMM analog power +20V or -10V failure | Meaning: PMM analog power +20V or -10V failure. Reason:PMI power distribution board defective , cable from PMI power distribution board to PMI board defective. Cure: 1) Check PMM cabling. 2) Check PMI power distribution board cabling. 3) Replace PMI power distribution board.
16962 | OPTICS_POWER_FAILURE | Error |  |  |  | Optics +12V or -12V failure | Meaning: Optics voltage error; Reason: PMI power supply fuse broken, cable power supply to PMI broken or PMI power supply defective; Cure: Replace broken fuses, replace cable power supply to PMI, replace PMI power supply
16963 | CALIBRATION_BOARD_POWER_FAILURE | Error |  |  |  | Calibration board power failure | Meaning: Calibration board power failure. Reason: PMI power distribution board defective, cable from PMI power distribution board to PMI board defective. Cure: 1) Check PMI power distribution board cabling. 2) Replace PMI power distribution board.
16964 | INITIALIZE_CAN | Error |  |  |  | Initialize CAN BUS failed | Meaning: Initialize CAN BUS failed; Reason: CAN hardware or software failure; Cure: 1) Check CAN BUS cabling; 2) Install new software; 3) Replace PMI board.
16965 | FATAL_INITIALIZATION_FAILED | Error |  |  |  | fatal error during initialization | Meaning: Fatal error occurred during initialization, digitizer is in the error state. Reason: A SW component detected a fatal error. Cure: Check underlying error description.
16966 | FATAL_CMD_CLEAN_UP | Error |  |  |  | PC command 'cleanUp' stopped with FATAL response | Meaning: Fatal digitizer error occured while PC command 'cleanUp' was executed. Reason: Check underlying error. Cure: Check underlying error.
16967 | CONFIRM_WRITING_TAG_FAILED | Error | 276 | Error occurred. Cassette chip data cannot be updated. Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. | Confirm | writing to RF-tag failed | Meaning: Writing data on RF-tag on IP failed. Reason: RF-tag on image plate is defective, tag reader is defective or not connected, cassette is not closed correctly. Cure: 1) Try again with another cassette. 2) Check cable between cassette handling board and tag reader. 3) Check mounting of tag reader 4) Replace tag reader. 5) Replace cassette handling board if error persists.
16968 | UNEXPECTED_PC_COMMAND | Error |  |  |  | received PC command which cannot be handled in the current state | Meaning: Digitizer Control received a PC command which cannot be handled in the current state of the digitizer (e.g. scan command if digitizer is in a fatal error state). The wrong PC command can be found in the logfile. Reason: SW error. Cure: Save logfile (do a complete backup), install new software; inform R&D team and send logfile to R&D
16969 | CONFIRM_CLEAN_UP_BEFORE_SCANNING | Error | 280 | Image plate detected during self-test in the digitizer. The image plate was moved back into the cassette. Please check whether the image is present at the workstation and re-scan the cassette, if necessary. If the image is present, the cassette was possibly not erased. Erase the cassette again in any case. Remove the cassette and press the confirm button to continue. | Confirm | digitizer received PC command 'cleanUp' before scanning started; the IP was not scanned and is transported back into the cassette | Meaning: Digitizer received PC command 'cleanUp' before scanning started; the IP was not scanned and is transported back into the cassette. Reason: PC connected to digitizer while digitizer was performing a cycle. Cure: Remove cassette and confirm message; insert cassette again to scan the image plate.
16970 | SCAN_WARNING_ICC | Error |  |  |  | ICC finished scan with warning | Meaning: The scan finished with a warning. The underlying error will be displayed as fatal error at end of the cycle. Reason: ICC (ImageCaptureControl) checked the scan and reported a scan unit (PMM, lens assembly) warning. Cure: Check description of underlying scan unit warning.
16971 | SUCK_IP_FAILED | Error | 221 | Image plate can not be sucked by the digitizer. The image plate was moved back into the cassette without scanning. Please insert cassette for scanning after repair again. |  | Sucking IP at stage failed. | Meaning: Sucking IP at stage failed. The vacuum is below the defined threshold. Reason: 1) IP is not fixed correctly by snaphooks. 2) Vacuum sensors do not work correctly. Cure: 1) Check fixation of IP by snaphooks. 2) Check seat of traverse on IP carrier plate. 3) Check hoses of vacuum sensors on IP handling board. 4) Replace IP handling board.
16972 | SECOND_PC_CONNECTS | Error |  |  |  | A second PC tries to connect to the digitizer. | Meaning: A second PC tries to connect to the digitizer. This is not supported. Reason: Two PCs are configured to connect to the same digitizer. Cure: The file 'd:\secondPCWhichConnects.txt' on digitizer's hard disk contains the name of the second PC. Change the digitizer configuration on this PC.
16973 | CANNOT_WRITE_2ND_PC_NAME_TO_FILE | Error |  |  |  | PC name cannot be written into file | Meaning: The name of the second PC which tries to connect to the digitizer cannot be written into the file 'd:\secondPCWhichConnects.txt'. Reason: Partition d: doesn't exist or is write protected. Cure: Check if the hard disk partition d: exists and check if it is write protected; unprotect it with the command 'unprot d:'.
16974 | CANNOT_CREATE_2ND_PC_NAME_FILE | Error |  |  |  | Cannot create file to save PC name | Meaning: The file 'd:\secondPCWhichConnects.txt' which saves the name of the second PC which tries to connect to the digitizer cannot be created. Reason: Partition d: doesn't exist or is write protected. Cure: Check if the hard disk partition d: exists and check if it is write protected; unprotect it with the command 'unprot d:'.
16975 | CANNOT_SEND_2ND_PC_CONNECTS_EVENT | Error |  |  |  | Cannot send the event 'second PC tries to connect' | Meaning: The Communicator callback method cannot send the event ''second PC tries to connect' to DigitizerControl. Reason: SW error. Cure: Install new software.
16976 | PREPARE_SCAN_ERROR | Error |  |  |  | prepareScan returned an error | Meaning: The preparation of the scan failed. Reason: Scan unit (PMT, lens assembly) error or incidence of light. Cure: Check underlying scan unit error description.
16977 | SET_IPDATA_ERROR | Error |  |  |  | setIPData passed over corrupted data | Meaning: the PC command 'setIPData' passed over corrupted data to the digitizer. Reason: SW error. Cure: Install new software.
16978 | PREPARE_SCAN_WARNING | Error |  |  |  | prepareScan returned a warning | Meaning: The preparation of the scan finished with a warning, the IP cannot be scanned. Reason: Scan unit (PMT, lens assembly) warning. Cure: Check underlying scan unit error description.
16979 | SCAN_REJECTED_SCAN_NOT_PREPARED | Error |  |  |  | scan rejected because prepare scan was not executed before | Meaning: Scan unit (PMT, optic module) was not prepared for scan, thus scaning cannot start and is rejected. Reason: Scan was not prepared, software error. Cure: 1) Re-start digitzer. 2) Install new software.
16980 | MOVE_IP_TO_STAGE_FAILED | Error | 293 | Error occurred. Image plate is not scanned and erased! Restart digitizer and try again. If error remains, contact service to fix the problem. |  | Moving IP to the stage failed. | Meaning: Moving the IP to the stage failed. The cassette was not opened and the IP is not scanned and not erased. Reason: The corresponding sensors or the vacuum unit do not work correctly. Cure: Restart the digitizer to perform a selftest of the digitizer. The selftest will display a specific error code, i.e. which sensor does not work.
16981 | MOVE_IP_TO_BOE_FAILED | Error | 294 | Error occurred. Image plate is not erased! Restart digitizer and try again. If error remains, contact service to fix the problem. |  | Moving IP to the BeginOfErasure (BOE) failed. | Meaning: Moving the IP to the BeginOfErasure (BOE) failed. The cassette was not opened and the IP is not erased. Reason: The corresponding sensors do not work correctly. Cure: Restart the digitizer to perform a selftest of the dgitizer. The selftest will display a specific error code, i.e. which sensor does not work.
16982 | DELAY_TIMER_ALREADY_IN_USE_WAIT_ICC_SCAN_END | Error |  |  |  | fatal SW error | Meaning: The DelayTimer is already running when the timer should be used for waiting a certain time to check if ICC (imageCaptureControl) finishes the scan in time after the slowscan drive stopped the scan. Reason: SW error. Cure: Install new software.
16983 | TIMEOUT_SCAN_END_ICC | Error |  |  |  | timeout for 'scan' response from ICC after slowscan stopped | Meaning: The scanning unit didn't finish the scan process although the slowscan motor has already stopped. Reason: ICC (ImageCaptureControl) is still waiting for end of the scan, i.e. the required number of pixels have not been stored in the SDRAM. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check supply voltages; check connection between PMI board and optics assembly; if OK, replace PMI; if still NOK, replace optics assembly
16984 | DIGITIZER_SW_VERSION_MISMATCH | Error | 298 | There is an inconsistency in the software packages installed. Contact service to finish installation. |  | The software versions of the internal digitizer modules do not match. | Meaning: Software versions loaded onto internal digitizer modules are not matching. Reason: Incomplete software installation, installation failed or hardware replacement (digitizer boards). Cure: Install digitizer software again.
16985 | INVALID_BUFFER_FROM_SPI | Error |  |  |  | spooler returned invalid buffer | Meaning: Software component 'spooler' returned invalid buffer (= NULL). Reason: Fatal software error. Cure: Install digitizer software again.
16986 | IMAGE_QUEUE_FULL | Error | 302 | Image queue full. Please wait. |  | spooler input queue is full | Meaning: Image queue of software component 'spooler' is full. Reason:The spooler's capacity is exhausted, no image can be scanned at the moment. Cure: Wait until images are transmitted to the processing stations. Be sure that the processing stations are connected and working properly.
16987 | UNKNOWN_DESTINATION | Error | 299 | Destination not known. Please identify the cassette at another ID station. | Confirm | destination is not known | Meaning: The destination is not known. Reason: The destination stored on the RF tag on the cassette is not known by the digitizer. Cure: Identify the cassette at another ID station.
16988 | CONFIRM_NOT_IDENTIFIED | Error | 262 | Cassette not identified. Identify the cassette. | Confirm | cassette is not identified | Meaning: The cassette is not identified. The cassette is rejected. Reason: The user did not identify the cassette, the erase mode is not activated. Cure: Identify the cassette at the ID station.
16989 | UNEXPECTED_SRVM_COMMAND | Error |  |  |  | received a command from the ServiceModule which cannot be handled in the current state | Meaning: Digitizer Control received a ServiceModule command which cannot be handled in the current state of the digitizer. The wrong command can be found in the logfile. Reason: SW error. Cure: Save logfile (do a complete backup), install new software; inform R&D team and send logfile to R&D
16990 | DELAY_TIMER_ALREADY_IN_USE_ENDURANCE_RUN | Error |  |  |  | fatal SW error | Meaning: The DelayTimer is already running when the timer should be used for waiting for a new cassette during the endurance run. Reason: SW error. Cure: Install new software.
16991 | CONFIRM_CANCEL_CALIBRATE_IDENTIFIED_CASSETTE | Error | 303 | Cassette already identified and possibly exposed - Press confirm to calibrate or cancel to return the image plate. | Confirm | an identified cassette was detected in calibration mode | Meaning: An identified cassette was detected in calibration mode. The user is asked if the image plate should be calibrated or returned into the cassette. Reason: An identified cassette was detected in calibration mode. Cure: Press 'confirm' to calibrate the image plate, press 'cancel' to return the image plate into the cassette.
16992 | CONFIRM_CALIBRATION_TERMINATED | Error | 304 | Remove the cassette and press the confirm button. | Confirm | calibration cycle terminated by user | Meaning: The calibration of the image plate was terminated by the user because an identified cassette was detected. The cassette was ejected into the intermediate position. Reason: User terminated the calibration of an identified cassette. Cure: Remove the cassette and press 'confirm'. Use an unidentified cassette for calibration.
16993 | EXCHANGED_COMPONENT_DETECTED | Error | 305 | Exchanged component detected. Please confirm exchange via Service Menu. |  | a new module serial number was detected at start-up, i.e. a component was exchanged. Input needed via the service menu | Meaning: At least one module serial number has changed since last start-up. The start-up is halted and waits for an input from the service technician. Reason: The service technician has exchanged a complete module or a board. Cure: Login into service menu and execute 'Configuration - Exchanged Components'.
16994 | CONFIRM_WRONG_SIZE_FOR_FLFS | Error | 110 | Wrong cassette type (wrong application). Remove cassette. | Confirm | wrong cassette format for FL/FS | Meaning: Cassette is rejected because FLFS (Full Leg/Full Spine) is not possible with this cassette format. FLFS can be only done with cassette format 35x43 cm. Reason: Wrong examination selected or wrong cassette format used for FLFS. Cure: Use only cassette format 35x43 cm for FLSF.
16995 | WRONG_ICN_PROTVER | Error |  |  |  | Protocol version of ICN is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ICN (Infocounter) is wrong. Cure: Install new software.
16996 | WRONG_ELC_PROTVER | Error |  |  |  | Protocol version of ELC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ELC (ELabelControl) is wrong. Cure: Install new software.
16997 | WRONG_IPC_PROTVER | Error |  |  |  | Protocol version of IPC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of IPC (ImageProcessing) is wrong. Cure: Install new software.
16998 | WRONG_SPI_PROTVER | Error |  |  |  | Protocol version of SPI is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of SPI (SpoolerInput) is wrong. Cure: Install new software.
16999 | WRONG_PRODUCT_TYPE | Error |  |  |  | The product type is wrong | Meaning: The product type read from the eLabels is not supported by this software. Reason: 1. This is not a DX-G or DX-M digitizer. 2. A wrong product type is stored in the eLabels. Cure: 1. Check if this digitizer is really a DX-G or DX-M. The name plate must show the type 5170. 2. Replace digitizer by a DX-G or DX-M.
17000 | ERASURE_UNIT_TIMEOUT | Error |  |  |  | A timeout had been triggered as the erasure unit had not been switched off in time. | Meaning: The standard workflow software seems to have stalled at least partly. Reason: Communication error (hardware or software) or software error Cure: Reset the digitizer, thus running the selftest. If the selftest is passed, there must have been a communication failure, escalate this error and provide the logfile.
17001 | ACTION_TIMER_ALREADY_IN_USE_START_OF_ERASURE | Error |  |  |  | fatal SW error | Meaning: The ActionTimer is already running when the timer should be used to check the erasure unit's on-time. Reason: SW error. Cure: Install new software.
17002 | GRIP_IP_ABORTED_SLOWSCAN_ERROR | Error | 265 | Image plate is not scanned and erased! Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. | Confirm | Grip IP in cassette aborted because the slowscan did not reach the home positon correctly after last scan. | Meaning: Grip IP in cassette aborted because the slowscan did not reach the home positon correctly after last scan. Reason: Any mechanical or electronic problem of slowscan drive. Cure: 1) Check for mechanic problems blocking the movement of the slowscan. 2) Check slowscan Reference sensor, motor and cabling. 3) Replace defective part(s), reference sensor, motor, cabeling or IPHandling Board, if error persists.
17408 | INVALID_CASSDET_STATUS | Error |  |  |  | invalid cassette detection status in input slot | Meaning: invalid status combination of locking cassette contact sensor and edge sensor. Edge sensor is pressed and locking cassette contact sensor is free. Reason: Either cassette slot or edge sensor defective and/or cabeling or cassette handling board. Cure: Replace defective part (sensor, cabeling, cassetteHandling board)
17409 | WRONG_CHC_PROTVER | Error |  |  |  | protocol version of CHC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of CHC (CassetteHandlingControl) is wrong. Cure: 1) Restart digitizer. 2) Install new software.
17410 | WRONG_IPHC_PROTVER | Error |  |  |  | protocol version of IPHC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of IPHC (ImagePlateHandlingControl) is wrong. Cure: 1) Restart digitizer. 2) Install new software.
18944 | INIT_FAILED | Error |  |  |  | The initialization of the Display failed | Meaning: A problem occurred during the initialization of the Display Reason: see logging Cure: have the bug fixed by RDE
19456 | TXFLDSENDMSG | Error |  |  |  | transmitFolder() - sendMsg failed |  | COM_DicomErr
19457 | TXFLDABORTASSOCIATION | Error |  |  |  | transmitFolder() - abortAssociation() failed |  | COM_DicomErr
19458 | CREATEOBJECT | Error |  |  |  | assembleDICOMObject() - createObject failed |  | COM_SWError
19459 | ADDPATNAME | Error |  |  |  | assembleDICOMObject() - adding attr PatientName failed |  | COM_SWError
19460 | ADDPATID | Error |  |  |  | assembleDICOMObject() - adding attr PatientBirthDatefailed |  | COM_SWError
19461 | ADDPATBIRTHDATE | Error |  |  |  | assembleDICOMObject() - adding attr PatientName failed |  | COM_SWError
19462 | ADDPATSEX | Error |  |  |  | assembleDICOMObject() - adding attr PatientSex failed |  | COM_SWError
19463 | ADDSTUDYINSTUID | Error |  |  |  | assembleDICOMObject() - adding attr StudyInstanceUID failed |  | COM_SWError
19464 | ADDSTUDYDATE | Error |  |  |  | assembleDICOMObject() - adding attr StudyDate failed |  | COM_SWError
19465 | ADDSTUDYTIME | Error |  |  |  | assembleDICOMObject() - adding attr StudyTime failed |  | COM_SWError
19466 | ADDACCESSIONNUMBER | Error |  |  |  | assembleDICOMObject() - adding attr AccessionNumber failed |  | COM_SWError
19467 | ADDREFPHYSICIAN | Error |  |  |  | assembleDICOMObject() - adding attr ReferringPhysiciansName failed |  | COM_SWError
19468 | ADDSTUDYID | Error |  |  |  | assembleDICOMObject() - adding attr StudyID failed |  | COM_SWError
19469 | ADDMODALITY | Error |  |  |  | assembleDICOMObject() - adding attr Modality failed |  | COM_SWError
19470 | OFFLINESERIESINSTUID | Error |  |  |  | assembleDICOMObject() - adding attr PatientName failed |  | COM_SWError
19471 | ADDSERIESNUMBER | Error |  |  |  | assembleDICOMObject() - adding attr SeriesNumber failed |  | COM_SWError
19472 | ADDLATERALITY | Error |  |  |  | assembleDICOMObject() - adding attr Laterality failed |  | COM_SWError
19473 | ADDBODYPART | Error |  |  |  | assembleDICOMObject() - adding attr BodyPart failed |  | COM_SWError
19474 | ADDPLATETYPE | Error |  |  |  | assembleDICOMObject() - adding attr PlateType failed |  | COM_SWError
19475 | ADDVIEWPOSITION | Error |  |  |  | assembleDICOMObject() - adding attr ViewPosition failed |  | COM_SWError
19476 | ADDMANUFACTURER | Error |  |  |  | assembleDICOMObject() - adding attr Manufacturer failed |  | COM_SWError
19477 | ADDMANUFACTURERMODELNAME | Error |  |  |  | assembleDICOMObject() - adding attr ManufacturerModelName failed |  | COM_SWError
19478 | ADDSERIALNUMBER | Error |  |  |  | assembleDICOMObject() - adding attr SerialNumber failed |  | COM_SWError
19479 | ADDSWVERSION | Error |  |  |  | assembleDICOMObject() - adding attr SoftwareVersion failed |  | COM_SWError
19480 | ADDIMAGETYPE | Error |  |  |  | assembleDICOMObject() - adding attr ImageType failed |  | COM_SWError
19481 | ADDACQDATE | Error |  |  |  | assembleDICOMObject() - adding attr AcquisitionDate failed |  | COM_SWError
19482 | ADDACQTIME | Error |  |  |  | assembleDICOMObject() - adding attr AcquisitionTime failed |  | COM_SWError
19483 | ADDINSTANCENUMBER | Error |  |  |  | assembleDICOMObject() - adding attr InstaceNumber failed |  | COM_SWError
19484 | ADDPATORIENT | Error |  |  |  | assembleDICOMObject() - adding attr PatientOrientation failed |  | COM_SWError
19485 | ADDIMAGECOMMENTS | Error |  |  |  | assembleDICOMObject() - adding attr ImageComments failed |  | COM_SWError
19486 | ADDPLATEID | Error |  |  |  | assembleDICOMObject() - adding attr PlateID failed |  | COM_SWError
19487 | ADDIMAGERPIXELSPACING | Error |  |  |  | assembleDICOMObject() - adding attr ImagerPixelSpacing failed |  | COM_SWError
19488 | ADDCASSETTESIZE | Error |  |  |  | assembleDICOMObject() - adding attr CassetteSize failed |  | COM_SWError
19489 | ADDEXPOSURESONPLATE | Error |  |  |  | assembleDICOMObject() - adding attr ExposuresOnPlate failed |  | COM_SWError
19490 | ADDRESCALEINTERC | Error |  |  |  | assembleDICOMObject() - adding attr RescaleIntercept failed |  | COM_SWError
19491 | ADDRESCALESLOPE | Error |  |  |  | assembleDICOMObject() - adding attr RescaleSlope failed |  | COM_SWError
19492 | ADDRESCALETYPE | Error |  |  |  | assembleDICOMObject() - adding attr RescaleType failed |  | COM_SWError
19493 | ADDPXLSAMPLES | Error |  |  |  | assembleDICOMObject() - adding attr PixelsPerSample failed |  | COM_SWError
19494 | ADDPHOTOMETRICINTERP | Error |  |  |  | assembleDICOMObject() - adding attr PhotometricInterpretation failed |  | COM_SWError
19495 | ADDROWS | Error |  |  |  | assembleDICOMObject() - adding attr Rows failed |  | COM_SWError
19496 | ADDCOLUMNS | Error |  |  |  | assembleDICOMObject() - adding attr Columns failed |  | COM_SWError
19497 | ADDALLOCEDBITS | Error |  |  |  | assembleDICOMObject() - adding attr AllocatedBits failed |  | COM_SWError
19498 | ADDSTOREDBITS | Error |  |  |  | assembleDICOMObject() - adding attr StoredBits failed |  | COM_SWError
19499 | ADDHIGHBIT | Error |  |  |  | assembleDICOMObject() - adding attr HighBit failed |  | COM_SWError
19500 | ADDPXLREPR | Error |  |  |  | assembleDICOMObject() - adding attr PixelRepresentation failed |  | COM_SWError
19501 | ADDSPECCHARSET | Error |  |  |  | assembleDICOMObject() - adding attr SpecificCharacterSet failed |  | COM_SWError
19502 | ADDINSTCREDATE | Error |  |  |  | assembleDICOMObject() - adding attr InstanceCreationDate failed |  | COM_SWError
19503 | ADDINSTCRETIME | Error |  |  |  | assembleDICOMObject() - adding attr InstanceCreationTime failed |  | COM_SWError
19504 | ADDSOPCLASSUID | Error |  |  |  | assembleDICOMObject() - adding attr SOPClassUID failed |  | COM_SWError
19505 | ADDSOPINSTUID | Error |  |  |  | assembleDICOMObject() - adding attr SOPInstanceUID failed |  | COM_SWError
19506 | ADDPRIVCREATOR | Error |  |  |  | assembleDICOMObject() - adding attr PrivateCreator failed |  | COM_SWError
19507 | ADDCASSETTEDATA | Error |  |  |  | assembleDICOMObject() - adding attr CassetteData failed |  | COM_SWError
19508 | ADDSCANDIRECTION | Error |  |  |  | assembleDICOMObject() - adding attr ScanDirection failed |  | COM_SWError
19509 | ADDIMG | Error |  |  |  | addImageFile() failed |  | COM_SWError
19510 | ADDCALLBACK | Error |  |  |  | addImageCallback() failed |  | COM_SWError
19511 | TXFLDRXMSG | Error |  |  |  | failed to receive a response from the SCP |  | COM_SWError
19512 | TXNDISRELEASE | Error |  |  |  | failed to release the association |  | COM_SWError
19513 | ADDIDSTUDYDESCRIPTION | Error |  |  |  | assembleDICOMObject() - adding attr StudyDescription failed |  | COM_SWError
19514 | ADDIDSERIESDESCRIPTION | Error |  |  |  | assembleDICOMObject() - adding attr SeriesDescription failed |  | COM_SWError
19515 | TXFINDPRESENTATIONCONTEXT | Error |  |  |  | transmitFolder - unable to find the presentation context ID |  | COM_DicomErr
19516 | DESTINATIONNOTFOUND | Error |  |  |  | destination not configured | the given destination name has not been found in the list of known destination. | COM_PSConfigurationErr
19517 | DESTINATIONNAMETOOLONG | Error |  |  |  | name of the destination is too long | the name of the destination exceeds 64 characters. | COM_PSConfigurationErr
19518 | DESTINATIONNOIPADDR | Error |  |  |  | This is not an ip address | The ipaddress exceeds the length of 15 characters. | COM_PSConfigurationErr
19519 | DESTINATIONCALLEDAETITLETOOLONG | Error |  |  |  | The called AE title of the destination is too long | The AE title of the destination exceeds 64 characters. | COM_PSConfigurationErr
19520 | NOUIDGENERATIONPOSSIBLE | Error |  |  |  | The generation of unique UIDs cannot be guaranteed | The system time used to generate unique identifiers as required by DICOM shows is not valid. Exchange the battery of the real-time clock or specify a valid timeserver (if applicable). | COM_SystemTimeErr
19521 | DESTINATIONPRECAETITLETOOLONG | Error |  |  |  | The AE title of the fast preview destination is too long | The AE title of the destination exceeds 64 characters. | COM_PSConfigurationErr
19522 | TXOFFADDIMGFILE | Error |  |  |  | failed to add imgfile hook | cannot add file reference to the image file into the DICOM object - image file existing? | COM_SWError
19523 | TXONADDIMGFILE | Error |  |  |  | failed to add img memory hook | cannot add hook to the image data in memory into the DICOM object - software failure! | COM_SWError
19524 | BUILDASSOCRQ | Error |  |  |  | failed to build the association request | check networking configuration (calling AE title, called AE title) | COM_PSConfigurationErr
19525 | SENDASSOCRQ | Error |  |  |  | failed to send association request | DTF library was not able to send the association request. | COM_SWError
19526 | CONNBLDABORTASSOCIATION | Error |  |  |  | cannot abort association | cannot abort assoc after problem when assembling the DICOM object | COM_SWError
19527 | RECVANSWERTIMEOUT | Error |  |  |  | error while waiting for the response | Error reported while waiting for the response of the DIMSE_sendMsg call. The image is queued again, rely on automatic recovery. | COM_DicomErr
19528 | ADDIMAGESINACQUISITION | Error |  |  |  | assembleDICOMObject() - adding attr images in acquisition failed |  | COM_SWError
19968 | IPDATA_NOTACCESSIBLE | Error |  |  |  | IPData NULL reference | The reference to IPData is NULL for SpoolerInput. | COM_SWError
19969 | SCANPAR_NOTACCESSIBLE | Error |  |  |  | ScanPar NULL reference | The reference to ScanParameters is NULL for SpoolerInput. | COM_SWError
19970 | CREATEEMPTYQ | Error |  |  |  | sc_qecreate failed for 'empty' queue | out of ressources or software bug | COM_SWError
19971 | CREATEQUEUEDQ | Error |  |  |  | sc_qecreate failed for 'queued' queue | out of ressources or software bug | COM_SWError
19972 | CREATETXEDQ | Error |  |  |  | sc_qecreate failed for 'transmitted' queue | out of ressources or software bug | COM_SWError
19973 | CREATEWARNQ | Error |  |  |  | sc_qecreate failed for 'warning' queue | out of ressources or software bug | COM_SWError
19974 | MALLOCBUFFERS | Error |  |  |  | cannot allocate memory for all buffers | out of memory | COM_SWError
19975 | ENQUEUEBUFFERINIT | Error |  |  |  | enqueuing a buffer to the 'empty' queue failed | either queue is full or QID error | COM_SWError
19976 | MARKQUEUED | Error |  |  |  | cannot push to 'queued' queue | software bug | COM_SWError
19977 | GETNEWBUFFEREMPTY | Error |  |  |  | error accessing the 'empty' queue | software bug | COM_SWError
19978 | GETNEWBUFFERTXED | Error |  |  |  | error accessing the 'transmitted' queue | software bug | COM_SWError
19979 | MARKEMPTY | Error |  |  |  | cannot push to 'empty' queue | software bug | COM_SWError
19980 | MARKTRANSMITTED | Error |  |  |  | cannot push to 'transmitted' queue | software bug | COM_SWError
19981 | MARKWARNING | Error |  |  |  | cannot push to 'warning' queue | software bug | COM_SWError
19982 | GETNEXTQUEUED | Error |  |  |  | cannot access the 'queued' queue | software bug | COM_SWError
19983 | GETNEXTWARNING | Error |  |  |  | cannot access the 'warning' queue | software bug | COM_SWError
19984 | MALLOCIMGBLOCK | Error |  |  |  | cannot allocate 64 kBytes for image block buffer | out of resources?!?! | COM_SWError
19985 | OPENIMGFILE | Error |  |  |  | cannot open image file | open imagefile to continue from disk failed | COM_SWErrorNeedsFormatFlash
19986 | IMAGELARGERTHANDEFINED | Error |  |  |  | the file found on disk is larger than the control data indicates | software bug | COM_SWError
19987 | CIS_OPENIMGFILE | Error |  |  |  | cannot open the image file in correctImageSize | software bug | COM_SWError
19988 | CIS_WRITEIMGFILE | Error |  |  |  | cannot fill incomplete image to required size | check hardware (CF-Card) | COM_SWErrorNeedsFormatFlash
19989 | DESTNOTAVAILABLE | Info | 301 | Image cannot be transferred. Destination {0} is not available. Check the connection to the destination. Use the routing functionality to send the image to another destination and press the confirm button. |  | destination is not available or returned error | Either the transmission failed due to a network problem or the target of an image transmission did not react or did reject the image for any reason. Check the destination. Image can be rerouted. | COM_PSConfigurationErr
19990 | WRITEPIXELSTODISK | Error |  |  |  | write error in imagefile | Meaning: writing the pixels to disk failed. Reason: CF-card defective or not enough memory left in partion e:/ Cure: replace CF-card. | COM_SWErrorNeedsFormatFlash
19991 | WRITECTRLTODISK | Error |  |  |  | write error in controlfile | Meaning: writing the controlfile to disk failed. Reason: CF-card defective or not enough memory left in partion e:/ Cure: replace CF-card. | COM_SWErrorNeedsFormatFlash
19992 | NODISKSPACE | Error |  |  |  | no enough space for spooler | Meaning: The available diskspace on e:/ is not sufficient for the spooler. Reason: Either too many files stored on e:/ or disk capacity decreased below an acceptable value. Cure: Exchange CF-card. | COM_SWErrorNeedsFormatFlash
19993 | DESTNOTAVAILABLE1DEST | Info | 314 | Image cannot be transferred. Destination {0} is not available. Check the connection to the destination and press the confirm button. | Confirm | destination is not available or returned error | Either the transmission failed due to a network problem or the target of an image transmission did not react or did reject the image for any reason. Check the destination. No other destination available. | COM_PSConfigurationErr
19994 | DEL_TIMEOUT | Error |  |  |  | Timeout deleting image. | Meaning: Waiting for the end of a previous delete image request timed out. New request will not be executed. Cure: Try again later or delete the image folder manually and reset. | COM_SWError
19995 | DEL_TIMEOUT2 | Error |  |  |  | Timeout deleting image. | Meaning: Waiting for the end of a delete image request timed out. Request was cancelled. Cure: Try again later or delete the image folder manually and reset. | COM_SWError
20480 | GET_PSDATA | Error |  |  |  | getCPFData_ps failed | Retrieving the PS data from the CPF file failed | COM_NotNeeded_ForServiceId
20481 | GET_APPLDATA | Error |  |  |  | getCPFData_appl failed | Retrieving the APPL data from the CPF file failed | COM_NotNeeded_ForServiceId
20482 | GET_NETDATA | Error |  |  |  | getCPFData_net failed | Retrieving the NET data from the CPF file failed | COM_NotNeeded_ForServiceId
20483 | GET_DIGDATA | Error |  |  |  | getCPFData_dig failed | Retrieving the DIG data from the CPF file failed | COM_SWError
20484 | ADD_DEST_TO_DS | Error |  |  |  | AddDataToStore failed for DS_CPF_DESTCOORD | Adding the Destination Coordinates to the Data Store failed | COM_SWError
20485 | ADD_STATIONNAME_TO_DS | Error |  |  |  | AddDataToStore failed for DS_CPF_DIGCOORD | Adding the Station Names Coordinates to the Data Store failed | COM_SWError
20486 | NOMEM_DEST | Error |  |  |  | malloc returned NULL | No memory can be allocated for the structure holding all Destinations | COM_SWError
20487 | GET_PRECDATA | Error |  |  |  | getCPFData_prec_failed | Retrieving the PREC data from the CPF file failed | COM_NotNeeded_ForServiceId
20488 | NOMEM_PREC | Error |  |  |  | malloc returned NULL | No memory can be allocated for the structure holding all precheck destinations | COM_SWError
20489 | MAINTENANCE_EXPECTED | Warning | 120 | Maintenance interval will expire within a short time. Please contact service. |  | Maintenance Interval will expire within a short time. Please contact service. |  | COM_NotNeeded_ForServiceId
20490 | MAINTENANCE_REQUIRED | Warning | 258 | Maintenance interval expired. Please contact service. |  | Maintenance Interval is expired. Please contact service as soon as possible. |  | COM_NotNeeded_ForServiceId
20491 | GET_IDENDATA | Error |  |  |  | getCPFData_iden failed | Retrieving the IDEN data from the CPF file failed | COM_NotNeeded_ForServiceId
20492 | INFOCOUNTER_RECREATED | Error |  |  |  | Infocounter data file recreated with default values. | Infocounter data file was recreated with default values, because original and backup data files were corrupted or lost. | COM_NotNeeded_ForServiceId
20992 | GROUPNAME_SMALLER_THAN_ZERO | Error |  |  |  | The group name is smaller than zero |  | COM_NotNeeded_ForServiceId
20993 | GROUPNAME_TOO_BIG | Error |  |  |  | The group name is too large |  | COM_NotNeeded_ForServiceId
20994 | NOBUFFER | Error |  |  |  | InitializeDataStore failed | Not enough memory could be allocated to hold the Data Store | COM_SWError
20995 | NOT_ENOUGH_MEMORY | Error |  |  |  | AddDataToStore failed | Not enough memory could be allocated | COM_SWError
20996 | NOT_INITIALIZED | Error |  |  |  | DataStore not initialized yet | Call 'InitializeDataStore' first. | COM_SWError
20997 | TIMEOUT | Error |  |  |  | lock already taken | A lock was taken and caused a waiting method to time out | COM_SWError
21504 | General | Error |  |  |  | General image processing error occurred | Meaning: A non-specific error occurred during image processing. Reason: Unknown. Cure: Unknown. | COM_SWError
21505 | MFS_FsLength | Error |  |  |  | Too few space for CRFileSystem. | Meaning: While trying to mount CRFileSystem, detected too few disk space. Reason: The flash memory on the photomultiplier unit is not existing or too small. Cure: Replace photomultiplier unit. | COM_PMMDataAccessErr
21506 | MFS_Mount | Error |  |  |  | Could not mount CRFileSystem. | Meaning: Could not mount CRFileSystem. Reason: . Cure: Perform a scan line calibration. | COM_PMMDataAccessErr
21507 | PRP_NoCRImage | Error |  |  |  | No valid CRImage given for preparing processing. | Meaning: The event to prepare processing contained no valid CRImage. Reason: Programming error. Either CRImage or one of its elements is a NULL pointer. Cure: Install another software version. | COM_SWError
21508 | PRP_NoDoseCalPIP | Error |  |  |  | Dose detection for PIP not calibrated. | Meaning: Dose detection for powder image plates not calibrated. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: NIP may be used. For PIP, replace photomultiplier module. | COM_PMMDataAccessErr
21509 | PRP_NoDoseCalNIP | Error |  |  |  | Dose detection for NIP not calibrated. | Meaning: Dose detection for needle image plates not calibrated. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: PIP may be used. For NIP, replace photomultiplier module. | COM_PMMDataAccessErr
21514 | PRO_NoCRImage | Error |  |  |  | No valid CRImage given for processing. | Meaning: The event to process an image contained no valid CRImage. Reason: Programming error. Either CRImage or one of its elements is a NULL pointer. Cure: Install another software version. | COM_SWError
21515 | PRO_WriteRaw | Error |  |  |  | Writing raw pixels to file failed. | Meaning: Writing raw pixels to file failed. Reason: Unknown. Cure: Ignore. | COM_NotNeeded_ForServiceId
21524 | INI_NoDoseCal | Error |  |  |  | Found no dose detection calibration at all. | Meaning: Found no dose detection calibration at all. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: Replace photomultiplier module. | COM_PMMDataAccessErr
21525 | INI_NoCRTools | Error |  |  |  | CRTools not started. | Meaning: At system start, the CRTools library was not found in memory. Reason: CRTools was not started. Cure: Install software again. | COM_SWError
21526 | INI_IncompatibleCRTools | Error |  |  |  | CRTools incompatible. | Meaning: The CRTools library is incompatible. Reason: The library CRTools does not fit to the rest of the application software. This is software configuration error and should never occur in the field. Cure: Try to install software again. If this doesn't help, install another software. | COM_SWError
21554 | IMB_INI_BufferCreation | Error |  |  |  | Creation of image buffer failed. | Meaning: Could not create a buffer for the image to scan. Reason: Memory for image buffer could not be allocated. Too few memory or problem with other parts of the software. Cure: Install the complete system again. If this fails, install another software version. Replace the CPU board. | COM_SWError
21555 | IMB_PRP_TooManyColumns | Error |  |  |  | Too many columns for preparing image buffer. | Meaning: At prepare operation, the number of coulumns and the block size have a wrong relation. Reason: Programming error. Cure: Install another software version. | COM_SWError
21556 | IMB_GNW_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21557 | IMB_GNP_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21558 | IMB_GNP_Timeout | Error |  |  |  | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21559 | IMB_GNR_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21560 | IMB_GNR_Timeout | Error |  |  |  | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21561 | IMB_GAR_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21562 | IMB_GAR_Timeout | Error |  |  |  | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21563 | IMB_GAR_Occupied | Error |  |  |  | All read blocks occupied. | Meaning: Could not get a new image block to read. All occupied. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21564 | IMB_RAR_InvalidID | Error |  |  |  | Tried to release an image block with an invalid ID. | Meaning: Tried to release an image block, which has not been reserved. Reason: Programming error. Cure: Ignore this error, if it does not cause any problem. | COM_NotNeeded_ForServiceId
21565 | IMB_GNW_Overflow | Error | 335 | Image buffer full, image scan stopped. Restart the digitizer and scan the image plate(s) again. |  | Image buffer overflow | Meaning: The image buffer has not been read out fast enough, so that it is not ready for further image data. Scanning has been stopped. Reason: Image data was not stored or sent fast enough. Most probably network problem, but may also be too slow flash memory. Cure: - Check and correct the network connection. Available image data will be retransmitted automatically. Note that by switching off the device parts of the image data may be lost. - If the network and the PC installation are ok and the error occurs again, replace the digitizer CPU board. - At devices with CF-card try if replacing the flash memory by a faster one helps. | COM_SWError
21566 | IMB_GAR_Released | Error |  |  |  | Image data no longer available. | Meaning: Image data is no longer available. The memory location is reused in circular buffer. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21567 | IMB_RBL_NoData | Error |  |  |  | Attempt to release more data than available. | Meaning: Attempt to release more data than available. Corrected to actual value. Reason: Programming error. Cure: Only warning, should be harmless. In case of problems install another software version. | COM_NotNeeded_ForServiceId
21604 | ISP_SignalRejected | Error |  |  |  | The incoming signal was rejected in this state. | Meaning: Image splitter got a signal in a state where it cannot be handled. Reason: Programming error. Cure: In case of problems install another software version. | COM_SWError
21605 | ISP_ImageQueueFull | Error | 302 | Image queue full. Please wait. |  | The spooler image queue is full. | Meaning: Image queue of software component spooler is full. Reason: The spooler's capacity is exhausted, no image can be scanned at the moment. Cure: Wait until images are transmitted to the processing station. Be sure that the processing station is connected, configured correctly and working properly. | COM_NotNeeded_ForServiceId
22016 | SYNCHRONIZING_FAILED | Error |  |  |  | Synchronizing failed.
22017 | DETECT_CONTROLLERS_FAILED | Error |  |  |  | detectControllers failed. | Meaning: detectControllers failed. Reason: Not all controllers were detected. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels.
22018 | DETECT_ELABELS_FAILED | Error |  |  |  | detectELabels failed. | Meaning: detectELabels failed. Reason: Not all eLabels were detected. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels.
22019 | ELABEL_MISSING | Error |  |  |  | Not all controllers and/or eLabels are connected. | Meaning: Not all controllers and/or eLabels are connected. Reason: Not all controllers and/or eLabels were detected. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels.
22020 | READ_ALL_ELABELS_FAILED | Error |  |  |  | readAllEntries failed. | Meaning: readAllEntries failed. Reason: Not all entries of all eLabels could be read. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels.
22021 | DATASTOREDATA_FILE_NOT_AVAILABLE | Error |  |  |  | DataStoreData file does not exist. | Meaning: DataStoreData file does not exist. Reason: DataStoreData file not stored on CF card. Cure: Please restore this file via Restore.
22022 | DATASTOREDATA_FILE_CORRUPT | Error |  |  |  | DataStoreData file is corrupt. | Meaning: DataStoreData file is corrupt. Reason: DataStoreData file could not opened. Cure: Please restore this file via Restore.
22023 | OrderRepairToolForOptics | Error |  |  |  | Parameter set of optic module is corrupted. | Meaning: Parameter set of optic module is corrupted. Reason: The structure of the eLabel on the optic module is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted optic module.
22024 | OrderRepairToolForPMT | Error |  |  |  | Parameter set of PMT/light collector module is corrupted. | Meaning: Parameter set of PMT/light collector module is corrupted. Reason: The structure of the eLabel on PMT is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted PMT/light collector module.
22025 | OrderRepairToolForHandling | Error |  |  |  | Parameter set of handling board is corrupted. | Meaning: Parameter set of handling board is corrupted. Reason: The structure of the eLabel on the handling board is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted handling board.
22026 | OrderRepairToolForScanner | Error |  |  |  | Parameter set of scanner module is corrupted. | Meaning: Parameter set of scanner module is corrupted. Reason: The structure of the eLabel on the scanner module is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted scanner module.
22027 | OrderRepairToolForCassClamping | Error |  |  |  | Parameter set of cassette clamping module is corrupted. | Meaning: Parameter set of cassette clamping module is corrupted. Reason: The structure of the eLabel on the cassette clamping module is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted cassette clamping module.
22028 | AUTOMATIC_REGISTRATION_FAILED | Error |  |  |  | Automatic Registration failed. | Meaning: Automatic Registration failed.
22029 | REPAIR_ELABEL_FAILED | Error |  |  |  | Repair of corrupted eLabel failed. | Meaning: Repair of corrupted eLabel failed.
22030 | WrongPMIModuleConfiguration | Error |  |  |  | Wrong PMI module configuration. | Meaning: Wrong PMI module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace PMI module.
22031 | WrongOpticsModuleConfiguration | Error |  |  |  | Wrong Optics module configuration. | Meaning: Wrong Optics module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace Optics module.
22032 | WrongPMTModuleConfiguration | Error |  |  |  | Wrong PMT module configuration. | Meaning: Wrong PMT module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace PMT module.
22033 | GenradPMTInMammoDigitizer | Error |  |  |  | Genrad PMT detected in DX-M. | Meaning: Genrad PMT detected in DX-M. Reason: Cross-exchange of the PMT with a DX-G. Cure: Replace PMT.
22034 | SparePartExistsNoDataStoreDataAvailable | Error |  |  |  | Sparepart(Module serialnumber 9999) exists but no data are available. | Meaning: Sparepart(Module serialnumber 9999) exists. Reason: Sparepart(Module serialnumber 9999) exists, but no registration was made. Cure: Replace sparepart with the old board and restart digitizer.
22528 | CRT_Base | Error |  |  |  | General error occurred in CRTools library | Meaning: A non-specific error occurred in the CRTools library. Reason: This error should normally not occur. Its value is used as base for all other CRTools errors. Cure: Ignore this error. | COM_SWError
22529 | PXC_InsufficientCalWidth | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Calibration line does not cover the complete width of the image to be corrected. Parts of the image will not be corrected. | Meaning: Only parts of the image will be corrected. The rest might have banding artefacts. Reason: Calibration line does not cover the complete width of the image to be corrected. Cure: Calibrate with the biggest used cassette. | COM_SWShadingCalibration
22530 | PXC_CalcCorrLines | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Calculating the correction lines failed. Using constant correction line. | Meaning: The image is uncorrected. Reason: An exception occurred during preparation of the correction line. Might be software bug. Cure: Check the log file to find out more information about the error. | COM_SWShadingCalibration
22531 | PXC_NoCalibrationLine | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | No calibration line loaded to correct the image. Using constant correction line. | Meaning: The image is uncorrected. Reason: Calibration line of this image plate type not available, invalid or could not be loaded. Cure: Calibrate the image plate type, which has generated the error. | COM_SWShadingCalibration
22532 | PXC_InvalidIPType | Error |  |  |  | An invalid IPType was given when clearing the calibration line. |  | COM_SWError
22533 | PXC_ExtrapolateCorrLine | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Extrapolating the correction lines failed. Using constant correction line. | Meaning: The image is uncorrected. Reason: An exception occurred during preparation of the correction line. Might be software bug. Cure: Check the log file to find out more information about the error. | COM_SWShadingCalibration
22534 | PXC_UnsupportedSamplesOrBPP | Error |  |  |  | Unsupported samples or bytes per pixel. |  | COM_SWError
22535 | PXC_UnsupportedSamplesOrBPP2 | Error |  |  |  | Unsupported samples or bytes per pixel. |  | COM_SWError
22536 | PXC_UnsupportedColumns | Error |  |  |  | Unsupported number of pixels per line. | Meaning: The requested number of pixels per line is negative or too small. Reason: Software bug. Cure: Install another software. | COM_SWError
22537 | PXC_UnsupportedLines | Error |  |  |  | Unsupported number of lines. | Meaning: The requested number lines is negative or too small. Reason: Software bug. Cure: Install another software. | COM_SWError
22538 | LCL_CreateCalibrationLine | Error |  |  |  | Creating a new CalibrationLine failed. |  | COM_SWError
22539 | LCL_WriteToCRFile | Error |  |  |  | Writing CalibrationLine to PMM failed. |  | COM_SWError
22540 | DCL_CreateCalibrationLine | Error |  |  |  | Creating a new dose CalibrationLine failed. |  | COM_SWError
22541 | DCL_WriteToCRFile | Error |  |  |  | Writing dose CalibrationLine to PMM failed. |  | COM_SWError
22542 | PXC_LoadIPCalNullResult | Error |  |  |  | PixelCorrector_LoadIPCal called with result null pointer. | Meaning: Null pointer in fuction call. Cannot execute. Reason: Software error. Cure: Install new software. | COM_SWError
22543 | PXC_CorrectLinesParameter | Error |  |  |  | PixelCorrector_CorrectLines called with invalid parameter. | Meaning: PixelCorrector_CorrectLines called with invalid parameter. Reason: Software error. Cure: Install new software. | COM_SWError
22544 | PXC_EOIBlocksPending | Error |  |  |  | There are still image blocks pending at end of image. Software error. |  | COM_SWError
22548 | CIM_UnknownImageFormat | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Reading IP calibration image failed because of unknown image file format. |  | COM_Image2DCalibrationErr
22549 | CIM_ImageNotLoaded | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Reading IP calibration image failed. |  | COM_Image2DCalibrationErr
22550 | CIM_ReadImageMemoryException | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Exception at reading IP calibration image from ID tag. |  | COM_Image2DCalibrationErr
22551 | CIM_ReadImageFileNameMismatch | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Wrong IPSerialNumber in IP calibration file. | Meaning: The name of the IP calibration image is different than the IPSerialNumber, stored in the file. Reason: Probably manual renaming of the calibration file. Cure: Never use an IP calibration file for another image plate. | COM_Image2DCalibrationErr
22552 | CIM_ReadImageFileException | Error |  |  |  | Exception at reading IP calibration image from file. |  | COM_Image2DCalibrationErr
22553 | CIM_WriteImageMemoryException | Error |  |  |  | Exception at writing IP calibration image to memory for storing on ID tag. |  | COM_Image2DCalibrationErr
22554 | CIM_WriteImageFileException | Error |  |  |  | Exception at writing IP calibration image to file. |  | COM_Image2DCalibrationErr
22555 | CIM_WriteNoImage | Error |  |  |  | Writing IP calibration image failed because no image present. Software error. |  | COM_Image2DCalibrationErr
22558 | CIM_RescaleNoImage | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | No image to rescale. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration or IP correction was disabled for this scan. Cure: Return the image plate to production for calibration. | COM_Image2DCalibrationErr
22559 | CIM_RescaleIPCalDotsPerMeter | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | The IP calibration image (TIFF file) contains wrong values for DotsPerMeter. | Meaning: Parameter error in the image plate calibration image. Reason: This could be prototype plate from DX-S development. But usually, this problem should not occur. Cure: Delete the calibration image from the hard disk and return the image plate to production for calibration. | COM_Image2DCalibrationErr
22560 | CIM_RescaleIPCalFailed | Error |  |  |  | Rescaling the given IP calibration image failed. |  | COM_Image2DCalibrationErr
22561 | CIM_RescaleException | Error |  |  |  | Exception at rescaling IP calibration image. |  | COM_Image2DCalibrationErr
22568 | CIM_ShiftAddCutNoCal | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | No IP calibration image available. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration. Cure: Return the image plate to production for calibration. | COM_Image2DCalibrationErr
22569 | CIM_ShiftAddCutException | Error |  |  |  | Exception at shifting / extending / cutting IP calibration image. |  | COM_Image2DCalibrationErr
22570 | CIM_ShiftAddCutTopExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many lines on top of the calibration image had to be extrapolated. | Meaning: Too many lines on top of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22571 | CIM_ShiftAddCutBottomExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many lines at the bottom of the calibration image had to be extrapolated. | Meaning: Too many lines at the bottom of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22572 | CIM_ShiftAddCutLeftExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many columns at the left side of the calibration image had to be extrapolated. | Meaning: Too many columns at the left side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22573 | CIM_ShiftAddCutRightExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many columns at the right side of the calibration image had to be extrapolated. | Meaning: Too many columns at the right side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22574 | CIM_AddIPCalFile | Error |  |  |  | Could not load IPCalFile. |  | COM_Image2DCalibrationErr
22575 | CIM_AddInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22576 | CIM_WriteMemoryReadFile | Error |  |  |  | Could not read IPCalFile to write to memory. |  | COM_Image2DCalibrationErr
22577 | CIM_WriteMemoryInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22578 | CIM_WriteFileInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22579 | CIM_WriteFileRescaledInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22580 | CIM_RescaleAndFlipException | Error |  |  |  | Error occurred while rescaling a calibration image. |  | COM_Image2DCalibrationErr
22581 | CIM_WriteFileTrailerRead | Error |  |  |  | Could not read IPCalFile trailer. |  | COM_Image2DCalibrationErr
22582 | CIM_WriteFileTrailerWrite | Error |  |  |  | Could not write IPCalFile trailer. |  | COM_Image2DCalibrationErr
22583 | CIM_ReadFileTrailerInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22584 | CIM_HeaderVersion | Error |  |  |  | Wrong IPCalibration header version. |  | COM_Image2DCalibrationErr
22585 | CIM_HeaderTruncated | Error |  |  |  | IPCal header truncated. |  | COM_Image2DCalibrationErr
22586 | CIM_DataTruncated | Error |  |  |  | IPCalibration data truncated. |  | COM_Image2DCalibrationErr
22587 | CIM_HeaderSize | Error |  |  |  | Wrong IPCal header size. |  | COM_Image2DCalibrationErr
22588 | CIM_HeaderType | Error |  |  |  | Wrong IPCal header type. |  | COM_Image2DCalibrationErr
22589 | CIM_IPCalHeaderVersion | Error |  |  |  | Wrong IPCal header version. |  | COM_Image2DCalibrationErr
22590 | CIM_IPCalHeaderTruncated | Error |  |  |  | IPCalFile header truncated. |  | COM_Image2DCalibrationErr
22591 | CIM_IPCalDataTruncated | Error |  |  |  | Not enough IPCalFile data. |  | COM_Image2DCalibrationErr
22592 | CIM_IPCalHeaderSize | Error |  |  |  | Wrong IPCalFile header size. |  | COM_Image2DCalibrationErr
22593 | CIM_Checksum | Error |  |  |  | IPCalibrationData checksum error. |  | COM_Image2DCalibrationErr
22594 | CIM_FileFormat | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Unknown image file format. |  | COM_Image2DCalibrationErr
22595 | CIM_SerialNumberMismatch | Error |  |  |  | The image file name does not match the IP serial number. |  | COM_Image2DCalibrationErr
22596 | CIM_WriteMNotValid | Error |  |  |  | The IPCalibrationData to be written to tag is not valid. | Meaning: The IP calibration data to be written to the tag is not valid. Reason: Probably software error. Cure: Install new software. | COM_Image2DCalibrationErr
22597 | CIM_WriteMTooBig | Error |  |  |  | The IPCalData to be written to tag is too big. | Meaning: Detected that the given IP calibration data does not fit onto the tag. Reason: Probably software error. Cure: Install new software. | COM_Image2DCalibrationErr
22598 | CIM_WriteMNotValid2 | Error |  |  |  | The IPCalibrationData to be written to tag is not valid. |  | COM_Image2DCalibrationErr
22599 | CIM_WriteNotValid | Error |  |  |  | The IPCalFile data to be written to file is not valid. |  | COM_Image2DCalibrationErr
22600 | CIM_NoFile | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Could not open CalibrationImage file. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: . Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22601 | CIM_IPCalDataVersion | Error |  |  |  | Wrong IPCalData version. |  | COM_Image2DCalibrationErr
22602 | CIM_IPCalDataVersion2 | Error |  |  |  | Wrong IPCalData version. |  | COM_Image2DCalibrationErr
22603 | CIM_IPCalDataType | Error |  |  |  | Wrong IPCalData type. |  | COM_Image2DCalibrationErr
22604 | CIM_IPCalDataTrailerLen | Error |  |  |  | IPCalFile trailer longer than data. |  | COM_Image2DCalibrationErr
22605 | CIM_IPCalFileTooSmall | Error |  |  |  | IPCalFile too small. |  | COM_Image2DCalibrationErr
22606 | CIM_IPCalFileChecksum | Error |  |  |  | IPCalFile checksum error. |  | COM_Image2DCalibrationErr
22607 | CIM_WriteThumbnailInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22608 | CIM_WriteThumbnailNoData | Error |  |  |  | CalibrationImage object contains no image data. |  | COM_Image2DCalibrationErr
22609 | CIM_WriteThumbnailRescale | Error |  |  |  | Rescaling the CalibrationImage failed. |  | COM_Image2DCalibrationErr
22610 | CIM_WriteThumbnailConvert | Error |  |  |  | Converting the CalibrationImage to 8 bits failed. |  | COM_Image2DCalibrationErr
22611 | CIM_AddNewAllocImage | Error |  |  |  | Could not allocate new image, when adding a new (empty) IP calibration image. |  | COM_Image2DCalibrationErr
22612 | CIM_AddNewIPCalFile | Error |  |  |  | Adding a new (empty) IP calibration image failed. |  | COM_Image2DCalibrationErr
22613 | CIM_AddNewInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22614 | CIM_CalibrateInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22615 | CIM_CalibrateNoData | Error |  |  |  | CalibrationImage object contains no image data. |  | COM_Image2DCalibrationErr
22616 | CIM_NormalizeInvalid | Error |  |  |  | Invalid image for normalization. |  | COM_Image2DCalibrationErr
22617 | CIM_Normalize0 | Error |  |  |  | Could not normalize image. All pixels are 0. |  | COM_Image2DCalibrationErr
22618 | CIM_NoBinaryData | Error |  |  |  | No data given to create a CalibrationImage from. |  | COM_Image2DCalibrationErr
22619 | CIM_NoSingleBlock | Error |  |  |  | ImageArea does not consist of one single block. | Meaning: The ImageArea for line artefact correction does not consist of one single block. No correction done. Reason: Software error. Cure: No cure. The image is ok, besides the fact that a possible line artefact is not corrected. | COM_Image2DCalibrationErr
22628 | CLN_CreateType | Error |  |  |  | Invalid CalibrationLine file type. |  | COM_ImageCalibrationLineErr
22629 | CLN_CreatePixelWidth | Error |  |  |  | Invalid pixelWidth in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22630 | CLN_CreateNLines | Error |  |  |  | Invalid number of calibration lines in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22631 | CLN_CreateNColumns | Error |  |  |  | Invalid number of columns in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22632 | CLN_ROILeft | Error |  |  |  | Invalid ROILeft in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22633 | CLN_CreateBitsPP | Error |  |  |  | Unsupported number of bits per pixel in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22634 | CLN_CreateOpenFile | Error |  |  |  | Opening the CalibrationLine file failed. |  | COM_ImageCalibrationLineErr
22635 | CLN_CreateOpenCRFile | Error |  |  |  | Opening the CalibrationLine file in CRFileSystem failed. |  | COM_ImageCalibrationLineErr
22636 | CLN_DeserializeInvalid | Error |  |  |  | Invalid CalibrationLine object. |  | COM_ImageCalibrationLineErr
22637 | CLN_DeserializeLength | Error |  |  |  | CalibrationLine file too small. |  | COM_ImageCalibrationLineErr
22638 | CLN_DeserializeFileType | Error |  |  |  | Invalid CalibrationLine file type. |  | COM_ImageCalibrationLineErr
22639 | CLN_DeserializeHeaderLength | Error |  |  |  | Invalid CalibrationLine file header length. |  | COM_ImageCalibrationLineErr
22640 | CLN_DeserializeHeaderLength2 | Error |  |  |  | Invalid CalibrationLine file header length. |  | COM_ImageCalibrationLineErr
22641 | CLN_DeserializeHeaderLength1 | Error |  |  |  | Invalid CalibrationLine file header length. |  | COM_ImageCalibrationLineErr
22642 | CLN_DeserializeVersion | Error |  |  |  | Invalid version of CalibrationLine file. |  | COM_ImageCalibrationLineErr
22643 | CLN_DeserializeLength2 | Error |  |  |  | CalibrationLine file too small. |  | COM_ImageCalibrationLineErr
22644 | CLN_DeserializePixelWidth1 | Error |  |  |  | Invalid pixelWidth in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22645 | CLN_DeserializePixelWidth | Error |  |  |  | Invalid pixelWidth in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22646 | CLN_DeserializeNColumns | Error |  |  |  | Invalid number of columns in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22647 | CLN_DeserializeBitsPP | Error |  |  |  | Unsupported number of bits per pixel in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22648 | CLN_DeserializeLength3 | Error |  |  |  | CalibrationLine file too small. |  | COM_ImageCalibrationLineErr
22649 | CLN_DeserializeNLines | Error |  |  |  | Invalid number of calibration lines in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22650 | CLN_DeserializeROILeft | Error |  |  |  | Invalid ROILeft in calibration file. |  | COM_ImageCalibrationLineErr
22651 | CLN_DeserializeLength4 | Error |  |  |  | Too few data to deserialize calibration line. |  | COM_ImageCalibrationLineErr
22652 | CLN_SerializeInvalid | Error |  |  |  | CalibrationLine_Serialize: Empty calibration line. |  | COM_ImageCalibrationLineErr
22653 | CLN_SerializeHeaderBytes | Error |  |  |  | CalibrationLine_Serialize: Wrong HeaderBytes (software bug). |  | COM_ImageCalibrationLineErr
22654 | CLN_NormalizeInvalid | Error |  |  |  | CalibrationLine_Normalize: Empty calibration line. |  | COM_ImageCalibrationLineErr
22655 | CLN_ClipInvalid | Error |  |  |  | Clip: Empty calibration line. |  | COM_ImageCalibrationLineErr
22656 | CLN_MaximumSlopeInvalid | Error |  |  |  | CalibrationLine_MaximumSlope: Empty calibration line. |  | COM_ImageCalibrationLineErr
22657 | CLN_MaximumSlopeNegative | Error |  |  |  | CalibrationLine_MaximumSlope: Unexpected negative value of maxDifference. |  | COM_ImageCalibrationLineErr
22658 | CLN_InterpolateInvalid | Error |  |  |  | CalibrationLine_Interpolate: Empty calibration line. |  | COM_ImageCalibrationLineErr
22659 | CLN_InterpolatePixelWidth | Error |  |  |  | CalibrationLine_Interpolate: Invalid new pixel width. |  | COM_ImageCalibrationLineErr
22660 | CLN_RescaleNNMatch | Error |  |  |  | CalibrationLine_RescaleNearestNeighbor: Length of input line and number of sections do not match. |  | COM_ImageCalibrationLineErr
22661 | CLN_RescaleCubicMatch | Error |  |  |  | CalibrationLine_RescaleCubic: Length of input line and number of sections do not match. |  | COM_ImageCalibrationLineErr
22662 | CLN_ExtrapolateInvalid | Error |  |  |  | CalibrationLine_Extrapolate: Empty calibration line. |  | COM_ImageCalibrationLineErr
22663 | CLN_ExtrapolateCut | Error |  |  |  | CalibrationLine_Extrapolate: Cannot cut more left columns than available. |  | COM_ImageCalibrationLineErr
22664 | CLN_ExtrapolateExtend | Error |  |  |  | CalibrationLine_Extrapolate: Cannot extend beyond begin of line. |  | COM_ImageCalibrationLineErr
22665 | CLN_ExtrapolateColumns | Error |  |  |  | CalibrationLine_Extrapolate: Invalid number of new columns. |  | COM_ImageCalibrationLineErr
22666 | CLN_MultiplyInvalid | Error |  |  |  | CalibrationLine_Multiply: Empty calibration line. |  | COM_ImageCalibrationLineErr
22667 | CLN_MultiplyLinesInvalid | Error |  |  |  | CalibrationLine_Multiply: Empty calibration line. |  | COM_ImageCalibrationLineErr
22668 | CLN_MultiplyLinesInvalid2 | Error |  |  |  | CalibrationLine_Multiply: Empty calibration line. |  | COM_ImageCalibrationLineErr
22669 | CLN_MultiplyLinesBitsPP | Error |  |  |  | CalibrationLine_Multiply: PixelBits do not match. |  | COM_ImageCalibrationLineErr
22670 | CLN_MultiplyLinesNCalLines | Error |  |  |  | CalibrationLine_Multiply: NumberOfCalLines do not match. |  | COM_ImageCalibrationLineErr
22671 | CLN_DivideInvalid | Error |  |  |  | CalibrationLine_Divide: Empty calibration line. |  | COM_ImageCalibrationLineErr
22672 | CLN_DivideZero | Error |  |  |  | CalibrationLine_Divide: Division by zero. |  | COM_ImageCalibrationLineErr
22673 | CLN_SubtractInvalid | Error |  |  |  | CalibrationLine_Subtract: Empty minuend line. |  | COM_ImageCalibrationLineErr
22674 | CLN_SubtractInvalid2 | Error |  |  |  | CalibrationLine_Subtract: Empty subtrahend line. |  | COM_ImageCalibrationLineErr
22675 | CLN_ExLineInvalid | Error |  |  |  | CalibrationLine_ExtrapolateLine: Line to be extrapolated is NULL! |  | COM_ImageCalibrationLineErr
22676 | CLN_ExLineStart | Error |  |  |  | CalibrationLine_ExtrapolateLine: startIndex not valid for line. |  | COM_ImageCalibrationLineErr
22677 | CLN_ExLineLength | Error |  |  |  | CalibrationLine_ExtrapolateLine: lineLength not valid for line. |  | COM_ImageCalibrationLineErr
22678 | CLN_ExLineFirstvalid | Error |  |  |  | CalibrationLine_ExtrapolateLine: firstValid is out of range. |  | COM_ImageCalibrationLineErr
22679 | CLN_ExLineLastValid | Error |  |  |  | CalibrationLine_ExtrapolateLine: lastValid is out of range. |  | COM_ImageCalibrationLineErr
22680 | CLN_MinMaxScanValInvalidLine | Error |  |  |  | CalibrationLine_CalculateMinMaxScanVal: Empty calibration line. |  | COM_ImageCalibrationLineErr
22681 | CLN_SplineDivByZero | Error |  |  |  | Bad input to routine splint leads to division by 0. |  | COM_ImageCalibrationLineErr
22682 | CLN_SetCalLinesLength | Error |  |  |  | Could not set CalLines due to wrong length. |  | COM_ImageCalibrationLineErr
22683 | CLN_SetCalLinesMalloc | Error |  |  |  | CalibrationLine_set_CalLines failed. malloc error. |  | COM_ImageCalibrationLineErr
22684 | CLN_ConsistencyInvalid | Error |  |  |  | CalibrationLine_CheckConsistency: Empty calibration line. |  | COM_ImageCalibrationLineErr
22685 | CLN_ConsistencyNotNormalized | Error |  |  |  | CalibrationLine_CheckConsistency: Gain line is not normalized. |  | COM_ImageCalibrationLineErr
22686 | CLN_ConsistencyRange | Error |  |  |  | CalibrationLine_CheckConsistency: Gain line contains too low pixel values. |  | COM_ImageCalibrationLineErr
22687 | CLN_Checksum | Error |  |  |  | CalibrationLine checksum error. | Meaning: There was a checksum error in a read calibration line. Reason: Transmission problem from photomultiplier module (PMM) to CPU board (PMI). Cure: The software automatically retries to read the calibration line. If the problem persists, perform a shading calibration. | COM_ImageCalibrationLineErr
22688 | CLN_DenormalizeInvalid | Error |  |  |  | CalibrationLine_Denormalize: Empty calibration line. |  | COM_ImageCalibrationLineErr
22689 | CLN_DenormalizeNotNormalized | Error |  |  |  | CalibrationLine_Denormalize: Calibration line is not normalized. |  | COM_ImageCalibrationLineErr
22708 | MED_USNoData | Error |  |  |  | SlidingMedian_UShort: No data given to calculate median. |  | COM_SWError
22709 | MED_USRangeTooSmall | Error |  |  |  | SlidingMedian_UShort: The median range is too small. |  | COM_SWError
22710 | MED_USRangeTooBig | Error |  |  |  | SlidingMedian_UShort: The median range is bigger than the data length. |  | COM_SWError
22711 | MED_FNoData | Error |  |  |  | SlidingMedian_Float: No data given to calculate median. |  | COM_SWError
22712 | MED_FRangeTooSmall | Error |  |  |  | SlidingMedian_Float: The median range is too small. |  | COM_SWError
22713 | MED_FRangeTooBig | Error |  |  |  | SlidingMedian_Float: The median range is bigger than the data length. |  | COM_SWError
22714 | ART_DetNoData | Error |  |  |  | LineArtefact_Detect: No image data given. |  | COM_SWError
22715 | ART_DetAreaSize | Error |  |  |  | LineArtefact_Detect: Columns or line is less or equal 0 |  | COM_SWError
22716 | ART_CorrNoData | Error |  |  |  | LineArtefact_Correct: No image data given. |  | COM_SWError
22717 | ART_CorrAreaSize | Error |  |  |  | LineArtefact_Correct: Columns or line is less or equal 0 |  | COM_SWError
22718 | RTC_QFull | Error |  |  |  | Retrace correction block queue is full. Software error. |  | COM_SWError
22719 | RTC_NotSuccessiveBlocks | Error |  |  |  | Retrace correction got not successive blocks. Software error. |  | COM_SWError
22720 | RTC_BlockTooSmall | Error |  |  |  | First image block for retrace correction is too small. Software error. |  | COM_SWError
22721 | RTC_NoInLine1 | Error |  |  |  | Retrace correction did not find correct input line. Software error. |  | COM_SWError
22722 | RTC_NoInLine2 | Error |  |  |  | Retrace correction did not find correct input line. Software error. |  | COM_SWError
22723 | RTC_NoInLine3 | Error |  |  |  | Retrace correction did not find correct input line. Software error. |  | COM_SWError
22724 | RTC_NoOutLine | Error |  |  |  | Retrace correction did not find correct output line. Software error. |  | COM_SWError
22728 | IMA_AddLines | Error |  |  |  | Could not add lines to ImageArea. |  | COM_SWError
22729 | GRA_Allocate | Error |  |  |  | Could not allocate graph image. |  | COM_SWError
22730 | GRA_Rescale | Error |  |  |  | Could not rescale graph image. |  | COM_SWError
22731 | GRA_Save | Error |  |  |  | Could not save graph image. |  | COM_SWError
22748 | SYN_CRFSNotMounted | Error |  |  |  | SyncCalLines: CRFileSystem not mounted. |  | COM_SWError
22749 | SYN_CRFSDir | Error |  |  |  | SyncCalLines: Could not get CRFileSystem directory. |  | COM_SWError
22750 | SYN_CRFSOpen | Error |  |  |  | SyncCalLines: Could not open calibration file from CRFileSystem. |  | COM_SWError
22751 | SYN_CRFSNoData | Error |  |  |  | SyncCalLines: Calibration file on CRFileSystem contains no data. |  | COM_SWError
22752 | SYN_Malloc | Error |  |  |  | SyncCalLines: Could not allocate buffer for CRFile. |  | COM_SWError
22753 | SYN_FOpen | Error |  |  |  | SyncCalLines: Could not open calibration file for writing. |  | COM_SWError
22754 | SYN_Dir | Error |  |  |  | SyncCalLines: Could not get directory. |  | COM_SWError
22755 | SYN_CreateCL | Error |  |  |  | SyncCalLines: Could not create CalibrationLine from file. |  | COM_SWError
22756 | SYN_CRFSWrite | Error |  |  |  | SyncCalLines: Could not write CalibrationLine to CRFileSystem. |  | COM_SWError
22757 | SYN_ExchangedAcq | Error | 159 | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. |  | SyncCalLines: Exchanged PMM detected. | Meaning: Detected an exchanged photomultiplier module. Shading calibration not good enough any more. Reason: Replacement by spare part. Cure: Perform a shading calibration. | COM_SWShadingCalibration
22758 | SYN_ExchangedStim | Error | 159 | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. |  | SyncCalLines: Exchanged optics module detected. | Meaning: Detected an exchanged optics module. Shading calibration not good enough any more. Reason: Replacement by spare part. Cure: Perform a shading calibration. | COM_SWShadingCalibration
22759 | SYN_AcqSerNo | Error | 159 | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. |  | SyncCalLines: Wrong PMM serial number. | Meaning: Detected wrong photomultiplier module serial number in calibration line. Shading calibration not valid. Reason: PMM serial number was changed or wrong calibration line was manually copied. Cure: Perform a shading calibration. | COM_SWShadingCalibration
22760 | SYN_GetLastWriteTime | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | SyncCalLines: Could not get the CalibrationLine write time (disk). |  | COM_SWError
22761 | SYN_CRFSSetLastWriteTime | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | SyncCalLines: Could not update the CalibrationLine write time in CRFileSystem. |  | COM_SWError
22762 | SYN_FWrite | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | SyncCalLines: Could not write CalibrationLine to disk. |  | COM_SWError
23040 | SetError | Error |  |  |  | Setting an error failed.
23041 | RemoveError | Error |  |  |  | Removing an error failed.
23042 | StartDiagnoseCycle | Info | 283 | Start diagnose cycle. | Confirm | Start Diagnose Cycle
23043 | StartEnduranceRunCycle | Info | 284 | Start endurance test cycle. | Confirm | Start Endurancerun Cycles
23044 | StartShadingCalibration | Info | 285 | Start shading calibration. | Confirm | Start shading clibration
23045 | MoveIPCarrierToMaintenance | Info | 296 | Device ready to move IP carrier to maintenance position. | Confirm | Device ready to move IP Carrier to maintenance position
23046 | MoveIPCarrierToPark | Info | 290 | Device ready to move IP carrier to park position | Confirm | Device ready to move IP Carrier to park position
23047 | READY4TRANSPORTLOCK | Info | 281 | Ready to fix transport lock. Switch off the digitizer and fix transport lock. |  | IP Carrier ready to attach the transport lock.
23048 | REMOVETRANSPORTLOCK | Info | 282 | Transport locks still attached. Remove transport locks. |  | Remove IP Carrier transport lock
23049 | StartStallCalibration | Info | 297 | Device ready to start stall calibration. | Confirm | Start Stall Calibration
23050 | StartCalBoardPosition | Info | 295 | Device ready to test Laser Beam Position. | Confirm | Start Cal. Board Position
23051 | NOT_IMPLEMENTED | Error |  |  |  | This function has no implementation
23552 | START_SERVER_FAILED | Error |  |  |  | Start Server failed | Starting of DirectID Server failed | COM_SWError
26112 | SW_VERSION_MISMATCH | Error |  |  |  | The software of the digitizer and the ErrorViewer on PC do not match. | Meaning: The software of the digitizer and the ErrorViewer on NX workstation do not match. The ErrorViewer is not able to display messages of the digitizer. Reason: 1. An old version of ErrorViewer is installed on NX workstation. Cure: Please install the ErrorViewer on NX wordkstation released for actual version of digitizer software. | COM_SWErrorViewerVersionMismatch
26113 | CREATE_SENDER_SOCKET | Error |  |  |  | Could not create sender socket. | Meaning: Internal software error connecting to NX workstation. Reason: TCP/IP communication problem with NX workstation Cure: Please restart the digitizer. If the problem occurs after restart again, please call service. | COM_NotNeeded_ForServiceId
26114 | CONNECT_FAILED | Error |  |  |  | Could not connect to NX workstation. | Meaning: Connection of ErrorViewer to NX failed Reason: 1. The NX workstation is not available in the network. 2. The NX workstation is not started or user is not logged in at NX workstation 3. The ErrorViewer on NX workstation is not started. 4. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 5. ErrorViewer version on NX workstation is not supported 6. The digitizer is configured for another NX workstation than the one actually used. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. 3. Start the ErrorViewer on NX workstation. 4. Exit the ErrorViewer and restart it. 5. Install an ErrorViewer with version 2.0 or higher and start the ErrorViewer on NX workstation. 5. Reconfigure NX workstation and digitizer. | COM_SWErrorViewerCommunicationProblem
26115 | TIMEOUT_SEND | Error |  |  |  | No response from NX workstation | Meaning:NX workstation does not respond. Reason: 1. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 2. ErrorViewer version on NX workstation is not supported Cure: 1. Exit the ErrorViewer and restart it. 2. Install an ErrorViewer with version 2.0 or higher | COM_SWErrorViewerCommunicationProblem
26116 | FATAL_ERROR | Error |  |  |  | Fatal error building message string. |  | COM_SWError
26117 | ACK_NOT_SYNC | Error |  |  |  | Acknoledge asynchron | Meaning: Communication problem between between digitizer and Erroviewer on NX wordkstation Reason: Communication between digitizer and Erroviewer on NX wordkstation is not synchron. Cure: Restart digitizer and ErrorViewer on NX workstation. If the problem occures again reinstall SW. | COM_SWErrorViewerCommunicationProblem
26118 | ERROR_SENDING_MESSAGE | Error |  |  |  | Error sending a message to NX workstation | Meaning: An error message could not be sent to ErrorViewer on NX wrokstation. Reason: 1. The NX workstation has actually been removed from the network. 2. The NX workstation has actually been shut dwon or user has logged out from NX workstation 3. The ErrorViewer has actually been shut down on NX workstation. 4. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 6. The digitizer configuration has actually been modified over WEB insterface , the digitizer now is configured for another NX workstation than the one actually used. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. 3. Start the ErrorViewer on NX workstation. 4. Exit the ErrorViewer and restart it. 5. Install an ErrorViewer with version 2.0 or higher and start the ErrorViewer on NX workstation. 5. Check configuration of NX workstation and digitizer. | COM_SWErrorViewerCommunicationProblem
26119 | PING_FAILED | Error |  |  |  | Could not connect to NX workstation. | Meaning: Ping of NX workstation failed Reason: 1. The NX workstation is not available in the network. 2. The NX workstation is not started. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. | COM_SWErrorViewerCommunicationProblem