## CR30SecondErrors
ErrorCode | Identifier | Priority | IdOfTextToDisplay | TextToDisplay | ConfirmRequired | Description | ExtendedDescription | ServID | ServCode | ServiceMessage
1024 | DTC_PERFTX_RDSYNC | Error |  |  |  | Error while waiting for sync frame | Timeout or error occured while initiating the transmission of identification data from the PMI Board to the TagReader component at the IP Handling Board. Either the CAN bus is defective or at least the TagReader component at the IP Handling Board (or the complete IP Handling Board) does not respond anymore.
1025 | DTC_PERFTX_BKLWRITE | Error |  |  |  | Error during block transfer of ipdata | Either the CAN bus is defective or at least the TagReader component at the IP Handling Board (or the complete IP Handling Board) does not respond anymore.
1026 | DTC_INITX_FCHOPEN | Error |  |  |  | Cannot open CAN channel for data transmission | Unable to open a channel to transmit the identification data from the PMI Board to the TagReader component at the IP Handling Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1027 | DTC_PERFRX_FCHOPEN | Error |  |  |  | Cannot open CAN channel for data retrieval | Unable to open a channel to transmit the identification data from the IP Handling Board to the PMI Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1028 | DTC_PERFRX_INITBLKRX | Error |  |  |  | Cannot initialize CAN channel for block retrieval | Timeout or error occured while initializing a block transfer over the CAN bus from the IP Handling Board to the PMI Board. This error is related to the PMI Board itself or to the software running on this PCB.
1029 | DTC_PERFRX_WRSYNC | Error |  |  |  | Error while writing sync frame | An error occured when trying to start the transmission of identification data from the IP Handling Board to the PMI Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1030 | DTC_PERFRX_BLKRX | Error |  |  |  | Error while performing the block read | An error occured when reading the identifiaction data that are sent from the IP Handling Board to the PMI Board. This error is related to one or more of the following reasons: - The CAN bus is defective. - The IP Handling Board aborted the transmission (e.g. software crash, hardware fault, ...) - The PMI Board itself reported an error condition in the CAN logic.
1031 | VHLDLL_INITUART | Error |  |  |  | DLL: Initialization of the UART failed | Initializing the serial line that connects the IP Handling Board with the Chip Reader failed. This is an internal error of the IP Handling Board and a string indication that this PCB is defective and should be replaced.
1032 | VHLDLL_RCF_RCT1 | Error |  |  |  | DLL: Timeout while reading a character from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1033 | VHLDLL_RCF_RCT2 | Error |  |  |  | DLL: Timeout while reading a character from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1034 | VHLDLL_SCF_RCT | Error |  |  |  | DLL: Timeout while waiting for Sync ACK from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1035 | VHLDLL_RA_RCT | Error |  |  |  | DLL: Timeout while waiting for Sync ACK from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1036 | VHLDLL_RX_UNKNOWN | Error |  |  |  | DLL: Unknown event in method receive() | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1037 | VHLDLL_TX_UNKNOWN | Error |  |  |  | DLL: Unknown event in method send() | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1038 | VHLDLL_STT_RXINITIAL_UNKNOWN | Error |  |  |  | DLL: Unknown event in initial transition in state OK::receiving | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1039 | VHLAL_POWERON_TX | Error |  |  |  | AL: Error sending command powerON() | Sending the command 'antenna on' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1040 | VHLAL_POWERON_RX | Error |  |  |  | AL: Error receiving answer for powerON() | Receiving the answer to the command 'antenna on' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1041 | VHLAL_POWERON_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to powerON() | The answer to the command 'antenna on' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1042 | VHLAL_POWERON_SEQID | Error |  |  |  | AL: Sequence error in answer for powerON() | The answer to the command 'antenna on' from the Chip Reader carries a wrong sequence ID.
1043 | VHLAL_SELECT_TX | Error |  |  |  | AL: Error sending command SELECT() | Sending the command 'select tag' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1044 | VHLAL_SELECT_RX | Error |  |  |  | AL: Error receiving answer for SELECT() | Receiving the answer to the command 'select tag' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1045 | VHLAL_SELECT_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to SELECT() | The answer to the command 'select tag' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1046 | VHLAL_SELECT_SEQID | Error |  |  |  | AL: Sequence error in answer for SELECT() | The answer to the command 'select tag' from the Chip Reader carries a wrong sequence ID.
1047 | VHLAL_GETINFO_TX | Error |  |  |  | AL: Error sending command GETINFO() | Sending the command 'get info' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1048 | VHLAL_GETINFO_RX | Error |  |  |  | AL: Error receiving answer for GETINFO() | Receiving the answer to the command 'get info' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1049 | VHLAL_GETINFO_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to GETINFO() | The answer to the command 'get info' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1050 | VHLAL_GETINFO_SEQID | Error |  |  |  | AL: Sequence error in answer for GETINFO() | The answer to the command 'get info' from the Chip Reader carries a wrong sequence ID.
1051 | VHLAL_READ_TX | Error |  |  |  | AL: Error sending command READ() | Sending the command 'read' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1052 | VHLAL_READ_RX | Error |  |  |  | AL: Error receiving answer for READ() | Receiving the answer to the command 'read' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1053 | VHLAL_READ_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to READ() | The answer to the command 'read' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1054 | VHLAL_READ_SEQID | Error |  |  |  | AL: Sequence error in answer for READ() | The answer to the command 'read' from the Chip Reader carries a wrong sequence ID.
1055 | VHLAL_WRITE_TX | Error |  |  |  | AL: Error sending command WRITE() | Sending the command 'write' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1056 | VHLAL_WRITE_RX | Error |  |  |  | AL: Error receiving answer for WRITE() | Receiving the answer to the command 'write' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1057 | VHLAL_WRITE_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to WRITE() | The answer to the command 'write' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1058 | VHLAL_WRITE_SEQID | Error |  |  |  | AL: Sequence error in answer for WRITE() | The answer to the command 'write' from the Chip Reader carries a wrong sequence ID.
1059 | VHLAL_POWEROFF_TX | Error |  |  |  | AL: Error sending command powerOFF() | Sending the command 'antenna off' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1060 | VHLAL_POWEROFF_RX | Error |  |  |  | AL: Error receiving answer for powerOFF() | Receiving the answer to the command 'antenna off' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1061 | VHLAL_POWEROFF_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to powerOFF() | The answer to the command 'antenna off' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1062 | VHLAL_POWEROFF_SEQID | Error |  |  |  | AL: Sequence error in answer for powerOFF() | The answer to the command 'antenna off' from the Chip Reader carries a wrong sequence ID.
1063 | VHLAL_SETMIF_TX | Error |  |  |  | AL: Error sending command setupMIFARE() | Sending the command 'select Mifare access mode' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1064 | VHLAL_SETMIF_RX | Error |  |  |  | AL: Error receiving answer for setupMIFARE() | Receiving the answer to the command 'select Mifare access mode' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1065 | VHLAL_SETMIF_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to setupMIFARE() | The answer to the command 'select Mifare access mode' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1066 | VHLAL_SETMIF_SEQID | Error |  |  |  | AL: Sequence error in answer for setupMIFARE() | The answer to the command 'select Mifare access mode' from the Chip Reader carries a wrong sequence ID.
1067 | VHLAL_GETSNR_TX | Error |  |  |  | AL: Error sending command getSNR() | Sending the command 'getSerialNumberOfTag' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1068 | VHLAL_GETSNR_RX | Error |  |  |  | AL: Error receiving answer for getSNR() | Receiving the answer to the command 'getSerialNumberOfTag' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1069 | VHLAL_GETSNR_CHECKSUM | Error |  |  |  | AL: Checksum error in answer to getSNR() | The answer to the command 'getSerialNumberOfTag' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1070 | VHLAL_GETSNR_SEQID | Error |  |  |  | AL: Sequence error in answer for getSNR() | The answer to the command 'getSerialNumberOfTag' from the Chip Reader carries a wrong sequence ID.
1071 | APP_TEST_ALINIT | Error |  |  |  | APP: Initializing the application layer failed | Cannot initialize the communication layer between the IP Handling Board and the Chip Reader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1072 | APP_TEST_GETINFO | Error |  |  |  | APP: Cannot retrieve firmware information | Cannot execute command 'getFirmwareVersion' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1073 | APP_GETINFO | Error |  |  |  | APP: Cannot retrieve firmware information | Cannot execute command 'getFirmwareVersion' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1074 | APP_TESTTAG_PWRON | Error |  |  |  | APP: Cannot switch on RF-power in testTAG() | Cannot execute command 'testForTag/AntennaOn' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1075 | APP_TESTTAG_PWROFF | Error |  |  |  | APP: Cannot switch off RF-power in testTAG() | Cannot execute command 'testForTag/AntennaOff' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1076 | APP_ENABLE_SETMIF | Error |  |  |  | APP: Cannot setupMifare access mode in method ENABLE | Cannot execute command 'Enable/setMifareAccMode' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1077 | APP_ENABLE_GETSNR | Error |  |  |  | APP: Cannot retrieve serial number of the tag in method ENABLE | Cannot execute command 'Enable/GetSerialNumberOfTag' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1078 | APP_ENABLE_SELECT | Error |  |  |  | APP: Cannot select tag in method ENABLE | Cannot execute command 'Enable/SelectTag' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1079 | APP_ENABLE_PWRON | Error | 117 | Error during writing on cassette chip after successful scan. Erase the image plate again. |  | APP: Cannot switch RF-power on in method ENABLE | Cannot execute command 'Enable/AntennaOn' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1536 | INFO_VERS_TERM | Error |  |  |  | IPInfo::Version: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::Version attribute read. | COM_SWCorruptIDData
1537 | INFO_VERS_CORRUPTED | Error |  |  |  | IPInfo::Version: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Version. | COM_SWCorruptIDData
1538 | INFO_IDNUM_TERM | Error |  |  |  | IPInfo::IDNumber: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IDNumber attribute read. | COM_SWCorruptIDData
1539 | INFO_IDNUM_CORRUPTED | Error |  |  |  | IPInfo::IDNumber: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IDNumber. | COM_SWCorruptIDData
1540 | INFO_INITDATE_TERM | Error |  |  |  | IPInfo::InitDate: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::InitDate attribute read. | COM_SWCorruptIDData
1541 | INFO_INITDATE_CORRUPTED | Error |  |  |  | IPInfo::InitDate: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::InitDate. | COM_SWCorruptIDData
1542 | INFO_IPSIZE_TERM | Error |  |  |  | IPInfo::IPSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSize attribute read. | COM_SWCorruptIDData
1543 | INFO_IPSIZE_CORRUPTED | Error |  |  |  | IPInfo::IPSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSize. | COM_SWCorruptIDData
1544 | INFO_SCANSIZE_TERM | Error |  |  |  | IPInfo::ScanSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::ScanSize attribute read. | COM_SWCorruptIDData
1545 | INFO_SCANSIZE_CORRUPTED | Error |  |  |  | IPInfo::ScanSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::ScanSize. | COM_SWCorruptIDData
1546 | INFO_IPTYPE_TERM | Error |  |  |  | IPInfo::IPType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPType attribute read. | COM_SWCorruptIDData
1547 | INFO_IPTYPE_CORRUPTED | Error |  |  |  | IPInfo::IPType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPType. | COM_SWCorruptIDData
1548 | INFO_IPSENS_TERM | Error |  |  |  | IPInfo::IPSensitivity: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSensitivity attribute read. | COM_SWCorruptIDData
1549 | INFO_IPSENS_CORRUPTED | Error |  |  |  | IPInfo::IPSensitivity: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSensitivity. | COM_SWCorruptIDData
1550 | INFO_IPERA_TERM | Error |  |  |  | IPInfo::IPErasibility: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPErasibility attribute read. | COM_SWCorruptIDData
1551 | INFO_IPERA_CORRUPTED | Error |  |  |  | IPInfo::IPErasibility: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPErasibility. | COM_SWCorruptIDData
1552 | INFO_IPTHICK_TERM | Error |  |  |  | IPInfo::IPThickness: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPThickness attribute read. | COM_SWCorruptIDData
1553 | INFO_IPTHICK_CORRUPTED | Error |  |  |  | IPInfo::IPThickness: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPThickness. | COM_SWCorruptIDData
1554 | INFO_CHECKSUM_TERM | Error |  |  |  | IPInfo::Checksum: unexpected attribute terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: No section terminator after IPInfo::Checksum attribute read. | COM_SWCorruptIDData
1555 | INFO_CHECKSUM_CORRUPTED | Error |  |  |  | IPInfo::Checksum: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Checksum. | COM_SWCorruptIDData
1556 | INFO_CHECKSUM_WRONG | Error |  |  |  | IPInfo::Checksum: checksum error | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Checksum error for section IPInfo. | COM_SWCorruptIDData
1557 | STATUS_VERS_TERM | Error |  |  |  | IPStatus::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Version attribute read. | COM_SWCorruptIDData
1558 | STATUS_VERS_NOTSUPPORTED | Error |  |  |  | IPStatus::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPStatus::Version not supported. | COM_SWCorruptIDData
1559 | STATUS_VERS_CORRUPTED | Error |  |  |  | IPStatus::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Version. | COM_SWCorruptIDData
1560 | STATUS_STATUS_TERM | Error |  |  |  | IPStatus::Status: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Status attribute read. | COM_SWCorruptIDData
1561 | STATUS_STATUS_CORRUPTED | Error |  |  |  | IPStatus::Status: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Status. | COM_SWCorruptIDData
1562 | STATUS_CYCLES_TERM | Error |  |  |  | IPStatus::Cycles: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Cycles attribute read. | COM_SWCorruptIDData
1563 | STATUS_CYCLES_CORRUPTED | Error |  |  |  | IPStatus::Cycles: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Cycles. | COM_SWCorruptIDData
1564 | STATUS_CHECKSUM_TERM | Error |  |  |  | IPStatus::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPStatus::Checksum attribute read. | COM_SWCorruptIDData
1565 | STATUS_CHECKSUM_CORRUPTED | Error |  |  |  | IPStatus::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Checksum. | COM_SWCorruptIDData
1566 | STATUS_CHECKSUM_WRONG | Error |  |  |  | IPStatus::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPStatus. | COM_SWCorruptIDData
1567 | CTRL_VERS_TERM | Error |  |  |  | IPCtrl::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Version attribute read. | COM_SWCorruptIDData
1568 | CTRL_VERS_NOTSUPPORTED | Error |  |  |  | IPCtrl::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPCtrl::Version not supported. | COM_SWCorruptIDData
1569 | CTRL_VERS_CORRUPTED | Error |  |  |  | IPCtrl::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Version. | COM_SWCorruptIDData
1570 | CTRL_SENSITIVITY_TERM | Error |  |  |  | IPCtrl::Sensitivity: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Sensitivity attribute read. | COM_SWCorruptIDData
1571 | CTRL_SENSITIVITY_CORRUPTED | Error |  |  |  | IPCtrl::Sensitivity: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Sensitivity. | COM_SWCorruptIDData
1572 | CTRL_EXPDOSE_TERM | Error |  |  |  | IPCtrl::ExposureDose: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::ExposureDose attribute read. | COM_SWCorruptIDData
1573 | CTRL_EXPDOSE_CORRUPTED | Error |  |  |  | IPCtrl::ExposureDose: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::ExposureDose. | COM_SWCorruptIDData
1574 | CTRL_OPCODE_TERM | Error |  |  |  | IPCtrl::OperationCode: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::OpCode attribute read. | COM_SWCorruptIDData
1575 | CTRL_OPCODE_CORRUPTED | Error |  |  |  | IPCtrl::CoperationCode: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::OpCode. | COM_SWCorruptIDData
1576 | CTRL_CHECKSUM_TERM | Error |  |  |  | IPCtrl::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPCtrl::Checksum attribute read. | COM_SWCorruptIDData
1577 | CTRL_CHECKSUM_CORRUPTED | Error |  |  |  | IPCtrl::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Checksum. | COM_SWCorruptIDData
1578 | CTRL_CHECKSUM_WRONG | Error |  |  |  | IPCtrl::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPCtrl. | COM_SWCorruptIDData
1579 | IPPSDATA_VERS_TERM | Error |  |  |  | IPPSData::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::Version attribute read. | COM_SWCorruptIDData
1580 | IPPSDATA_VERS_NOTSUPPORTED | Error |  |  |  | IPPSData::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPPSData::Version not supported. | COM_SWCorruptIDData
1581 | IPPSDATA_VERS_CORRUPTED | Error |  |  |  | IPPSData::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::Version. | COM_SWCorruptIDData
1582 | IPPSDATA_MPPSINSTUID_TERM | Error |  |  |  | IPPSData::MPPSInstanceUID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::MPPSInstanceUID attribute read. | COM_SWCorruptIDData
1583 | IPPSDATA_MPPSINSTUID_CORRUPTED | Error |  |  |  | IPPSData::MPPSInstanceUID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::MPPSInstanceUID. | COM_SWCorruptIDData
1584 | IPPSDATA_SOPINSTUID_TERM | Error |  |  |  | IPPSData::SOPInstanceUID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::SOPInstanceUID attribute read. | COM_SWCorruptIDData
1585 | IPPSDATA_SOPINSTUID_CORRUPTED | Error |  |  |  | IPPSData::SOPInstanceUID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::SOPInstanceUID. | COM_SWCorruptIDData
1586 | IPPSDATA_PROCSTAT_TERM | Error |  |  |  | IPPSData::ProcessingStation: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ProcessingStation attribute read. | COM_SWCorruptIDData
1587 | IPPSDATA_PROCSTAT_CORRUPTED | Error |  |  |  | IPPSData::ProcessingStation: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ProcessingStation. | COM_SWCorruptIDData
1588 | IPPSDATA_STIMIDENTID_TERM | Error |  |  |  | IPPSData::StimIdentID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::StimIdentID attribute read. | COM_SWCorruptIDData
1589 | IPPSDATA_STIMIDENTID_CORRUPTED | Error |  |  |  | IPPSData::StimIdentID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::StimIdentID. | COM_SWCorruptIDData
1590 | IPPSDATA_EXPROOMID_TERM | Error |  |  |  | IPPSData::ExposureRoomID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ExposureRoom attribute read. | COM_SWCorruptIDData
1591 | IPPSDATA_EXPROOMID_CORRUPTED | Error |  |  |  | IPPSData::ExposureRoomID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ExposureRoom. | COM_SWCorruptIDData
1592 | IPPSDATA_CHECKSUM_TERM | Error |  |  |  | IPPSData::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPPSData::Checksum attribute read. | COM_SWCorruptIDData
1593 | IPPSDATA_CHECKSUM_CORRUPTED | Error |  |  |  | IPPSData::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::Checksum. | COM_SWCorruptIDData
1594 | IPPSDATA_CHECKSUM_WRONG | Error |  |  |  | IPPSData::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPPSData. | COM_SWCorruptIDData
1595 | INFO_VERSION_NOTSUPPORTED | Error |  |  |  | IPInfo::Version: this version is not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPInfo::Version not supported. | COM_SWCorruptIDData
1596 | INFO_APPTYPE_TERM | Error |  |  |  | IPInfo::ApplicationType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::AppType attribute read. | COM_SWCorruptIDData
1597 | INFO_APPTYPE_CORRUPTED | Error |  |  |  | IPInfo::ApplicationType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::AppType. | COM_SWCorruptIDData
1598 | IPCTRL_PATNAME_TERM | Error |  |  |  | IPCtrl::PatientName: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::PatientName attribute read. | COM_SWCorruptIDData
1599 | IPCTRL_PATNAME_CORRUPTED | Error |  |  |  | IPCtrl::PatientName: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::PatientName. | COM_SWCorruptIDData
1600 | INFO_OPTION_TERM | Error |  |  |  | IPInfo::Option: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::Option attribute read. | COM_SWCorruptIDData
1601 | INFO_OPTION_CORRUPTED | Error |  |  |  | IPInfo::Option: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Option. | COM_SWCorruptIDData
4096 | IDENTIFICATION_TIMEOUT | Error | 132 | ID button not pressed within the timeout period. Image plate was not scanned. Cassette stays clamped until ID data are supplied by the operator. |  | user did not identify at NX in time | Meaning: The digitizer is waiting for the identification data and the timeout (60 sec.) exceeded. Reason: User did not identify at NX in time. Cure: Identify at the NX. If this is not working, check ethernet connection to the processing station and the network configuration.
4097 | WRONG_PRODUCT_TYPE | Error |  |  |  | the product type is wrong | Meaning: The product type read from the eLabels is not supported by this software. Reason: 1. This is not a 2nd generation CR30 digitizer or the wrong software version (for Genrad or Mammo) is installed. 2. A wrong product type is stored in the eLabels. Cure: 1. Check if this digitizer is really a 2nd generation CR30. The name plate must show the type 5175/2xx for CR30-X 2nd (Genrad) or 5179 for CR30-Xm (Mammo). 2. Check if the correct software version is installed: CIR_xxxx runs only on CR30-X 2nd (Genrad), CIRMxxxx runs only on CR30-Xm (Mammo). 3. Replace digitizer by a 2nd generation CR30.
4098 | IDENTIFIED_CASS_CANNOT_READ_DESTINATION | Error |  |  |  | cannot read destination name from data store | Meaning: The destination name cannot be read from the data store. Reason: An identified cassette was inserted and the digitizer tried to send the image to the configured destination, but the destination (= name of the processing station) cannot be read. Cure: Try to set the name of the destination again in the digitizer settings (on the NX via ADC.cpf or by hand in the service client) and reset the digitizer. Then insert the cassette again. If this is still not working, install new software and try again.
4099 | CONFIRM_READ_TAGDATA_FAILED | Error | 100 | Error during read-out of cassette chip data. Remove cassette. | Confirm | Reading of tag data failed | Meaning: Reading the data from the RF-tag failed. Reason: RF-tag on cassette tray is defective or chip reader is defective or not connected. Cure: Try it again with another cassette; check cable between HANDLING CONTROL BOARD and chip reader; check if chip reader is mounted correctly; if still NOK, replace chip reader, then HANDLING CONTROL BOARD.
4100 | FAILED_TO_FATAL_NO_ERROR_CODE | Error |  |  |  | DigitizerControl (DIC) received an error message with status 'FAILED' which leads to a fatal error. No specific error code was passed over. | Meaning: SW component DIC (DigitizerControl) received an error message which leads to a fatal digitizer error. No specific error code was passed over. Reason: The error code is missing resp. the error should be handled specifically. This is a SW error. Cure: Remove cassettes and IP from digitizer and restart digitizer to start the self test. If error persists, update SW.
4101 | IMAGE_QUEUE_FULL | Error | 302 | Image queue full. Please wait. |  | spooler input queue is full | Meaning: Image queue of software component 'spooler' is full. Reason:The spooler's capacity is exhausted, no image can be scanned at the moment. Cure: Wait until images are transmitted to the processing station. Be sure that the processing station is connected and working properly.
4102 | CONFIRM_REMOVE_CASSETTE_AGAIN_AT_START_UP | Error | 304 | Remove the cassette and press the confirm button. | Confirm | Start-Up: Cassette still detected after user was requested to remove cassette | Meaning: Remove the cassette, because the cassette is still detected. Reason: Cassette found during start-up. Reading the RF-tag failed. User was requested to remove cassette, but after 5 sec. the cassette detection sensor still detects the cassette. User is requested again to remove the cassette. Cure: Confirm message, remove cassette.
4103 | CONFIRM_TAG_ERR_OR_NO_TRAY_INSERT_CASS_18X24_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4104 | CONFIRM_TAG_ERR_OR_NO_TRAY_INSERT_CASS_15X30_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4105 | CONFIRM_CASSETTE_WITH_IP_FOUND_AT_START_UP | Error | 191 | Image plate during self-test in the digitizer detected. The image plate was moved back into the cassette. Please check whether the image is present at the PC and re-scan the cassette, if necessary. If the image is present, the cassette was not erased possibly. Erase the cassette again in any case. | Confirm | Start-Up: Cassette with IP found, no IP is in slowscan or transport unit | Meaning: Cassette with an IP found during start-up. Reason: Digitizer was switched off during cycle. Cure: Confirm message, remove cassette; check if image was sent to the NX; if not, scan cassette again; erase cassette again in any case.
4106 | CONFIRM_EMPTY_CASSETTE_FOUND_AT_START_UP | Error | 106 | Empty cassette found at start-up. Remove cassette. | Confirm | Start-Up: Cassette without an IP found, no IP is in slowscan or transport unit | Meaning: Cassette found during start-up, but no IP. Reason: Cassette without IP was inserted and digitizer was switched off. Cure: Confirm message, remove cassette; put a tray and an IP into the cassette.
4107 | IDENTIFICATION_FAILED | Error | 293 | Error occurred. Image plate is not scanned and erased! Restart digitizer and try again. If error remains, contact service to fix the problem. |  | The identification on the NX failed | Meaning: The identiication on the processing station NX failed. Reason: NX configuration error or software error. Cure: Check that NX is correctly configured for this digitizer; check configuration of examination studies; if everything is correctly configured, save the log files and inform both R&D teams of NX and digitizer
4108 | WRONG_IHC_PROTVER | Error |  |  |  | Protocol version of IHC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of IHC (IP-HandlingControl) is wrong. Cure: Install new software; inform R&D team.
4109 | WRONG_CHC_PROTVER | Error |  |  |  | Protocol version of CHC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of CHC (CassetteHandlingControl) is wrong. Cure: Install new software; inform R&D team.
4110 | WRONG_TRC_PROTVER | Error |  |  |  | Protocol version of TRC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of TRC (TagReaderControl) is wrong. Cure: Install new software; inform R&D team.
4111 | WRONG_ICC_PROTVER | Error |  |  |  | Protocol version of ICC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ICC (ImageCaptureControl) is wrong. Cure: Install new software; inform R&D team.
4112 | WRONG_ELC_PROTVER | Error |  |  |  | Protocol version of ELC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ELC (ElabelControl) is wrong. Cure: Install new software; inform R&D team.
4113 | WRONG_ICN_PROTVER | Error |  |  |  | Protocol version of ICN is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ICN (Infocounter component) is wrong. Cure: Install new software; inform R&D team.
4114 | WRONG_DID_PROTVER | Error |  |  |  | Protocol version of DID is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of DID (DirectID component) is wrong. Cure: Install new software; inform R&D team.
4115 | WRONG_IPC_PROTVER | Error |  |  |  | Protocol version of IPC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of IPC (ImageProcessing component) is wrong. Cure: Install new software; inform R&D team.
4116 | WRONG_SPI_PROTVER | Error |  |  |  | Protocol version ofSPI is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of SPI (SpoolerInput component) is wrong. Cure: Install new software; inform R&D team.
4117 | INVALID_BUFFER_FROM_SPI | Error |  |  |  | spooler returned invalid buffer | Meaning: Software component 'spooler' returned invalid buffer (= NULL). Reason: Fatal software error. Cure: Install new software; inform R&D team.
4118 | IDENTIFICATION_FAILED_BUSY | Error | 293 | Error occurred. Image plate is not scanned and erased! Restart digitizer and try again. If error remains, contact service to fix the problem. |  | The identification on the NX failed because a previous ID request is still busy | Meaning: The identification on the NX failed because a previous ID request is still busy and the digitizer tried several times to get the identification data from the NX. Reason: NX hang-up, configuration error or software error. Cure: 1. Restart NX and digitizer. 2. Check that NX is correctly configured for this digitizer; check configuration of examination studies; if everything is correctly configured, save the log files and inform both R&D teams of NX and digitizer
4119 | IHC_SELFTEST_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4120 | TRC_SELFTEST_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4121 | PREPARE_SCAN_WARNING_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4122 | SCAN_WARNING_ICC_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4123 | CONFIRM_WRONG_DATAFLOW_SPEC | Error | 108 | Wrong cassette type (wrong cassette chip initialization). Remove cassette. | Confirm | cassette is initialized with wrong data flow spec; spec version must be >= 1.3 | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong data flow specification (< version 1.3) Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
4124 | CONFIRM_WRONG_IP_TYPE | Error | 109 | Wrong cassette type (wrong image plate type). Remove cassette. | Confirm | wrong image plate type; must be a flexible image plate | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong setting of the image plate type; the image plate type must be set to 'flexible'. Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
4125 | CONFIRM_WRONG_APP_TYPE | Error | 110 | Wrong cassette type (wrong application). Remove cassette. | Confirm | wrong application type; must be GenRad HR | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong setting of the application type; the application type must be set to 'GenRad HR'. Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
4126 | CONFIRM_UNSUPPORTED_IP_SIZE | Error | 111 | Wrong cassette type (unsupported size). Remove cassette. | Confirm | unsupported IP size | Meaning: Cassette is rejected because IP size is not supported or cassette is initialized with unsupported size. Note: The cassette format 35x35 cm is not supported if the product serial number is smaller than 30000 for standard digitizer (subtype 100) resp. 40000 for white label digitizer (subtype 110). Reason: Cassette is rejected due to a wrong setting of the IP size or digitizer does not support 35x35 cm because digitizer is too old (see above). Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette. If this error occurs with format 35x35 cm check the product serial number and the subtype.
4127 | CONFIRM_WRONG_IP_SIZE | Error | 112 | Wrong cassette size initialization. Remove cassette. | Confirm | cassette initialized with wrong IP size | Meaning: Cassette is rejected because IP size is wrong or cassette is initialized with wrong size. Reason: Cassette is rejected due to a wrong setting of the IP size. Cure: Check if cassette is designated to be used in this digitizer; check initialization of the cassette.
4128 | CONFIRM_EMPTY_CASSETTE_AND_IP_FOUND_AT_START_UP | Error | 191 | Image plate during self-test in the digitizer detected. The image plate was moved back into the cassette. Please check whether the image is present at the PC and re-scan the cassette, if necessary. If the image is present, the cassette was not erased possibly. Erase the cassette again in any case. | Confirm | Start-Up: Cassette without IP found, IP is in slowscan or transport unit | Meaning: Cassette and IP was found in digitizer during start-up. Reason: Cassette was inserted and digitizer was switched off during the cycle. Cure: Confirm message, remove cassette; check if image was sent to the NX; if not, scan cassette again; erase cassette again in any case.
4129 | CONFIRM_ENABLE_TAGREADER_FAILED | Error | 114 | Cassette chip not readable or cassette tray is missing. Remove cassette. | Confirm | Enabling of tag reader failed | Meaning: Enabling the RF-tag reader failed. Reason: RF-tag on cassette tray is defective or chip reader is defective or not connected. Cure: Try it again with another cassette; check cable between HANDLING CONTROL BOARD and chip reader; check if chip reader is mounted correctly; if still NOK, replace chip reader, then HANDLING CONTROL BOARD.
4130 | WRONG_UID_PROTVER | Error |  |  |  | Protocol version of UID is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of UID (UserInterfaceDispatcher) is wrong. Cure: Install new software; inform R&D team.
4131 | TIMEOUT_CHC_ENABLE_CASS_DETECTION | Error |  |  |  | response timeout for CHC command enableCassDetection | Meaning: Communication timeout between PMI BOARD and HANDLING CONTROL BOARD. Reason: The SW component CHC (CassetteHandlingControl) on HANDLING CONTROL BOARD didn't reply to the command 'enableCassDetection' in time. Cure: Check CAN connection between PMII board and HANDLING CONTROL BOARD.
4132 | TIMEOUT_CHC_DISABLE_CASS_DETECTION | Error |  |  |  | response timeout for CHC command disableCassDetection | Meaning: Communication timeout between PMI BOARD and HANDLING CONTROL BOARD. Reason: The SW component CHC (CassetteHandlingControl) on HANDLING CONTROL BOARD didn't reply to the command 'disableCassDetection' in time. Cure: Check CAN connection between PMI BOARD and HANDLING CONTROL BOARD.
4133 | EXCHANGED_COMPONENT_DETECTED | Error | 305 | Exchanged component detected. Please confirm exchange via Service Menu. |  | exchanged component detected | Meaning: An exchanged component is detected and it cannot be determined automatically if only a board or a complete module was exchanged. Reason: A component was exchanged because a different serial number of a component is detected. For further operation the input from the service technician is needed. Cure: Confirm exchange via Service Menu (select board or module exchange).
4134 | PREPARE_SCAN_FAILED | Error | 293 | Error occurred. Image plate is not scanned and erased! Restart digitizer and try again. If error remains, contact service to fix the problem. |  | preparing the scan failed | Meaning: An error occurred before the scan actually starts. The digitizer terminated the scan cycle and is in a fatal error state. The image plate was not scanned and not erased. Reason: A scanning software component detected an error (e.g. incidence of light). Cure: Restart digitizer to execute the selftest and to release the cassette. Scan the cassette again. Try another cassette. If error persists, save logfiles.
4135 | FATAL_RELEASE_CASSETTE_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4136 | CONFIRM_CASSETTE_WITH_IP_AND_IP_FOUND_AT_START_UP_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4137 | NOBODY_WAITED_ON_ERROR_CONFIRMATION_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4138 | UNEXPECTED_SIGNAL_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4139 | FATAL_ERROR_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4140 | FATAL_INITIALIZATION_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4141 | FATAL_ERASE_MODE_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4142 | CMD_CURRENTLY_NOT_ALLOWED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4143 | SEND_FATAL_INVALID_PORT_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4144 | FATAL_CMD_ENABLE_FEEDER_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4145 | FATAL_CMD_DISABLE_FEEDER_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4146 | FATAL_CMD_EJECT_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4147 | FATAL_CMD_ERASE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4148 | FATAL_CMD_FEED_IP_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4149 | FATAL_CMD_RETURN_IP_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4150 | FATAL_CMD_SCAN_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4151 | FATAL_CMD_START_SELFTEST_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4152 | CONFIRM_UNKNOWN_FORMAT_DETECTED | Error | 116 | Unknown cassette format detected (cassette inserted correctly?). Remove cassette and insert it again. | Confirm | cassette fixer detected unknown format | Meaning: The MOTOR CASSETTE FIXING M1 detected an unknown cassette format. The cassette is still clamped and a user message is displayed to press confirm to release the cassette. The digitizer is still working, but if this error occurs frequently, it requires service support. Reason: The MOTOR CASSETTE FIXING M1 measured a cassette width which doesn't match with a known cassette format. The determined cassette width is too small or too big due to: - The cassette was inserted wrongly by the user (skewed, upside down or not with the flap first) - The CASSETTE FIXATION STEEL BAND is not moving smoothly - The CASSETTE DETECTION HOOK is not properly fixed or not the improved version - Load detection for MOTOR CASSETTE FIXING M1 is imperfect, STALL CALIBRATION was not successfully - Drive mechanism of MOTOR CASSETTE FIXING M1 is blocked or defect Cure: 1. Remove the cassette and insert it correctly again. 2. Check if cassette can be pushed smoothly into the right edge of the input slot and that the wedge of the cassette fixer presses correctly into the left notch of the cassette. 3. Insert another cassette of the same (preferred) or any other size. 4. Check if the CASSETTE FIXATION STEEL BAND and the cassette fixing motor is moving smoothly. 5. Check if the CASSETTE DETECTION HOOK is properly fixed. If applicable mount the improved CASSETTE DETECTION HOOK, refer to Service Bulletin No.35 6. Perform a calibration of the MOTOR CASSETTE FIXING M1, refer to the Service Manual 7. Replace the MOTOR CASSETTE FIXING M1, refer to the Service Manual for order and replacement 8. Replace the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement 9. Replace the complete CASSETTE FIXATION UNIT, refer to the Service Manual for order and replacement
4153 | TRC_INITIALIZATION_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4154 | CLEANING_UP_CHC_IHC_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4155 | INIT_CHC_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4156 | NX_NOT_AVAILABLE | Error | 307 | Image processing software is not available. Restart the image processing software. |  | NX is not available | Meaning: The NX application resp. the processing station is not available anymore. Reason: NX application stopped at the processing station or a network problem occurred. Cure:Check if NX application is running properly on the processing station, check ethernet connection to the processing station and the network configuration.
4157 | CONFIRM_WRITING_TAG_FAILED | Error | 117 | Error during writing on cassette chip after successful scan. Erase the image plate again. | Confirm | writing to RF-tag failed | Meaning: Writing data on RF-tag failed. Reason: RF-tag on cassette tray is defective or chip reader is defective or not connected. Cure: Try it again with another cassette; check cable between HANDLING CONTROL BOARD and chip reader; check if chip reader is mounted correctly; if still NOK, replace chip reader, then HANDLING CONTROL BOARD.
4158 | CONFIRM_ENABLE_TAGREADER_FOR_WRITING_FAILED | Error | 117 | Error during writing on cassette chip after successful scan. Erase the image plate again. | Confirm | enabling tag reader for writing to RF-tag failed | Meaning: Enabling the tag reader for writing data on RF-tag failed. Reason: RF-tag on cassette tray is defective or chip reader is defective or not connected. Cure: Try it again with another cassette; check cable between HANDLING CONTROL BOARD and chip reader; check if chip reader is mounted correctly; if still NOK, replace chip reader, then HANDLING CONTROL BOARD.
4159 | CONFIRM_ENABLE_TAGREADER_FAILED_ERASE_MODE | Error | 119 | Error during writing on cassette chip after successful erasure cycle. Erase the image plate again. | Confirm | enabling tag reader for writing to RF-tag failed during erasure cycle | Meaning: Enabling the tag reader for writing data on RF-tag failed during erasure cycle. Reason: RF-tag on cassette tray is defective or chip reader is defective or not connected. Cure: Try it again with another cassette; check cable between HANDLING CONTROL BOARD and chip reader; check if chip reader is mounted correctly; if still NOK, replace chip reader, then HANDLING CONTROL BOARD.
4160 | CONFIRM_WRITING_TAG_FAILED_ERASE_MODE | Error | 119 | Error during writing on cassette chip after successful erasure cycle. Erase the image plate again. | Confirm | writing to RF- tag failed during erasure cycle | Meaning: Writing data on RF-tag failed during erasure cycle. Reason: RF-tag on cassette tray is defective or chip reader is defective or not connected. Cure: Try it again with another cassette; check cable between HANDLING CONTROL BOARDd and chip reader; check if chip reader is mounted correctly; if still NOK, replace chip reader, then HANDLING CONTROL BOARD.
4161 | WARNING_READ_TAGDATA_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4162 | DESERIALIZATION_IPSTATUS_FAILED | Error |  |  |  | could not deserialize IPStatus | Meaning: Could not deserialize IPStatus data. Reason: Check underlying error. Cure: Check underlying error.
4163 | DESERIALIZATION_IPINFO_FAILED | Error |  |  |  | could not deserialize IPInfo | Meaning: Could not deserialize IPInfo data. Reason: Check underlying error. Cure: Check underlying error.
4164 | DESERIALIZATION_IPCCTRL_FAILED | Error |  |  |  | could not deserialize IPCtrl | Meaning: Could not deserialize IPCtrl data. Reason: Check underlying error. Cure: Check underlying error.
4165 | DESERIALIZATION_IPPSDATA_FAILED | Error |  |  |  | could not deserialize IPPSData | Meaning: Could not deserialize IPPSData. Reason: Check underlying error. Cure: Check underlying error.
4166 | IP_ERASED_ALREADY_DUE_TO_ERASE_TIMEOUT_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4167 | CAN_ADAPTER_ERROR | Error |  |  |  | fatal error in CAN adapter layer | Meaning: Fatal error in CAN adapter layer Reason: CAN adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help replace PMI BOARD.
4168 | LOGGING_ADAPTER_ERROR | Error |  |  |  | fatal error in logging adapter layer | Meaning: Fatal error in logging adapter layer Reason: Logging adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help replace PMI BOARD.
4169 | ERASE_BUTTON_PRESSED_DURING_SCAN | Error | 306 | To proceed, identify the cassette at the processing station. To erase the image plate, press the erase button before inserting the cassette. |  | erase button was pressed during scan cycle | Meaning: The erase button was pressed during the scan cycle. This is not allowed! Reason: The erase button was pressed too late: To start an erasure cycle, the erase button must be pressed before the cassette is inserted! Cure: Now identify on the processing station, then the scan cycle continues and the IP will be scanned and erased.
4170 | CAN_VCC_FAILURE | Error |  |  |  | CAN Vcc failure | Meaning: CAN voltage error. Reason: POWER INTEGRATION BOARD fuses broken, cable POWER INTEGRATION BOARD to PMI broken or POWER INTEGRATION BOARD defect; Cure: Check fuses on POWER INTEGRATION BOARD, replace POWER INTEGRATION BOARD cable to PMI BOARD, replace POWER INTEGRATION BOARD | PMI_CANVccErr | CIR214 | Supply voltage +12V (CAN-bus line driver) on PMI Board low or missing. Meaning: Test of the supply voltage (+12V) for the CAN-bus line driver failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411) 2) Contact problem (connector P5) or defect on CPMI Power Distribution Board (B411) 3) Contact problem (connector P1#17) or defect on CPMI Board (B410). Cure: 1) Check the supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411). If required replace the supply cable, contact your Regional Support Network for further assistance 2) Check connector P5 on CPMI Power Distribution Board (B411) 3) Check connector P1#17 on CPMI Board (B410) 4) Exchange the CPMI Power Distribution Board (B411), refer to the Service Manual for order and replacement 5) Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement
4171 | PMM_PLUS20V_FAILURE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4172 | PMM_MINUS10V_FAILURE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4173 | ANALOG_VCC_FAILURE | Error |  |  |  | Analog Vcc failure | Meaning: The check of the +8 V supply voltage for A/D conversion on the PMI BOARD failed. The Digitizer is out of operation and requires service support. Reason: Check of the +8V supply voltage failed due to: - Defective flat cable between the POWER INTEGRATION BOARD and the PMI BOARD, broken wires or bad contacts - Defective POWER INTEGRATION BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between the POWER INTEGRATION BOARD and the PMI BOARD 3. Exchange the POWER INTEGRATION BOARD | PMI_AnalogVccErr | CIR213 | Supply voltage +8V on PMI Board low or missing. Meaning: Test of the supply voltage (+8V) for the A/D conversion on the CPMI Board failed. The Digitizer is out of operation and requires service support.. Reason: 1) Contact problem or defect at supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411) 2) Contact problem (connector P5) or defect on CPMI Power Distribution Board (B411) 3) Contact problem (connector P1#17) or defect on CPMI Board (B410). Cure: 1) Check the supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411). If required replace the supply cable, contact your Regional Support Network for further assistance 2) Check connector P5 on CPMI Power Distribution Board (B411) 3) Check connector P1#17 on CPMI Board (B410) 4) Exchange the CPMI Power Distribution Board (B411), refer to the Service Manual for order and replacement 5) Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement.
4174 | OPTICS_VCC_FAILURE | Error |  |  |  | optics +12V or -12V failure | Meaning: The check of the supply voltage for the optic module failed. The Digitizer is out of operation and requires service support. Reason: Check of the supply voltage failed due to: - Defective flat cable between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD, broken wires or bad contacts - Defective PMI POWER DISTRIBUTION BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD 3. Exchange the PMI POWER DISTRIBUTION BOARD | PMI_OpticsPwrErr | CIR215 | Supply voltage +12V / -12V (Optic module) on PMI Board low or missing. Meaning: Test of the supply voltage (+12V / -12V) for the optic module failed. The Digitizer is out of operation and requires service support.. Reason: 1) Contact problem or defect at supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411) 2) Contact problem (connector P5) or defect on CPMI Power Distribution Board (B411) 3) Contact problem (connector P1#17) or defect on CPMI Board (B410). Cure: 1) Check the supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411). If required replace the supply cable, contact your Regional Support Network for further assistance 2) Check connector P5 on CPMI Power Distribution Board (B411) 3) Check connector P1#17 on CPMI Board (B410) 4) Exchange the CPMI Power Distribution Board (B411), refer to the Service Manual for order and replacement 5) Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement.
4175 | OPTICS_MINUS12V_FAILURE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4176 | DISPLAY_VCC_FAILURE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4177 | WRONG_RAMP_PIXEL_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4178 | INIT_BOARD_FAILED_OBSOLETE2 | Error |  |  |  | OBSOLETE | OBSOLETE
4179 | CHC_COMMUNICATION_FAILED | Error |  |  |  | no communication possible with CHC on HANDLING CONTROL BOARD | Meaning: Communication error between PMI BOARD and HANDLING CONTROL BOARD. Cure: Check CAN connection
4180 | FATAL_IHC_NO_RECEIVER_FOUND_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4181 | FATAL_IHC_UNKNOWN_CHANNEL_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4182 | FATAL_IHC_ERR_READING_ACKNOWLEDGE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4183 | FATAL_IHC_CAN_ERROR_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4196 | FEEDER_DISABLED_REJECT_CASSETTE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4197 | FATAL_REJECT_CASSETTE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4198 | DIGITIZER_START_UP | Info | 123 | Digitizer start-up. |  | digitizer is starting-up | Meaning: Digitizer is starting-up or not yet enabled to process inserted cassettes after start-up. Reason: Digitizer is starting-up or not yet enabled by the processing station to take cassettes. Cure: Check if digitizer is perfoming the start-up (you can motor movements). If you hear movements, wait until they are finished. If you hear nothing, restart the digitizer; if this doesn't help or if message doesn't disappear a few minutes after end of digitizer's start-up, reboot the PC.
4199 | TIMEOUT_SCAN_END_ICC | Error | 129 | Problem during scan. The image may not have the required quality. |  | timeout for 'scan' response from ICC after slowscan stopped | Meaning: The scanning unit didn't finish the scan process although the slowscan motor has already stopped. Reason: ICC (ImageCaptureControl) is still waiting for end of the scan, i.e. the required number of pixels have not been stored in the SDRAM of the FPGA. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI BOARD and PMT/lens assembly; if OK, replace PMI; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly.
4200 | OPCODE_NOT_SUPPORTED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4201 | CONFIRM_CLEAN_UP_CASS_CLOSED_ENABLE_TAGREADER_FAILED | Error | 125 | Cassette chip not readable or cassette tray is missing. Remove cassette, if present. If no cassette is present in the input slot, please contact service to fix the problem. | Confirm | Enabling of tag reader failed during clean-up after closing the cassette. | Meaning: Enabling of tag reader failed during start-up after closing the cassette. Reason: RF-tag not found. Cure: If a cassette is present in the input slot during start-up, check if cassette contains a tray with a RF-tag. If no cassette is present in the input slot during start-up, check if cassette detection sensor (edge sensor, microswitch on cassette holder) is connected and working properly.
4202 | CONFIRM_IP_WITHOUT_CASSETTE_FOUND_AT_START_UP | Error | 126 | Start-up: An image plate was found in the digitizer, but no cassette. The image plate will be moved to the cassette slot as soon as you confirm. Please confirm and withdraw the image plate. | Confirm | Start-up: an IP was found in the Digitizer, but no cassette. The IP will be transported to the cassette slot as soon as you confirm. Please confirm and withdraw the IP. | Meaning: An IP was found in the Digitizer, but no cassette during start-up. Reason: IP-jam occured before or manipulation by hand. Cure:The IP will be transported to the cassette slot as soon as you confirm. Please confirm and withdraw the IP.
4203 | FATAL_CMD_CLEAN_UP_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4204 | CLEAN_UP_IN_ERASURE_MODE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4205 | CONFIRM_CLEAN_UP_BEFORE_SCANNING_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4206 | CLEAR_IMAGE_TX_ACTIVE_DURING_SCANNING_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4207 | INIT_CHC_OPENER_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4208 | CONFIRM_REMOVE_CASS_FROM_SLOT | Error | 255 | Cassette slot is blocked by something. Check cassette slot and remove object. | Confirm | initialization of CHC's opener drive failed; remove cassette out of input slot | Meaning: Initialization of CHC's (CassetteHandlingControl) opener drive failed. Reason: Cassette slot is blocked by something. Cure: Remove cassette or any other object out of input slot and confirm message. Then the opener drive will be tested again and a dedicated error message is displayed if initialization fails again.
4209 | FATAL_ERROR_ASAP_DISPLAYS_ERROR_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4210 | CASSETTE_DETECTION_FAILURE | Error |  |  |  | cassette is still detected in the input slot at start-up after user confirmed that he removed the cassette (checked 5 s after cassette is released) | Meaning: Failure of the cassette detection or user didn't remove cassette without a tray from the input slot at start-up. Reason: Cassette is still detected in the input slot at start-up after user confirmed that he removed the cassette (checked 5 s after cassette is released). Cure: Check that cassette has been removed from the input slot; if no cassette was in the input slot, check if microswitch on cassette holder is connected and working properly
4211 | AUTO_CLEAN_UP_INIT_CHC_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4212 | AUTO_CLEAN_UP_INIT_IHC_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4213 | AUTO_CLEAN_UP_NO_IP_FOUND | Error | 136 | Open digitizer to check for image plate jam or check for an empty cassette and restart the digitizer. |  | automatic clean-up failed because IP was not found. | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, because IHC (IP-HandlingControl) did not find an IP. Cure: Open digitizer to check for image plate jam or check for an empty cassette and restart the digitizer.
4214 | AUTO_CLEAN_UP_NO_CASS_FOUND | Error |  |  |  | automatic clean-up failed because cassette was not found. | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, because CHC (CassetteHandlingControl) did not find the cassette, which must be present in this situation. Cure: Check if cassette is clamped correctly; restart digitizer to execute selftest again.
4215 | AUTO_CLEAN_UP_IHC_FAILURE_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4216 | AUTO_CLEAN_UP_CHC_SELFTEST_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4217 | AUTO_CLEAN_UP_IHC_SELFTEST_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4218 | FEED_IP_AFTER_AUTO_CLEAN_UP_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4219 | ICC_DID_NOT_DETECT_BOS | Error |  |  |  | 'scan_IM_BOS' intermediate response from ICC still missing although slowscan has finished the scan already. | Meaning: The scanning unit didn't start the scan process although the slowscan motor has finished the scan already. Reason: ICC (ImageCaptureControl) is still waiting for begin of the scan, i.e. no pixels have been stored yet in the SDRAM of the FPGA. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI BOARD and PMT/lens assembly; if OK, replace PMI; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly.
4220 | IPSTATUS_NOT_EXPOSED | Error | 293 | Error occurred. Image plate is not scanned and erased! Restart digitizer and try again. If error remains, contact service to fix the problem. |  | IP cannot be scanned because IP status is wrong. | Meaning: The IP cannot be scanned because the IP status is wrong. Reason: Digitizer rejects the scan because the NX did not set the IP status to 'exposed'. Normally the IP status is set to 'exposed' by pressing the 'ID' button on the NX. This is a software error. Cure: Check if digitizer is configured correctly on the NX; check NX software installation; if OK, inform R&D team.
4221 | PMM_5V | Error |  |  |  | Check PMM 5 failed | Reason: Fuse for PMM 5V on PMI BOARD or POWER INTEGRATION BOARD defective. Cure: Replace PMM 5V fuse | PMI_PMM5VErr | CIR217 | Supply voltage +5V (PMT) on PMI Board low or missing. Meaning: New (proposal): Test of the supply voltage (+5V) for the PMT failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411) 2) Contact problem (connector P5) or defect on CPMI Power Distribution Board (B411) 3) Contact problem (connector P1#17) or defect on CPMI Board (B410) Cure: 1) Check the supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411). If required replace the supply cable, contact your Regional Support Network for further assistance 2) Check connector P5 on CPMI Power Distribution Board (B411) 3) Check connector P1#17 on CPMI Board (B410) 4) Exchange the CPMI Power Distribution Board (B411), refer to the Service Manual for order and replacement 5) Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement
4222 | CONFIRM_CASSETTE_COVER_BLOCKED | Error | 260 | Digitizer cannot open the cassette. The cassette may be damaged. Press confirm to release the cassette. Use another cassette. Call your local service. | Confirm | cassette cover could not have been opened because cover is blocked | Meaning: Digitizer cannot open the cassette because the cassette cover is blocked. Reason: Cassette is damaged, cassette mechanics is defective. Cure: Confirm message; remove cassette; check if cassette cover can be opened by hand; if not, replace cassette
4223 | SCAN_WARNING_SLOWSCAN_SPEEDFACTOR | Error | 267 | Image geometry may be out of tolerance. Possible bad image quality. Please contact service. |  | Slowscan finished scan with warning that default slowscan speed factor was used for scanning | Meaning: The scan finished with a warning. Reason: Slowscan reported a warning that the default slowscan speed factor was used for scanning, because the slowscan factor which is stored on the slowscan's ELabel is out of the valid range. Therefore the image geometry may be out of tolerance. Cure: Replace slowscan board with a board which contains the correct speed factor for this device serial number. Ask service headquarter how to do this.
4224 | CONFIRM_AUTO_RETURN_RECONNECTED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4225 | ERASE_MODE_ERASE_IP_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4226 | SCAN_MODE_ERASE_IP_FAILED_OBSOLETE | Error |  |  |  | OBSOLETE | OBSOLETE
4227 | PMM_ANALOG_POWER_FAILURE | Error |  |  |  | PMM analog power +20V or -10V failure | Meaning: PMM analog power +20V or -10V failure. Reason:PMI power distribution board defective , cable from PMI power distribution board to PMI BOARD defective. Cure: 1) Check PMM cabling. 2) Check PMI power distribution board cabling. 3) Replace PMI power distribution board. | PMI_PMMAnalogPwrErr | CIR216 | Supply voltage +20V / -10V (PMT) on PMI Board low or missing. Meaning: Test of the supply voltage (+20V / -10V) for the PMT failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411) 2) Contact problem (connector P5) or defect on CPMI Power Distribution Board (B411) 3) Contact problem (connector P1#17) or defect on CPMI Board (B410) Cure: 1) Check the supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411). If required, replace the supply cable, contact your Regional Support Network for further assistance 2) Check connector P5 on CPMI Power Distribution Board (B411) 3) Check connector P1#17 on CPMI Board (B410) 4) Exchange the CPMI Power Distribution Board (B411), refer to the Service Manual for order and replacement 5) Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement
4228 | CALIBRATION_BOARD_POWER_FAILURE | Error |  |  |  | Calibration board power failure | Meaning: Calibration board power failure. Reason: PMI power distribution board defective, cable from PMI power distribution board to PMI BOARD defective. Cure: 1) Check PMM cabling. 2) Check PMI power distribution board cabling. 3) Replace PMI power distribution board. | PMI_OpticsPwrErr | CIR215 | Supply voltage +12V / -12V (Optic module) on PMI Board low or missing. Meaning: Test of the supply voltage (+12V / -12V) for the optic module failed. The Digitizer is out of operation and requires service support.. Reason: 1) Contact problem or defect at supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411) 2) Contact problem (connector P5) or defect on CPMI Power Distribution Board (B411) 3) Contact problem (connector P1#17) or defect on CPMI Board (B410). Cure: 1) Check the supply cable between connector P1#17 on CPMI Board (B410) and connector P5 on CPMI Power Distribution Board (B411). If required replace the supply cable, contact your Regional Support Network for further assistance 2) Check connector P5 on CPMI Power Distribution Board (B411) 3) Check connector P1#17 on CPMI Board (B410) 4) Exchange the CPMI Power Distribution Board (B411), refer to the Service Manual for order and replacement 5) Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement.
4229 | INITIALIZE_CAN | Error |  |  |  | Initialize CAN BUS failed | Meaning: Initialize CAN BUS failed; Reason: CAN hardware or software failure; Cure: 1) Check CAN BUS cabling; 2) Install new software; 3) Replace PMI BOARD.
4230 | DIGITIZER_SW_VERSION_MISMATCH | Error | 308 | There is an inconsistency in the software packages installed or an electronic board is disconnected. Contact service to finish installation. |  | The software versions of the internal digitizer modules do not match. | Meaning: Software versions loaded onto internal digitizer modules are not matching. Reason: Incomplete software installation, installation failed or hardware replacement (digitizer boards). Cure: Install digitizer software again.
4231 | UNKNOWN_DESTINATION | Error |  |  |  | destination is not known | Meaning: The destination name is not known. Reason: The configured destination (= name of the processing station) is not known by the digitizer. Cure: Set the name of the destination correctly in the digitizer settings (on the NX via ADC.cpf or by hand in the service client) and reset the digitizer. Then insert the cassette again.
4232 | CONFIRM_UNLOCK_AND_REMOVE_CASS_CLOSE_FAILED | Error | 323 | Error occurred. Cassette cannot be closed. Press confirm to unlock the cassette. To remove cassette, pull at left hand side. Do not use the cassette again before an inspection. | Confirm | Cassette cannot be closed; remove open cassette out of input slot. | Meaning: Cassette cannot be closed by opener drive. Reason: Cassette slot is blocked by something. Cure: Confirm message. Then cassette will be unfixed. Remove empty cassette with open flap by turning out on left hand side.
4233 | CONFIRM_MOVE_IP_OUT_TO_SLOT | Error | 324 | The image plate will be moved to the cassette slot as soon as you confirm. Press confirm and withdraw the image plate. | Confirm | IP in digitizer without cassette. Moves IP out to input slot. | Meaning: Image plate is in digitizer but no cassette. Reason: Empty cassette was removed by user . Cure: Confirm message. Then image plate is moved out through open input slot. Remove image plate.
4234 | CONFIRM_IP_WITHDRAWN_FROM_SLOT | Error | 325 | Press confirm after you have withdrawn the image plate. | Confirm | Image plate is in open input slot to withdraw. | Meaning: Image plate in digitzer was moved out to open input slot. Reason: No cassette in digitizer. Cure: Withdraw image plate from input slot, then Confirm message.
5122 | WRONG_STATE | Error |  |  |  | The received event cannot be handled in the actual state of CaSSC. | Meaning: The received event cannot be handled in the actual state of CaSSC. The event and the state can be seen in the event logging. Reason: An event was sent at the wrong time to the CaSSC. Cure: Verify that the event is sent to the correct state of the CaSSC state table.
5123 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE_SLOW_SCAN | Error |  |  |  | The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScan. | Meaning: The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScan. Reason: A wrong state or has been set. Cure: Check the corresponding SW.
5124 | MOTOR_MOVEMENT_FAILED | Error |  |  |  | A movement of the SlowScan or the SlowScanLift was not accomplished correctly. The specific movement is coded in the previous error. | Meaning: A movement of the SlowScan or the SlowScanLift was not accomplished correctly. The specific movement is coded in the previous error. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
5125 | LIFT_CLOSE_PRESCAN_ROLLS_FOR_ALIGNMENT | Error |  |  |  | Slowscan lift motor movement to close prescan rollers for alignment of IP failed. | Meaning: Closing the slowscan prescan rollers with slowscan lift motor failed. Reason: Wrong steering parameters. Cure: Reset digitizer and try again. If error remains, reinstall digitizer Software. | SSU_SSLMovementError | CIR264 | MOTOR M7 (slow scan roller lift) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan roller lift (MOTOR M7) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan roller lift (MOTOR M7) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5126 | LIFT_MOVE_BLADE_DOWN | Error |  |  |  | Slowscan lift motor movement to move alignment blade down failed. | Meaning: Moving alignment blade down with slowscan lift motor failed. Reason: Wrong steering parameters. Cure: Reset digitizer and try again. If error remains, reinstall digitizer Software. | SSU_SSLMovementError | CIR264 | MOTOR M7 (slow scan roller lift) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan roller lift (MOTOR M7) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan roller lift (MOTOR M7) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5127 | LIFT_CLOSE_POSTSCAN_ROLLS | Error |  |  |  | Slowscan lift motor movement to close postscan rollers failed. | Meaning: Closing the slowscan postscan rollers with slowscan lift motor failed. Reason: Wrong steering parameters. Cure: Reset digitizer and try again. If error remains, reinstall digitizer Software. | SSU_SSLMovementError | CIR264 | MOTOR M7 (slow scan roller lift) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan roller lift (MOTOR M7) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan roller lift (MOTOR M7) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5128 | LIFT_CLOSE_BOTH_ROLLS | Error |  |  |  | Slowscan lift motor movement to close both roller pairs failed. | Meaning: Closing both slowscan roller pairs with slowscan lift motor failed. Reason: Wrong steering parameters. Cure: Reset digitizer and try again. If error remains, reinstall digitizer Software. | SSU_SSLMovementError | CIR264 | MOTOR M7 (slow scan roller lift) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan roller lift (MOTOR M7) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan roller lift (MOTOR M7) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5129 | LIFT_FIND_REFERENCE | Error |  |  |  | Slowscan lift motor movement find Reference failed. | Meaning: Find reference position of slowscan lift motor failed. Reason: Reference light barrier not found. Cure: Check refernce light barrier, cable and HandlingBoard, and change defective part. | SSU_SSLFindRefError | CIR265 | MOTOR M7 (slow scan roller lift) does not detect reference position. Meaning: The drive (MOTOR M7) of the slow scan roller lift does not reach the reference position 'anterior slow scan roller lifted'. The Digitizer is out of operation and requires service support. Reason: 1) Contact problems of cable or adapter plug of MOTOR M7 2) Contact problems of cable or adapter plug of light barrier Lift Ref Pos (B426) 3) Drive belt or the belt pulley is not proper fixed or defect 4) Light barrier or light barrier flag is not proper fixed or defect 5) Any obstacle in the travel range of the slow scan roller lift drive Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the cabling and the adapter plug of MOTOR M7 3) Check the cabling and the adapter plug of light barrier Lift Ref Pos (B426) 4) Perform the 'Prepare Cal. Board Replacement' via the Service User Interface and check visually at the ROLLER LIFT ASSEMBLY RIGHT, if the upper scan rollers are lifted 5) If the upper scan rollers are not lifted, check the drive belt, the belt pulley, the light barrier, the light barrier flag and the slow scan roller lift drive for any defect and proper fixing 6) For replacement of SLOW SCAN ROLLER LIFT contact your Regional Support Network for further assistance
5130 | LIFT_PREPARE | Error |  |  |  | Slowscan lift motor movement to prepare for feeding IP position failed. | Meaning: Moving to prepare for feeding IP position with slowscan lift motor failed. Reason: Wrong steering parameters. Cure: Reset digitizer and try again. If error remains, reinstall digitizer Software. | SSU_SSLMovementError | CIR264 | MOTOR M7 (slow scan roller lift) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan roller lift (MOTOR M7) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan roller lift (MOTOR M7) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5131 | SLOWSCAN_SCANNING | Error |  |  |  | Slowscan motor movement scan IP failed. | Meaning: Scanning movement with slowscan motor failed. Reason: End of Image Plate not detected in transport control light barrier. Cure: Check light barrie r and replace if defective. Check slowscan unit for proper IP transport.
5132 | SLOWSCAN_RETURN_IP_LB | Error |  |  |  | Slowscan motor movement return IP failed. | Meaning: Returning IP movement with slowscan motor failed. Reason: Image Plate was not detected in transport control light barrier. Cure: Check light barrier and replace if defective. Check slowscan unit for proper IP transport.
5133 | SLOWSCAN_IP_FAST_IN_SCANNER | Error |  |  |  | Slowscan motor movement of IP fast into scanner failed. | Meaning: Moving IP fast into scanner with slowscan motor failed. Reason: End of Image Plate not detected in transport control light barrier. Cure: Check light barrie rand replace if defective. Check slowscan unit for proper IP transport.
5134 | IP_NOT_AT_PRE_SCAN_ROLLERS | Error |  |  |  | The IP was not transported to the pre-scan rollers correctly. | Meaning: The IP was not transported to the pre-scan rollers correctly. 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction. | SSU_SSMovementError | CIR254 | MOTOR M2 (slow scan) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan (MOTOR M2) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5135 | IP_NOT_AT_POST_SCAN_ROLLERS | Error |  |  |  | The IP was not transported to the post-scan rollers correctly. | Meaning: The IP was not transported to the post-scan rollers correctly. 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction. | SSU_SSMovementError | CIR254 | MOTOR M2 (slow scan) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan (MOTOR M2) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5136 | REAR_PART_OF_IP_NOT_AT_POST_SCAN_ROLLERS | Error |  |  |  | The rear part of the IP did not reach the post-scan rollers correctly. | Meaning: The rear part of the IP did not reach the post-scan rollers correctly. 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction. | SSU_SSMovementError | CIR254 | MOTOR M2 (slow scan) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan (MOTOR M2) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5137 | IP_NOT_AT_EOS | Error |  |  |  | The IP was not transported to the end of scan correctly. | Meaning: The IP was not transported to the end of scan correctly. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction. | SSU_SSMovementError | CIR254 | MOTOR M2 (slow scan) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan (MOTOR M2) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5138 | IP_NOT_AT_BOE | Error |  |  |  | The IP was not transported to the begin of erasure correctly. | Meaning: The IP was not transported to the begin of erasure correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5139 | WRONG_COMMAND | Error |  |  |  | The received command is unknown and cannot be handled by the CaSSC. | Meaning: The received command is unknown and cannot be handled by the CaSSC. Reason: A wrong event was sent to the CaSSC. Cure: Verify that the corrrect event is sent to the CaSSC.
5140 | CALL_BACK_POINTER_ZERO | Error |  |  |  | The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. | Meaning: The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. Reason: -- Cure: Check the (initialization of) pointers.
5141 | MOVE_SLOW_SCAN_FOR_CLEAN_UP_1_FAILED | Error |  |  |  | The SlowScan movement 1 for cleaning up the scanner failed. | Meaning: The SlowScan movement 1 for cleaning up the scanner failed (1: direction into the cassette). 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction. | SSU_SSMovementError | CIR254 | MOTOR M2 (slow scan) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan (MOTOR M2) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5143 | STOP_SLOW_SCAN_FAILED | Error |  |  |  | The SlowScan did not stop correctly. | Meaning: The SlowScan did not stop correctly. 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction. | SSU_SSMovementError | CIR254 | MOTOR M2 (slow scan) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan (MOTOR M2) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
5144 | SPEED_FACTOR_OUT_OF_REQUIRED_RANGE | Error |  |  |  | The speed factor is out of the required range. | Meaning: The speed factor is out of the required range. Reason: The corresponding e-label value was not correct or not correctly read out. Cure: Check the corresponding e-label value and the reading procedure.
5145 | SLOWSCAN_RETURN_IP | Error |  |  |  | Slowscan motor movement return IP failed. | Meaning: Returning IP movement with slowscan motor after scan failed. Reason: Image Plate was not detected in transport control light barrier. Cure: Check light barrier and replace if defective. Check slowscan unit for proper IP transport.
5632 | INIT_STEPPERS_RETRY_MAX_LIMIT | Error |  |  |  | Unspecified movement error during initialisation. | Meaning:Unspecified movement error during initialisation. Positioning of any drive failed. Reason: unknown. Cure: 1) Restart Digitizer. 2)If error persists, check all cassette handling drives with corresponding mechanics and replace defective part(s).
5633 | OPENER_CLOSE_FLAP | Error |  |  |  | Close cassette flap movement of opener drive failed. | Meaning: The cassette flap could not be closed correctly. The CASSETTE OPENER MOTOR M3 did not reached the correct target position for cassette closed.;The cassette is still clamped and could not be released.The Digitizer is out of operation and requires service support. Reason: The CASSETTE OPENER MOTOR M3 could not close the cassette flap successfully due to;- Tray or Image Plate is still outside of the cassette and blocks the shutter.- Image Plate is damaged or bended and blocks the shutter.- Drawer unit is not properly adjusted and blocks the shutter.- Shutter of the cassette is blocked or defect.- Drive mechanism of CASSETTE OPENER MOTOR M3 is blocked or defect. Cure: 1. Restart the Digitizer and check, if the problem persists. 2. If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual. 3. Check the tray and the Image Plate for any defects. 4. Check the function of the cassette shutter and the cassette inside for any particle or removed fleece. 5. Check the proper position of the DRAWER UNIT, refer to the Service Manual, adjustments. 6. Check the drive mechanism of CASSETTE OPENER MOTOR M3 for proper function and smooth running. 7. Exchange the CASSETTE OPENER UNIT, refer to the Service Manual for order and replacement. 8 Check the drive mechanism of CASSETTE FIXATION UNIT for proper function and smooth running | CHM_OpenerCloseError | CIR525 | MOTOR M3 (cassette opener) does not detect stop position. Meaning: The drive (MOTOR M3) of the cassette opener does not detect the reference position 'cassette flap closed' (lower stopper). The Digitizer is out of operation and requires service support. Reason: 1) Cassette flap can not be closed due to a blocked tray or Image Plate or due to a defect at the cassette 2) Drive belt or the belt pulley is not proper fixed or defect 3) Any obstacle in the travel range of the cassette opener drive or defects at the gear rack 4) Sluggishness of the cassette opener drive. Cure: 1. Restart the Digitizer and check, if the problem persists. 2. If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual. 3. Check the tray, the Image Plate and the cassette flap for any defects. 4. Check the function of the cassette shutter and the cassette inside for any particles or removed fleece. 5. Check the drive belt, the belt pulley on the motor shaft and the gear rack for any defects 6. Check the drive mechanism of Cassette Opener for proper function and smooth running. 7. Perform the 'Diagnose Cycle' in the Service User Interface, refer to the Service Manual 8. Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
5634 | OPENER_OPEN_FLAP | Error |  |  |  | Open cassette flap movement of opener drive failed. | Meaning: Cassette flap open position not reached by opener drive. Reason: Drive is blocked by something or cassette is damaged. Cure: - Check cassette on damage if error occurres only with this cassette. Check mechanics and electronis of opener drive and repair/exchange defective parts. | CHM_OpenerOpenError | CIR524 | MOTOR M3 (cassette opener) moves with difficulty. Meaning: The drive (MOTOR M3) of the cassette opener does not reach the end position 'cassette flap open'. The Digitizer is out of operation and requires service support. Reason: 1) Cassette flap can not be opened due to a defect at the cassette 2) Drive belt or the belt pulley is not proper fixed or defect 3) Any obstacle in the travel range of the cassette opener drive or defects at the gear rack 4) Sluggishness of the cassette opener drive Cure: 1) Restart the Digitizer and check, if the problem persists 2) If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual. 3) Check the cassette flap for any defects 4) Check the drive belt, the belt pulley on the motor shaft and the gear rack for any defects 5) Check the drive mechanism of Cassette Opener for proper function and ease of movement 6) Perform the 'Diagnose Cycle' in the Service User Interface, refer to the Service Manual 7) Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
5635 | FIXER_UNCLAMP_CASS | Error |  |  |  | Unclamp Cassette Movement of fixer drive failed. | Meaning: Cassette unclamp position not reached by Fixer drive. Reason: Drive is blocked by something or movement parameters are wrong. Cure:Repeat Stall Calibration of Fixer drive, if error persists, check mechanics and electronis of Fixer drive and repair/exchange defective parts. | CHM_FixerUnclampError | CIR585 | MOTOR M1 (cassette fixing) does not detect stop position. Meaning: The drive (MOTOR M1) of the cassette fixing does not reach the reference position 'cassette unclamped'. The Digitizer is out of operation and requires service support. Reason: 1) Coupling on the motor shaft is not proper fixed or defect 2) Towing bracket on the spindel bush or the spindel bush itself is not proper fixed or broken 3) Infrequent contact problems of cable or MOTOR M1 4) STALL-Calibration line out of range Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the coupling on the motor shaft for proper fixation and function 4) Check the towing bracket and the spindel bush for any defects like stress cracks 3) Perform the 'Diagnose Cycle' in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 4) Perform a STALL-Calibration of MOTOR M1 (cassette fixing), refer to the Service Manual 5) If the problem remains, exchange the CASSETTE FIXATION UNIT, refer to the Service Manual for order and replacement.
5636 | CLEAN_UP_FAILED | Error |  |  |  | Unspecified failure during cleanUp | Meaning: Any unspecified error occurred during CleanUp cycle. Reason: unknown Cure: 1) If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual. 2. Restart the Digitizer, if the problem persists reinstall software 3. Check all drives, mechanical and electronic parts of the cassette unit and replace defective parts.
5637 | OPENER_ABS_POS_INIT_MODULE_RETRY_MAX_LIMIT_OBSOLETE | Error |  |  |  | OpenerAbsPos Move in initModule, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Initialisation of Stepper Motor Opener in Cassette Handling incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service. 1. Check Opener Mechanics for obstacles, Check the driving elements for getting stuck or blocked, Ensure easy smooth movement. 2. Check Repair Part 'TOOTHED BELT FOR CASSETTE UNIT' Related Repair Parts are also 'CASSETTE OPENER' and 'GS CR30 IP-HANDLING'
5638 | LOCKER_LOCK_IP | Error |  |  |  | Lock IP Movement of Locker drive failed. | Meaning: Lock IP position in cassette not reached by Locker drive. Reason: Drive is blocked by something or cassette is damaged. Cure: 1) Check cassette on damage if error occurres only with this cassette. 2) Check mechanics and electronis of Locker drive and repair/exchange defective parts. | CHM_LockerLockIPError | CIR564 | Drive of MOTOR M17 (cassette unlatch assembly) moves with difficulties. Meaning: The drive (MOTOR M17) of the cassette unlatch assembly does not reach the end position 'unlock pin pushed in'. The Digitizer is out of operation and requires service support. Reason: 1) Cassette unlock pin can not push in the cassette unlocker due to a defect at the cassette 2) Cam disk on the motor shaft or pin on the cam disk not proper fixed or defekt 3) Sluggishness of the cassette unlatch drive, e.g. guide rail get jammed 4) Infrequent contact problems of cable or MOTOR M17 Cure: 1) Restart the Digitizer and check, if the problem persists 2) If the cassette is still clamped, unclamp the cassette manually and remove it 3) Check the cassette unlocker for proper function and ease of movement 4) Check the cam shaft with the pin for proper fixation and function 5) Check the drive mechanism of the CASSETTE UNLATCH ASSEMBLY for proper function and ease of movement 6) Perform the 'Test Cycle' in the Service User Interface with different cassettes, refer to the Service Manual. If the error is related to specific cassettes, replace the affected cassettes. 7) Perform the 'Diagnose Cycle' in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 8) Exchange the CASSETTE UNLATCH ASSEMBLY, refer to the Service Manual for order and replacement.
5639 | LOCKER_UNLOCK_IP | Error |  |  |  | UnLock IP Movement of Locker drive failed.. | Meaning: Unock IP position out of cassette not reached by Locker drive. Reason: Drive is blocked by something or movement parameters are wrong. Cure: 1) Repeat Stall Calibration of Locker drive 2) check mechanics and electronis of Locker drive and repair/exchange defective parts. | CHM_LockerRefPosError | CIR565 | MOTOR M17 (cassette unlatch assembly) does not detect stop position. Meaning: Locker Reference position not found. Initialisation incomplete. Reason: Mechanical stop(reference position) not found. Drive is blocked by something or movement parameters are wrong. Cure: 1) check mechanics and electronis of Locker drive and repair/exchange defective parts. 2) Repeat Stall Calibration of Locker drive
5640 | INIT_MODULE_ABS_POS_RETRY_MAX_LIMIT_OBSOLETE | Error |  |  |  | AbsPos Move in initModule, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Initialisation incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Unspecific failure for motor Fixer, Locker, Opener. Cure: Restart Digitizer. If error persists please call Service. 1. Check cables of motor Fixer, Locker, Opener.
5641 | LOCKER_FIND_REFERENCE | Error |  |  |  | Find Reference Movement of Locker drive failed. | Meaning: Locker Reference position not found. Initialisation incomplete. Reason: Mechanical stop not found. Cure: 1) Repeat Stall Calibration of Locker drive. 2) Check Locker Mechanics if blocked by something or moving stiff and repair. | CHM_LockerRefPosError | CIR565 | MOTOR M17 (cassette unlatch assembly) does not detect stop position. Meaning: Locker Reference position not found. Initialisation incomplete. Reason: Mechanical stop(reference position) not found. Drive is blocked by something or movement parameters are wrong. Cure: 1) check mechanics and electronis of Locker drive and repair/exchange defective parts. 2) Repeat Stall Calibration of Locker drive
5642 | FIXER_FIND_REFERENCE | Error |  |  |  | Find Reference.Movement of Fixer drive failed. | Meaning: Fixer Reference position not found. Initialisation incomplete. Reason: Mechanical stop not found. Cure: 1) Repeat Stall Calibration of Fixer drive. 2) Check Fixer Mechanics if blocked by something or moving stiff and repair. | CHM_FixerUnclampError | CIR585 | MOTOR M1 (cassette fixing) does not detect stop position. Meaning: The drive (MOTOR M1) of the cassette fixing does not reach the reference position 'cassette unclamped'. The Digitizer is out of operation and requires service support. Reason: 1) Coupling on the motor shaft is not proper fixed or defect 2) Towing bracket on the spindel bush or the spindel bush itself is not proper fixed or broken 3) Infrequent contact problems of cable or MOTOR M1 4) STALL-Calibration line out of range Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the coupling on the motor shaft for proper fixation and function 4) Check the towing bracket and the spindel bush for any defects like stress cracks 3) Perform the 'Diagnose Cycle' in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 4) Perform a STALL-Calibration of MOTOR M1 (cassette fixing), refer to the Service Manual 5) If the problem remains, exchange the CASSETTE FIXATION UNIT, refer to the Service Manual for order and replacement.
5643 | CASSETTE_EDGE_SWITCH_NOT_CONNECTED | Error |  |  |  | The edge switch for cassette detection is not connected correctly. | Meaning: The edge switch for cassette detection is not connected correctly. Reason: -- Cure: Check the switch and the corresponding cabling. Replace the defective parts. | CHM_SensorCassEdgeConnectionError | CIR423 | Contact or cable between Handling Control Board and Cass. Edge Sensor problem. Meaning: Connection test between the Handling Control Board and the Cassette Edge Sensor (S1) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at cable between Handling Control Board and Cassette Edge Sensor (S1) 2) Contact problem or defect at connectors of the Cassette Edge Sensor (S1) 3) Contact problem (connector P15) or defect on the Handling Control Board. Cure: 1) Check the cable between Handling Control Board and Cassette Edge Sensor (S1). If required replace the cable, refer to the Service Manual for order and replacement 2) Check contacts on the Cassette Edge Sensor (S1) 3) Check connector P15 on the Handling Control Board 4) Exchange the Cassette Fixation Unit, refer to the Service Manual for order and replacement 5) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
5644 | OPENER_FIND_REFERENCE | Error |  |  |  | Find Reference.Movement of Opener drive failed. | Meaning: Opener Reference position not found. Initialisation incomplete. Reason: Mechanical stop not found. Cure: Check Opener Mechanics if blocked by something or moving stiff and repair. | CHM_OpenerCloseError | CIR525 | MOTOR M3 (cassette opener) does not detect stop position. Meaning: The drive (MOTOR M3) of the cassette opener does not detect the reference position 'cassette flap closed' (lower stopper). The Digitizer is out of operation and requires service support. Reason: 1) Cassette flap can not be closed due to a blocked tray or Image Plate or due to a defect at the cassette 2) Drive belt or the belt pulley is not proper fixed or defect 3) Any obstacle in the travel range of the cassette opener drive or defects at the gear rack 4) Sluggishness of the cassette opener drive. Cure: 1. Restart the Digitizer and check, if the problem persists. 2. If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual. 3. Check the tray, the Image Plate and the cassette flap for any defects. 4. Check the function of the cassette shutter and the cassette inside for any particles or removed fleece. 5. Check the drive belt, the belt pulley on the motor shaft and the gear rack for any defects 6. Check the drive mechanism of Cassette Opener for proper function and smooth running. 7. Perform the 'Diagnose Cycle' in the Service User Interface, refer to the Service Manual 8. Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
5645 | SENSOR_OPENER_OPENFLAP_NOT_CONNECTED | Error |  |  |  | The open control sensor of cassette opener drive is not connected. | Meaning: The open control sensor of cassette opener drive is not connected. Reason: -- Cure: Check the sensor and the corresponding cabling. Replace the defective parts. | CHM_SensorOpenerOpenConnectionError | CIR532 | Contact or cable to SENSOR B431 (cassette opener sensor) problem. Meaning: Selftest of sensor B431 (cassette opener sensor) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at cable between Handling Control Board and Cassette Opener Sensor (B432) 2) Contact problem at connectors or defect of the Cassette Opener Sensor (B432) 3) Contact problem (connector P20) or defect on the Handling Control Board Cure: 1) Check the data cable between Handling Control Board and Cassette Opener Sensor (B432). If required exchange the data cable, refer to the Service Manual for order and replacement 2) Check connector on the Cassette Opener Sensor (B432). If required exchange the Cassette Opener Sensor (B432), refer to the Service Manual for order and replacement 3) Check connector P20 on the Handling Control Board 4) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
5646 | OPENER_FIND_REF_STEPDEV_NEG | Error |  |  |  | Find Reference.Movement of Opener drive failed with less steps moved. | Meaning: Cassette Opener motor(M3) reference positioning failed with less steps moved. Reason: 1) Opener drive or mechanics damaged or runs stiff 2) Upper stopper not mounted correctly or fixed at worng position. Cure: 1) Check opener mechanics for any damage or blockade (fixing screws of upper stopper ..). 2) Replace Opener Unit. | CHM_Opener_StepDeviationNeg | CIR526 | MOTOR M3 (cassette opener) detected less steps. Meaning: The drive (MOTOR M3) of the cassette opener detected less steps between the upper and the lower stop position. The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley is not proper fixed or defect 2) Defects at the gear rack or gearwheels 3) Any obstacle in the travel range of the cassette opener drive, e.g. upper stopper with fixing screws not proper mounted 4) Sluggishness of the cassette opener drive. Cure: 1) Check the drive belt, the belt pulley on the motor shaft and the gear rack for any defects 2) Check, if the rubber cushions at the upper and lower stopper are propper fixed 3) Check, if the two fixing screws of the upper stopper does not touch the cassette opener bracket 4) Check the whole drive mechanism of Cassette Opener for proper function and smooth running. 5) Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
5647 | OPENER_FIND_REF_STEPDEV_POS | Error |  |  |  | Find Reference.Movement of Opener drive failed with too much steps moved. | Meaning: Cassette Opener motor(M3) reference positioning failed with too much steps moved. Reason: 1) Upper or lower rubber damper mounted not correctly 2) Upper stopper not fixed correctly or at worng position. 3) Lower stopper not fixed correctly. Cure: 1) Check both rubber damper. 2) Check upper stopper if correct mounted. 3) Check Lower stopper if corect mounted. 4) Insert washer (max 1mm) between upper stopper and rubber damper. 5) Replace Opener Unit. | CHM_Opener_StepDeviationPos | CIR527 | MOTOR M3 (cassette opener) detected too much steps. Meaning: The drive (MOTOR M3) of the cassette opener detected too much steps between the upper and the lower stop position. The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley is not proper fixed or defect 2) Rubber cushions at the upper or lower stopper are not proper fixed 3) Upper or lower stopper not proper fixed Cure: 1. Check the drive belt, the belt pulley on the motor shaft for any defects 2. Check, if the rubber cushions at the upper and lower stopper are proper fixed 3. Check, if the upper and lower stopper are proper mounted 4. Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
5732 | WARN_CLAMP_CASS_INVALID_WIDTH | Warning | 116 | Unknown cassette format detected (cassette inserted correctly?). Remove cassette and insert it again. |  | Warning: Clamp Cassette Movement detected invalid Cassette Width | Meaning: Clamping cassette failed. Reason: The clamped Cassette has a invalid width. Cure: 1) Check if input of cassette was correct. 2) Check cassette on mechanical damage if only this cassette failes. 3) Check fixer drive and mechanics for defective parts and replace.
6145 | WRONG_STATE | Error |  |  |  | The received event cannot be handled in the actual state of CaIPH. | Meaning: The received event cannot be handled in the actual state of CaSSC. The event and the state can be seen in the event logging. Reason: Software error. Cure: Software bugfix.
6146 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE_OBSOLETE | Error |  |  |  | The received event cannot be handled in the actual state of a low level state machine (in a passive class). | Meaning: The received event cannot be handled in the actual state of a low level state machine (in a passive class). Reason: A wrong state has been set. Cure: Check the corresponding SW.
6147 | MOVE_DRAWER_OUT_OF_CASSETTE_FAILED | Error |  |  |  | Drawer unit movement out of the cassette failed. | Meaning: The Drawer unit movement out of cassette to the reference stopper was not finished successfully. Reason: 1) Stall detection parameters are wrong. 2) The drawer unit is blocked by something, is running stiff or not moving at all. Cure: 1) Do stall calibration of Drawer unit. 2) Check Drawer mechanics and repair defective parts. 3) Check Drawer motor and cables and replace defective parts. | IPHM_DrawerRefPosError | CIR675 | MOTOR M4 (drawer unit) does not detect stop position. Meaning: The drive of the drawer unit does not detect the reference position (stopper). The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley is not proper fixed or defect 2) Any obstacle in the travel range of the drawer unit drive, stopper not proper fixed 3) Sluggishness of the drawer unit drive 4) Infrequent contact problems of cable or MOTOR M3 5) STALL-Calibration line out of range. Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the drive belt and the belt pulley on the motor shaft for any defects 3) Check it the rubber buffer at the stop position is proper fixed 4) Check the drive mechanism of DRAWER UNIT for proper function and ease of movement 5) Perform the 'Endurance Run Cycle' (e.g. 20 cycles) in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 6) Perform a STALL-Calibration of MOTOR M4 (drawer unit), refer to the Service Manual 7) If the problem remains, exchange the DRAWER UNIT, refer to the Service Manual for order and replacement.
6149 | MOVE_DRAWER_INTO_CASSETTE_FAILED | Error |  |  |  | Drawer unit movement into the cassette failed. | Meaning: The Drawer unit movement into cassette was not finished successfully. Reason: 1) Stall detection parameters are wrong. 2) The Drawer unit is blocked by something, is running stiff or not moving at all. Cure: 1) Do stall calibration of drawer unit. 2) Check Drawer mechanics and repair defective parts. 3) Check drawer motor and cables and replace defective parts. | IPHM_DrawerPositionError | CIR674 | MOTOR M4 (drawer unit) moves with difficulty. Meaning: The drive (MOTOR M4) of the of the drawer unit does not reach one of the end positions. The Digitizer is out of operation and requires service support. Reason: 1) Cassette flap can not be opened completely due to a defect at the cassette 2) Drive belt or the belt pulley is not proper fixed or defect 2) Any obstacle in the travel range of the drawer unit drive. 3) Sluggishness of the drawer unit drive 4) Infrequent contact problems of cable or MOTOR M4 5) STALL-Calibration line out of range. Cure: 1) Restart the Digitizer and check, if the problem persists 2) If the cassette is still clamped, unclamp the cassette manually and remove it 3) Check the cassette flap for any defects 4) Check the drive belt and the belt pulley on the motor shaft for any defects 5) Check the drive mechanism of DRAWER UNIT for proper function and ease of movement 6) Perform the 'Diagnose Cycle' or the 'Endurance Run Cycle' (e.g. 20 cycles) in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 7) Perform a STALL-Calibration of the drawer unit, refer to the Service Manual 8) If the problem remains, exchange the DRAWER UNIT, refer to the Service Manual for order and replacement.
6151 | MOVE_DRAWER_OUTOF_CASS_PARTLY_FAILED | Error |  |  |  | The Drawer movement out of cassette partly failedy. | Meaning: The Drawer movement out of cassette partly was not finished successfully. Reason: wrong movement parameters. Cure: 1) Reset digitizer. 2) If error persists, reinstall software. | IPHM_DrawerMovementError
6152 | MOVE_IP_TO_SCANNER_FAILED_OBSOLETE | Error | 136 | Open digitizer to check for image plate jam or check for an empty cassette and restart the digitizer. |  | Moving the IP to the scanner with transport rolls failed. | Meaning: Moving the IP to the scanner with transport rolls failed. Transport control lightbarrier was not interrupted. Reason: 1) IP-Jam. 2) empty cassette. 3) Any problem with Drawer, Roller or Roller Clamper drive unit. 4) IP-transport control sensor defective. Cure: 1) Open digitizer and remove IP. 2) test cassette for IP inside. 3) Check described drive units for proper work and repair defective mechanical or electronical parts 4) Check IP-transport control sensor with cable and handling board and replace defective parts.
6153 | IP_JAM_ROLLERS_CLEANUP | Error |  |  |  | IP jam occurred during cleanUp cycle. | Meaning: The IP cannot be transported out of control lightbarrier with rollers and slowscan drive. Reason: 1) IP is blocked by something. 2) IP-Transport control sensor not working correctly. 3) Any problem with rollers drive. 4) Any problem with RollerClamper drive. 5) Any problem with Slowscan drive. 6) Any problem with Slowscan lift rolls drive. Cure: 1) Check if IP is inside and blocked by anything, and correct failure.2) Check IP-transport control sensor with cable and replace defective parts. 3) Check if all the drives work correctly and replace defective parts.
6156 | ROLLERCLAMPER_CLAMP_IP | Error |  |  |  | The Roller Clamper movement clamping IP failed. | Meaning: The Roller Clamper movement clamping IP was not finished successfully. Reason: The Roller Clamper drive is blocked by something or not moving at all. Cure: Check Roller Clamper drive mechanics and electronics for proper work and replace defective parts. | IPHM_RollClamperClampError | CIR604 | MOTOR M6 (roller clamping) moves with difficulty. Meaning: The drive (MOTOR M6) of the roller clamping does not reach the end position 'roller clamped'. The Digitizer is out of operation and requires service support. Reason: 1) Any obstacle in the travel range of the drive and lifting mechanism of the roller clamping 2) Drive shaft is not proper mounted: Gear wheels at the left and right side on the drive shaft are meshing the gear racks left and right with an offset (not in sync) 3) Infrequent contact problems of cable or MOTOR M6 Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the gear wheels and gear racks on the left and right side of the drive shaft for proper function and ease of movement 3) Check the lifting mechanism of the roller clamping for proper function and ease of movement 4) Check, if the gear wheel drive on the left and right side are in sync to lift the rollers on both sides in the same way 5) Perform the 'Endurance Run Cycle' (e.g. 20 cycles) in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 6) For replacement of the Input Tray contact your Regional Support Network for further assistance
6157 | ROLLERCLAMPER_UNCLAMP_IP | Error |  |  |  | The Roller Clamper movement unclamping IP failed. | Meaning: The Roller Clamper movement unclamping IP to stopper was not finished successfully. Reason: 1) Stall detection parameters are wrong. 2) The Roller Clamper drive is blocked by something or not moving at all. Cure: 1) Do stall calibration of Roller Clamper unit. 2)Check Roller Clamper drive mechanics and electronics for proper work and replace defective parts. | IPHM_RollClamperUnclampError | CIR605 | MOTOR M6 (roller clamping) does not detect stop position Meaning: The drive (MOTOR M6) of the roller clamping does not reach the reference position 'roller unclamped'. The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley on the motor shaft is not proper fixed or defect 2) Gear wheels or gear racks on the left and right side of the drive shaft are not proper fixed or broken 3) Any obstacle in the travel range of the drive and lifting mechanism of the roller clamping 4) Gear wheels at the left and right side on the drive shaft are meshing the gear racks left and right with an offset (not in sync) 5) Leaf spring at stop position of the roller lift mechanism not in position or broken 6) Infrequent contact problems of cable or MOTOR M6 7) STALL-Calibration line out of range. Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the drive belt or the belt pulley on the motor shaft for proper fixation and function 3) Check the gear wheels and gear racks on the left and right side of the drive shaft for proper fixation and function 4) Check, if the gear wheel drive on the left and right side are in sync to lift the rollers on both sides in the same way 5) Check the drive and lifting mechanism of the roller clamping for proper function and ease of movement 6) Check the leaf spring at stop position of the roller lift mechanism 7) Perform the 'Endurance Run Cycle' (e.g. 20 cycles) in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 8) Perform a STALL-Calibration of MOTOR M6 (roller clamping), refer to the Service Manual 9) For replacement of the roller clamping contact your Regional Support Network for further assistance.
6161 | RETURN_IP_TO_TRAY_FAILED | Error |  |  |  | The IP transport back to the tray failed. | Meaning: The IP could not be returned to the tray correctly. IP did not leave transport control lightbarrier. Reason:1) IP is blocked by something. 2) Any problem with Roller drive or RollerClamper drive. Cure: 1) Open digitizer and remove IP. 2) Check Roller and RollerClamper units for proper work and repair defective mechanical or electronical parts.
6162 | MOVE_IP_INTO_SLOW_SCAN_UNIT_FAILED | Error | 133 | Image plate jam. Please restart the digitizer and erase the image plate. |  | The IP transport into the Slowscan unit failed. | Meaning: The IP could not be transported into SlowScanUnit correctly. IP-transport control lightbarrier did not become free. Reason: 1) IP-Jam. IP is blocked by something. 2) Any problem with Slowscan drive, SlowscanLift drive or RollerClamper drive Cure: 1) Open digitizer and remove IP 2) Check described drive units for proper work and repair defective mechanical or electronical parts .
6165 | ERASURE_UNIT_DATA_CABLE | Error | 137 | Failure at erasure module. Please restart digitizer. If error remains, please contact service. If an image has been scanned, the image plate must be erased again after the repair of the digitizer. |  | The data cable of the erasure unit is not connected. | Meaning: The data cable of the erasure unit is not connected. Reason: -- Cure: 1) Check/replace the cable connection between handling control board and erasure unit. 2) Check/replace erasure unit. 3) Check/replace handling control board. | IPHM_ErasureUnitDataCableError | CIR424 | Contact or cable between Handling Control Board and Erasure Unit problem. Meaning: Connection test between the Handling Control Board and the Erasure Unit failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at data cable between Handling Control Board and Erasure LED Control Board (B424) 2) Contact problem (connector P2) or defect on the Erasure LED Control Board (B424) 3) Contact problem (connector P16) or defect on the Handling Control Board. Cure: 1) Check the data cable between Handling Control Board and Erasure LED Control Board (B424). If required replace the data cable, refer to the Service Manual for order and replacement 2) Check connector P2 on the Erasure LED Control Board (B424). 3) Check connector P16 on the Handling Control Board 4) Exchange the Erasure Unit, refer to the Service Manual for order and replacement 5) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
6168 | FEED_IP_FAILED | Error | 136 | Open digitizer to check for image plate jam or check for an empty cassette and restart the digitizer. |  | Moving the IP from cassette to scanner failed. | Meaning: Moving the IP to the scanner with transport rolls failed. Transport control lightbarrier was not interrupted. Reason: 1) IP-Jam. 2) empty cassette. 3) Any problem with Drawer, Roller or Roller Clamper drive unit. 4) IP-transport control sensor defective. Cure: 1) Open digitizer and remove IP. 2) test cassette for IP inside. 3) Check described drive units for proper work and repair defective mechanical or electronical parts 4) Check IP-transport control sensor with cable and handling board and replace defective parts.
6170 | WRONG_STATE_LOW_LEVEL_STATE_OBSOLETE | Error |  |  |  | The received event cannot be handled in the actual state of the low level state machine (passive class). | The received event cannot be handled in the actual state of the low level state machine of the Drawer (passive class). Reason: A wrong state has been set. Cure: Check the corresponding SW.
6171 | LED_CIRCUIT_ERASURE_UNIT_NOT_OPERATIONAL | Error | 137 | Failure at erasure module. Please restart digitizer. If error remains, please contact service. If an image has been scanned, the image plate must be erased again after the repair of the digitizer. |  | One of the LED circuits of the erasure unit is not operational. | Meaning: One of the LED circuits of the erasure unit is not operational, i.e. the corresponding self test failed. Reason: 1) Erasure unit defective. 2) Handling control board defective. Cure: 1) Check/replace the erasure unit. 2) Check/replace the handling control board. | IPHM_ErasureUnitLEDError | CIR362 | On or more LED's of the Erasure Unit defective. Meaning: One of the LED circuits of the erasure unit is not operational, i.e. the corresponding self test failed. The Digitizer is out of operation and requires service support. Reason: 1) Erasure unit defective. 2) Handling control board defective. Cure: 1) Check/replace the erasure unit. 2) Check/replace the handling control board.
6173 | WRONG_VALUE_STATE_ERASURE_UNIT | Error |  |  |  | The erasure unit status has a wrong value. | Meaning: The erasure unit status has not the expected value (enabled or disabled). Reason: 1) voltage supply cable to erasure unit defective. 2) Erasure unit defective. 3) Handling control board defective. Cure: 1. Check the voltage supply connection and the cabling between erasure unit and handling control board and replace if defective. 2) Check/replace the erasure unit. 3) Check/replace the handling control board. | IPHM_ErasureUnitStatusError | CIR361 | Erasure unit wrong status. Meaning: The erasure unit status has not the expected value (enabled or disabled). The Digitizer is out of operation and requires service support. Reason: 1) voltage supply cable to erasure unit defective. 2) Erasure unit defective. 3) Handling control board defective. Cure: 1) Check the voltage supply connection and the cabling between erasure unit and handling control board and replace if defective. 2) Check/replace the erasure unit. 3) Check/replace the handling control board.
6178 | WRONG_COMMAND | Error |  |  |  | The received command is unknown and cannot be handled by the CaIHC. | Meaning: The received command is unknown and cannot be handled by the CaIHC. Reason: Software error. A wrong event was sent to the CaIHC. Cure: Software bugfix.
6185 | CALL_BACK_POINTER_ZERO | Error |  |  |  | The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. | Meaning: The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. Reason: Software error Cure: Software bugfix.
6187 | WRONG_VALUE_IRUN_ADJUSTABLE_DRAWER | Error |  |  |  | The adjustable run current value of the Drawer read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable run current value of the Drawer read out of the eLabel is not within the required range. Therefore the default value is used. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Do stall calibration of Drawer drive unit.
6189 | WRONG_CHC_PROTVER | Error |  |  |  | Protocol version of CHC is wrong | Meaning: Incompatible SW components (protocol version of CassetteHandlingControl CHC is wrong). Reason: SW error. Cure: Reinstall Software.
6190 | WRONG_SSC_PROTVER | Error |  |  |  | Protocol version of SSC is wrong | Meaning: Incompatible SW components (protocol version of SlowscanControl SSC is wrong). Reason: Different SW versions are installed on Handling Control board and on Slowscan board. Cure:Reinstall Software.
6191 | CAN_ADAPTER_ERROR | Error |  |  |  | fatal error in CAN adapter layer | Meaning: Fatal error in CAN communication layer. Reason: CAN error or SW Error. Cure: Check CAN connections; if OK, reinstall software.
6192 | LOGGING_ADAPTER_ERROR | Error |  |  |  | fatal error in logging adapter layer | Meaning: Fatal error in logger layer. Reason: CAN error or SW Error. Cure: Check CAN connections; if OK, reinstall software.
6200 | LIGHT_BARRIER_FOR_IP_DETECTION_NOT_CONNECTED | Error |  |  |  | IP-transport control sensor is not connected correctly. | Meaning: IP-transport control sensor(reflection light barrier) is not connected correctly. Reason: 1) Sensor not connected. 2) Sensor, cable or handling board defective. Cure: 1) Reconnect sensor. 2) Check the sensor, the corresponding cabling and the handling board, and replace defective parts. | IPHM_LBIPDetectionConnectionError | CIR422 | Contact or cable between Handling Control Board and Roller IP Sensor problem. Meaning: Connection test between the Handling Control Board and the Roller IP Sensor (B2) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at data cable between Handling Control Board and Roller IP Sensor (B2) 2) Contact problem or defect at connector on the Roller IP Sensor (B2) 3) Contact problem (connector P18) or defect on the Handling Control Board. Cure: 1) Check the data cable between Handling Control Board and Roller IP Sensor (B2). If required replace the data cable, refer to the Service Manual for order and replacement 2) Check connector on the Roller IP Sensor (B2). If required replace the Roller IP Sensor, refer to the Service Manual for order and replacement 3) Check connector P18 on the Handling Control Board 4) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
6201 | SENSOR_SLOWSCANLIFT_REFERENCE_NOT_CONNECTED | Error |  |  |  | The sensor for Slowscan Lift Reference position is not connected correctly. | Meaning: The sensor for Slowscan Lift Reference position is not connected correctly. Reason: 1) Sensor not connected. 2) Sensor, cable or handling board defective. Cure: 1) Reconnect sensor. 2) Check the sensor, the corresponding cabling and the handling board, and replace defective parts. | IPHM_LBSSLReferenceConnectionError | CIR267 | Contact or cable to SENSOR B426 (scan roller lift position) problem Meaning: Selftest of sensor B426 (scan roller lift position) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at cables between Handling Control Board, Adapter Board (B430) and Scan Roller Lift Position Sensor (B426) 2) Contact problem or defect of the Scan Roller Lift Position Sensor (B426) 3) Contact problem or defect of the Adapter Board (B430) 4) Contact problem (connector P14) or defect on the Handling Control Board Cure: 1) Check the data cables between Handling Control Board, Adapter Board (B430) and Scan Roller Lift Position Sensor (B426). If required exchange the data cables and the Adapter Board (B430), refer to the Service Manual for order and replacement 2) Check connectors on the Adapter Board (B430) 3) Check connectors on the Scan Roller Lift Position Sensor (B426). If required exchange the Scan Roller Lift Position Sensor (B426), for replacement contact your Regional Support Network for further assistance 4) Check connector P14 on the Handling Control Board 5) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
6207 | FEED_IP_OPEN_COVER_FAILED | Error |  |  |  | Feed the IP failed because the cover could not be opened. | Meaning: Feeding the IP to scanner failed because the cassette cover could not be opened. Reason: 1) The cassette is damaged. 2) Cassette opener drive is blocked by something or not moving at all. Cure: 1) Check the used cassette and replace if damaged. 2) Check cassette opener mechanics and electronics and replace defective parts. | CHM_OpenerOpenError | CIR524 | MOTOR M3 (cassette opener) moves with difficulty. Meaning: The drive (MOTOR M3) of the cassette opener does not reach the end position 'cassette flap open'. The Digitizer is out of operation and requires service support. Reason: 1) Cassette flap can not be opened due to a defect at the cassette 2) Drive belt or the belt pulley is not proper fixed or defect 3) Any obstacle in the travel range of the cassette opener drive or defects at the gear rack 4) Sluggishness of the cassette opener drive Cure: 1) Restart the Digitizer and check, if the problem persists 2) If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual. 3) Check the cassette flap for any defects 4) Check the drive belt, the belt pulley on the motor shaft and the gear rack for any defects 5) Check the drive mechanism of Cassette Opener for proper function and ease of movement 6) Perform the 'Diagnose Cycle' in the Service User Interface, refer to the Service Manual 7) Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
6208 | WRONG_UIC_PROTVER | Error |  |  |  | Protocol version of UIC is wrong | Meaning: Incompatible SW components (protocol version of UserInterface UIC is wrong). Reason: SW error. Cure:Reinstall Software.
6209 | WRONG_VALUE_IRUN_ADJUSTABLE_LOCKER | Error |  |  |  | The adjustable run current value of the Locker read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable run current value of the Locker read out of the eLabel is not within the required range. Therefore the default value is used. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Do stall calibration of Locker drive unit.
6210 | WRONG_VALUE_IRUN_ADJUSTABLE_ROLLERCLAMPER | Error |  |  |  | The adjustable run current value of the RollerClamper read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable run current value of the RollerClamper read out of the eLabel is not within the required range. Therefore the default value is used. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Do stall calibration of Roller Clamper drive unit.
6211 | WRONG_VALUE_IRUN_ADJUSTABLE_ROLLER_OBSOLETE | Error |  |  |  | The adjustable run current value of the Roller read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable run current value of the Roller read out of the eLabel is not within the required range. Therefore the default value is used. The wrong value is given as parameter. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Check the eLabel access (SW and HW).
6212 | WRONG_VALUE_IRUN_ADJUSTABLE_FIXER | Error |  |  |  | The adjustable run current value of the Fixer read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable run current value of the Fixer read out of the eLabel is not within the required range. Therefore the default value is used. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Do stall calibration of Fixer drive unit.
6214 | ERASURE_UNIT_POWER_CABLE | Error | 137 | Failure at erasure module. Please restart digitizer. If error remains, please contact service. If an image has been scanned, the image plate must be erased again after the repair of the digitizer. |  | The power cable to the erasure unit is not connected or the power supply is not working. | Meaning: The power cable to the erasure unit is not connected. Reason: 1) Power cable to erasur unit is unplugged. 2) erasure unit defective. 3) Power supply defective. 4) handling control board defective. Cure: 1) Check the cabling between power supply and erasure unit. Reconnect or replace. 2. Check the erasure unit and replace if defective. 3) Check/replace power supply 4) Check/replace handling control board. | IPHM_ErasureUnitPowerCableError | CIR363 | Supply voltage +24V on Erasure Unit low or missing Meaning: Test of the supply voltage (+24V) for erasure unit failed on the Handling Control Board. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect of the power cable between Flat Connector Strip (XK3) and Erasure LED Control Board (B424) 2) Contact problem (connector P2 or P3) or defect on the Erasure LED Control Board (B424) 3) Contact problem or defect of the cable between Erasure Unit and Handling Control Board 4) Contact problem (connector P16) or defect on the Handling Control Board Cure: 1) Check the supply voltage (+24V) at connector P3 on the Erasure LED Control Board (B424). If supply voltage (+24V) is not present, check power cable between Flat Connector Strip (XK3) and Erasure LED Control Board (B424). 2) Check connector P2 on the Erasure LED Control Board (B424) 3) Check the data cable between Erasure Unit and Handling Control Board. If required replace the data cable, refer to the Service Manual for order and replacement 4) Check connector P16 on the Handling Control Board 5) Exchange the Erasure Unit, refer to the Service Manual for order and replacement 6) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
6215 | ADC_CONVERSION_TIMEOUT | Error |  |  |  | The timeout for the A/D conversion exceeded. | Meaning: The A/D (analog to digital) conversion was not finished in the given time. It is not possible to measure the voltage. Reason: The A/D converter on the handling control board is defective. Cure: Replace handling control board.
6216 | DRAWER_TO_REFPOS | Error |  |  |  | Movement drawer drive to reference stopper failed. | Meaning: The Drawer unit movement to the reference stopper was not finished successfully. Reason: 1) Stall detection parameters are wrong. 2) The drawer unit is blocked by something, is running stiff or not moving at all. Cure: 1) Do stall calibration of Drawer unit. 2) Check Drawer mechanics and repair defective parts. 3) Check Drawer motor and cables and replace defective parts. | IPHM_DrawerRefPosError | CIR675 | MOTOR M4 (drawer unit) does not detect stop position. Meaning: The drive of the drawer unit does not detect the reference position (stopper). The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley is not proper fixed or defect 2) Any obstacle in the travel range of the drawer unit drive, stopper not proper fixed 3) Sluggishness of the drawer unit drive 4) Infrequent contact problems of cable or MOTOR M3 5) STALL-Calibration line out of range. Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the drive belt and the belt pulley on the motor shaft for any defects 3) Check it the rubber buffer at the stop position is proper fixed 4) Check the drive mechanism of DRAWER UNIT for proper function and ease of movement 5) Perform the 'Endurance Run Cycle' (e.g. 20 cycles) in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 6) Perform a STALL-Calibration of MOTOR M4 (drawer unit), refer to the Service Manual 7) If the problem remains, exchange the DRAWER UNIT, refer to the Service Manual for order and replacement.
6217 | WRONG_VALUE_CASS_SLOT_WIDTH_FIXER | Error |  |  |  | The calibrated cassSlotWidth value of the Fixer read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The calibrated cassSlotWidth value of the Fixer read out of the eLabel is not within the required range. Therefore the default value is used. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Do adjustment of cassette slot width or replace cassette unit.
6218 | WRONG_VALUE_REPETITION_PRECLAMP_FIXER | Error |  |  |  | The adjustable pre-clamp repetition value of the Fixer read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable for pre-clamp repetition value of the Fixer read out of the eLabel is not within the required range. Therefore the default value is used. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Do stall calibration of Fixer drive unit.
6219 | TAKE_IP_OUTCASSETTE_LOCKER_FAILED | Error |  |  |  | Take IP out of cassette failed because Locker drive error. | Meaning: Take IP out of cassette failed because of Locker drive error. Reason: One of movements unlock cover or unlock IP could not performed correctly. Cure: 1) Do stall calibration of Locker drive. 2) Check locker drive mechanics and electronics for proper work and replace defective part(s). 3) Check used cassette for damage and proper functionality. | CHM_LockerLockIPError | CIR564 | Drive of MOTOR M17 (cassette unlatch assembly) moves with difficulties. Meaning: The drive (MOTOR M17) of the cassette unlatch assembly does not reach the end position 'unlock pin pushed in'. The Digitizer is out of operation and requires service support. Reason: 1) Cassette unlock pin can not push in the cassette unlocker due to a defect at the cassette 2) Cam disk on the motor shaft or pin on the cam disk not proper fixed or defekt 3) Sluggishness of the cassette unlatch drive, e.g. guide rail get jammed 4) Infrequent contact problems of cable or MOTOR M17 Cure: 1) Restart the Digitizer and check, if the problem persists 2) If the cassette is still clamped, unclamp the cassette manually and remove it 3) Check the cassette unlocker for proper function and ease of movement 4) Check the cam shaft with the pin for proper fixation and function 5) Check the drive mechanism of the CASSETTE UNLATCH ASSEMBLY for proper function and ease of movement 6) Perform the 'Test Cycle' in the Service User Interface with different cassettes, refer to the Service Manual. If the error is related to specific cassettes, replace the affected cassettes. 7) Perform the 'Diagnose Cycle' in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 8) Exchange the CASSETTE UNLATCH ASSEMBLY, refer to the Service Manual for order and replacement.
6220 | TEMPERATURE_ERASURE_UNIT_TOO_HIGH | Error | 135 | The temperature of the erasure unit is too high. Please wait until the cooling unit reduces the temperature. |  | The temperature of the erasure unit is too high. | Meaning: The temperature of the erasure unit is too high. Reason: -- Cure: Wait until the temperature is again inside valid range.
6221 | ROLLER_RELEASE_TRAY | Error |  |  |  | Release tray movement of Roller drive failed. | Meaning: Release tray position not reached by Roller drive. Reason: Wrong movement parameters. Cure: Restart digitizer. If error persists, reinstall software. | IPHM_RollerMovementError | CIR454 | MOTOR M5 (transport roller) movement drive parameter out of range. Meaning: Movement parameter to drive the transport roller (MOTOR M5) are out of range. The Digitizer is out of operation and requires service support. . Reason: 1) Parameter, which defines the movement of the transport roller (MOTOR M5) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
6222 | ROLLER_PREP_FEED | Error |  |  |  | Prepare feed movement of Roller drive failed. | Meaning:Prepare feed position not reached by Roller drive. Reason: Wrong movement parameters. Cure: Restart digitizer. If error persists, reinstall software.defective parts. | IPHM_RollerMovementError | CIR454 | MOTOR M5 (transport roller) movement drive parameter out of range. Meaning: Movement parameter to drive the transport roller (MOTOR M5) are out of range. The Digitizer is out of operation and requires service support. . Reason: 1) Parameter, which defines the movement of the transport roller (MOTOR M5) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
6223 | ROLLER_INTO_SCANNER | Error |  |  |  | Into Scanner movement of Roller drive failed. | Meaning:Into Scanner position not reached by Roller drive. Reason: Wrong movement parameters. Cure: Restart digitizer. If error persists, reinstall software.defective parts.defective parts. | IPHM_RollerMovementError | CIR454 | MOTOR M5 (transport roller) movement drive parameter out of range. Meaning: Movement parameter to drive the transport roller (MOTOR M5) are out of range. The Digitizer is out of operation and requires service support. . Reason: 1) Parameter, which defines the movement of the transport roller (MOTOR M5) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
6224 | ROLLER_INTO_SCANNER_COMPLETE | Error |  |  |  | Into Scanner complete movement of Roller drive failed. | Meaning:Into Scanner complete position not reached by Roller drive. Reason:Wrong movement parameters. Cure: Restart digitizer. If error persists, reinstall software.defective parts. | IPHM_RollerMovementError | CIR454 | MOTOR M5 (transport roller) movement drive parameter out of range. Meaning: Movement parameter to drive the transport roller (MOTOR M5) are out of range. The Digitizer is out of operation and requires service support. . Reason: 1) Parameter, which defines the movement of the transport roller (MOTOR M5) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
6225 | DRAWER_PREP_FEED | Error |  |  |  | Prepare feed movement of Drawer drive failed. | Meaning:Prepare feed position not reached by Drawer drive Reason: wrong movement parameters. Cure: 1) Reset digitizer. 2) If error persists, reinstall software. | IPHM_DrawerMovementError
6226 | ROLLERCLAMPER_CLAMP_LIGHT | Error |  |  |  | Clamp light movement of RollerClamper drive failed. | Meaning:Clamp light position not reached by RollerClamper drive. Reason:wrong movement parameters. Cure: 1) Reset digitizer. 2) If error persists, reinstall software. | IPHM_RollClamperMovementError | CIR608 | MOTOR M6 (roller clamping) movement drive parameter out of range Meaning: Movement parameter to drive the roller clamping (MOTOR M6) to position 'IP slightly clamped' are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the roller clamping (MOTOR M6) to position 'IP slightly clamped' are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance
6227 | ROLLERCLAMPER_UNCLAMP_IP_SLOWLY | Error |  |  |  | Unclamp IP slowly movement of RollerClamper drive failed. | Meaning:Unclamp IP position not reached by RollerClamper drive. Reason: 1) Stall detection parameters are wrong. 2) The Roller Clamper drive is blocked by something or not moving at all. Cure: 1) Do stall calibration of Roller Clamper unit. 2)Check Roller Clamper drive mechanics and electronics for proper work and replace defective parts. | IPHM_RollClamperUnclampError | CIR605 | MOTOR M6 (roller clamping) does not detect stop position Meaning: The drive (MOTOR M6) of the roller clamping does not reach the reference position 'roller unclamped'. The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley on the motor shaft is not proper fixed or defect 2) Gear wheels or gear racks on the left and right side of the drive shaft are not proper fixed or broken 3) Any obstacle in the travel range of the drive and lifting mechanism of the roller clamping 4) Gear wheels at the left and right side on the drive shaft are meshing the gear racks left and right with an offset (not in sync) 5) Leaf spring at stop position of the roller lift mechanism not in position or broken 6) Infrequent contact problems of cable or MOTOR M6 7) STALL-Calibration line out of range. Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the drive belt or the belt pulley on the motor shaft for proper fixation and function 3) Check the gear wheels and gear racks on the left and right side of the drive shaft for proper fixation and function 4) Check, if the gear wheel drive on the left and right side are in sync to lift the rollers on both sides in the same way 5) Check the drive and lifting mechanism of the roller clamping for proper function and ease of movement 6) Check the leaf spring at stop position of the roller lift mechanism 7) Perform the 'Endurance Run Cycle' (e.g. 20 cycles) in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 8) Perform a STALL-Calibration of MOTOR M6 (roller clamping), refer to the Service Manual 9) For replacement of the roller clamping contact your Regional Support Network for further assistance.
6244 | WARN_CLAMP_CASS_INVALID_WIDTH | Warning |  |  |  | Warning: Clamp Cassette Movement detected invalid Cassette Width | Reason: The clamped Cassette has a invalid width. Meaning: 1) Cassette wrong inserted. 2) Cassette damaged. 3) Cassette fixer drive not working correctly. 4) Cassette slot width parameter wrong adjusted. Cure: 1) Retry cassette input. 2) Check cassette for damage and replace. 3) Check cassette fixer mechanics and electronics for proper work and replace defective part(s). 4) Redo adjustment of cassette slot width parameter or replace cassette unit.
7168 | CAS_LATENCY_NOT_SUPPORTED | Error |  |  |  | FPGA initialization failed because CAS_LATENCY is not supported | Meaning: FpgaAccess: FPGA initialization failed because CAS_LATENCY is not supported. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7169 | BURST_LENGTH_NOT_SUPPORTED | Error |  |  |  | FPGA initialization failed because BURST_LENGTH is not supported | Meaning: FpgaAccess: FPGA initialization failed because BURST_LENGTH is not supported. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7170 | SDRAM_CONTROLLER_START_FAILED | Error |  |  |  | FPGA initialization failed because SDRAM controller do not start | Meaning: FpgaAccess: FPGA initialization failed because SDRAM controller do not start. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7171 | TEST_FPGA_NOT_LOADED | Error |  |  |  | FPGA test failed FPGA is not loaded yet | Meaning: FpgaAccess: FPGA test failed FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_NotNeeded_ForServiceId
7172 | PATTGEN_FPGA_NOT_LOADED | Error |  |  |  | FPGA pattern generation failed because FPGA is not loaded yet | Meaning: FpgaAccess: FPGA pattern generation failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_NotNeeded_ForServiceId
7173 | PATTGEN_FLUSH_FIFO_UNDERRUN | Error |  |  |  | FPGA pattern generation failed because flush FIFO underrun occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO underrun occured. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7174 | PATTGEN_FLUSH_FIFO_OVERFLOW | Error |  |  |  | FPGA pattern generation failed because flush FIFO overflow occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO overflow occured. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7175 | PATTGEN_FLUSH_FIFO_DRAIN_NOK | Error |  |  |  | FPGA pattern generation failed because flush FIFO cannot be drained | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7176 | FLUSH_FIFO_DRAIN_FAILED | Error |  |  |  | flush FIFO cannot be drained; HW problem, fatal error | Meaning: FpgaAccess: flush FIFO cannot be drained; HW problem, fatal error. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7177 | FPGA_TEST_FAILED | Error |  |  |  | FPGA test failed (one of the RAMs was not correctly accessible), fatal error | Meaning: FpgaAccess: FPGA test failed (one of the RAMs was not correctly accessible). Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7178 | TIMEOUT_FILLING_SCAN_FIFO | Error |  |  |  | scan fifo write pump did not fill scan fifo with SDRAM data, fatal error | Meaning: FpgaAccess: scan fifo write pump did not fill scan fifo with SDRAM data. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7179 | UNKNOWN_FLUSH_FIFO_INPUT | Error |  |  |  | unknown input for flush fifo selected, fatal software error | Meaning: FpgaAccess: unknown input for flush fifo selected. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7180 | FILLSDRAM_FPGA_NOT_LOADED | Error |  |  |  | filling the SDRAM failed because FPGA is not loaded yet | Meaning: FpgaAccess: filling the SDRAM failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_NotNeeded_ForServiceId
7181 | FILLSDRAM_IMAGE_TOO_BIG | Error |  |  |  | filling the SDRAM failed because desired image is too big | Meaning: FpgaAccess: filling the SDRAM failed because desired image is too big. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7182 | PATTGEN_SDRAM_SIZE_UNKNOWN | Error |  |  |  | FPGA pattern generation failed because size of SDRAM is not determined yet (call initFpga first!) | Meaning: FpgaAccess: FPGA pattern generation failed because size of SDRAM is not determined yet. Reason: SW error (call initFpga first!). Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7183 | PATTGEN_EXCEEDS_SDRAM | Error |  |  |  | FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM | Meaning: FpgaAccess: FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM. Reason: SW error. Cure: Install new software; inform R&D team. | COM_NotNeeded_ForServiceId
7184 | LOAD_FPGA_CONFIGURATION | Error |  |  |  | Can't load FPGA configuration, fatal error | Meaning: FpgaAccess: Can't load FPGA configuration. Reason: Hardware defect or software mismatch. Cure: Upload PMI software and FPGA configuration; if still NOK, replace PMI board | COM_NotNeeded_ForServiceId
7185 | LOAD_FPGA_HARDWARE | Error |  |  |  | Load FPGA: can't set FPGA into configuration mode, fatal error | Meaning: FpgaAccess: Load FPGA can't set FPGA into configuration mode. Reason: HW error. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
7186 | STORED_FPGA_CONFIGURATION | Error |  |  |  | No valid FPGA version found on PMI | Meaning: FPGA version doesn't fit to hard- or software version. Reason: SW or HW error. Cure: Update software including FPGA configuration; if still NOK, replace PMI board | COM_NotNeeded_ForServiceId
7187 | READ_PMI_CPU | Error |  |  |  | Can't read from FPGA CPU | Meaning: Can't read from FPGA CPU. Reason: Hardware error on PMI. Cure: Retry switch digitizer off/on; if still NOK, replace PMI board. | COM_NotNeeded_ForServiceId
7188 | CAL_ERROR_FLAG | Error |  |  |  | PMI CPU Error flag is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Polygon or Laser or PMI board defective. Cure: Retry switch digitizer off/on, replace lens assembly CR30X, replace PMI board. | COM_NotNeeded_ForServiceId
7189 | CAL_EOL | Error |  |  |  | EOL error counter is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Optical path is partly interrupted, calibration board is defective. Cure: Clean scan line with antistatic brush, check mounting of light guide, switch digitizer off/on, replace calibration plate | COM_GeoCalibrationErr
7190 | CAL_CAL_SENSORS | Error |  |  |  | Calibration sensor error counter is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Optical path is partly interrupted, calibration board is defective. Cure: Clean scan line with antistatic brush, switch digitizer off/on, replace calibration plate | COM_GeoCalibrationErr
7680 | IMAGE_SERVER_CREATION_FAILED | Error |  |  |  | ImageTransferControl cannot create image data server; fatal! | Meaning: ImageTransferControl cannot create image data server. Reason: SW error. Cure: Install new software. | COM_ImageTransferControlErr
7681 | INVALID_MAX_PACKET_SIZE | Error |  |  |  | ImageTransferControl received wrong FW packet size from PC; fatal! | Meaning: ImageTransferControl received wrong FW packet size from PC. Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7682 | INVALID_FW_SPEED_CODE | Error |  |  |  | ImageTransferControl determined invalid FW speed; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&D team in any case. | COM_ImageTransferControlErr
7683 | BLOCK_TRANSFER_TIMEOUT_EXCEEDED | Error |  |  |  | ImageTransferControl: block transfer takes too long; either FW cable was disconnected during the block transfer or a fatal HW error occured! | Meaning: ImageTransferControl: block transfer takes too long. Reason: The transmission of an image data block takes too long. Cure: Check if FW cable is connected correctly; if still NOK, replace FW cable; if still NOK, replace PMI board; if still NOK, replace FW-board on PC. | COM_ImageTransferControlErr
7684 | INVALID_PC_NODE_ID | Error |  |  |  | ImageTransferControl: block transfer failed because PC's node ID is invalid; fatal error! | Meaning: ImageTransferControl: block transfer failed because PC's node ID is invalid. Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7685 | SCAN_FIFO_OVERFLOW | Error |  |  |  | ImageTransferControl: scan fifo overflow during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo overflow during block transfer. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7686 | SCAN_FIFO_UNDERRUN | Error |  |  |  | ImageTransferControl: scan fifo underrun during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo underrun during block transfer. Reason: HW error. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI board and PMT/lens assembly; if OK, replace PMI board; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly | COM_ImageTransferControlErr
7687 | FPGA_BUS_ERROR | Error |  |  |  | ImageTransferControl: FPGA bus error occured; fatal! | Meaning: ImageTransferControl: FPGA bus error occured. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7688 | INVALID_NR_OF_REQUESTED_BYTES | Error |  |  |  | ImageTransferControl: invalid number of requested bytes ( <=0); fatal SW bug! | Meaning: ImageTransferControl: invalid number of requested bytes ( <=0). Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7689 | TIMEOUT_GET_PACKET_SIZE | Error |  |  |  | ImageTransferControl: timeout: cannot determine used FW packet size; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&D team in any case. | COM_ImageTransferControlErr
7690 | TIMEOUT_DATA_NOT_AVAILABLE | Error |  |  |  | ImageTransferControl: timeout: scanner didn't store enough data in SDRAM yet to start block transfer; wait, then try again! | Meaning: ImageTransferControl: timeout. Reason: Scanner didn't store enough data in SDRAM yet to start block transfer. Cure: No problem, PC software has to wait and try again. | COM_ImageTransferControlErr
7691 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump. Reason: HW error (DM (DataMover) pump or LLC (LinkLayerController) failed). Cure: Replace PMI board. | COM_ImageTransferControlErr
7692 | RECOVER_PACKET_FAILED | Error |  |  |  | ImageTransferControl: recovering the FW packet failed; fatal error! | Meaning: ImageTransferControl: recovering the FW packet failed. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7693 | TOO_MANY_PACKETS_PER_BLOCK | Error |  |  |  | ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block); request smaller block! | Meaning: ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block). Reason: SW error (PC must request smaller block!). Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7694 | FLUSH_FIFO_DRAIN_FAILED | Error |  |  |  | flush FIFO cannot be drained; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7695 | PIXELVIEWER_FLUSH_FIFO_OVERFLOW | Error |  |  |  | flush FIFO overflow occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO overflow occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7696 | PIXELVIEWER_FLUSH_FIFO_UNDERRUN | Error |  |  |  | flush FIFO underrun occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO underrun occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7697 | PIXELVIEWER_CRC1_AND_2_NOT_EQUAL | Error |  |  |  | after scanning in PixelViewer mode, the CRC1 and 2 are not equal; HW problem, fatal error! | Meaning: ImageTransferControl: after scanning in PixelViewer mode, the CRC1 and 2 are not equal. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7698 | PIXELVIEWER_START_NEXT_SCAN_FAILED | Error |  |  |  | starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running; HW failure (e.g. BOL missing) or SW synchronization problem, fatal error! | Meaning: ImageTransferControl: starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running. Reason: HW failure (e.g. BOL missing) or SW synchronization problem. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case. | COM_ImageTransferControlErr
7699 | DM_ERROR_HANDLER_SETUP_FAILED | Error |  |  |  | registering the callback function for DM errors failed | Meaning: ImageTransferControl: registering the callback function for DM errors failed. Reason: SW error. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7700 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_START | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7701 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_OVFL | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7702 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_URUN | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7703 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_RETRANS | Error |  |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission. Reason: HW error. Cure: Replace PMI board. | COM_ImageTransferControlErr
7704 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER | Error |  |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost. | COM_ImageTransferControlErr
7705 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_START | Error |  |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. Actual transfer was not started yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost. | COM_ImageTransferControlErr
7706 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_PUMP | Error |  |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. DM pump was not enabled yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image transmission automatically again, image is not lost. | COM_ImageTransferControlErr
7707 | IMAGE_BUFFER_CREATION_FAILED | Error |  |  |  | ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. | Meaning: ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. Reason: not enough memory resources (probably software bug) Cure: Install new software. | COM_ImageTransferControlErr
7708 | MORE_DATA_REQUESTED_THAN_SCANNED | Error |  |  |  | ImageTransferUDPControl: more data requested than scanned | Meaning: ImageTransferUDPControl: more data requested than scanned. Reason: Index out of range. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7709 | SERVER_DISCONNECTED_DURING_BLOCK_TRANSFER | Error |  |  |  | ImageTransferUDPControl: block transfer stopped because server was disconnected! | Meaning: ImageTransferUDPControl: block transfer stopped because server was disconnected! Reason: Cable was disconnected during transmission of image data to PC. Cure: Connect cable again; PC will request the image trasnmission automatically again, image is not lost. | COM_ImageTransferControlErr
7710 | NO_DATA_IN_DRAM_TIMEOUT | Error |  |  |  | ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. | Meaning: ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. Reason: Copying data into DRAM takes too long, or, no data anymore in FPGA. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7711 | DMA_NOT_PREPARED_TIMEOUT | Error |  |  |  | ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. | Meaning: ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. Reason: Timing problem. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7712 | PATGEN_ENABLE_PROBLEM | Error |  |  |  | ImageTransferUDPControl: The Pattern Generator cannot be enabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be enabled. Reason: Could be software or hardware problem. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7713 | PATGEN_DISABLE_PROBLEM | Error |  |  |  | ImageTransferUDPControl: The Pattern Generator cannot be disabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be disabled. Reason: Could be software or hardware problem. Cure: Install new software; inform R&D team. | COM_ImageTransferControlErr
7714 | IMAGE_TRANSFER_CANCELLED | Error |  |  |  | ImageTransferUDPControl: the image transfer was cancelled | Meaning: ImageTransferUDPControl: the image transfer was cancelled. | COM_ImageTransferControlErr
7715 | DM_ERROR | Error |  |  |  | ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error | Meaning: ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error. Reason: A DM error occurred during the transmission of an image block. Cure: The PC will automatically request the image block again (up to 10 times); if all repetitions fail an image transer error is displayed; in the latter case reboot the PC, the PC will start the image transmission automatically again, the image is not lost as long as the digitizer is not switched off. | COM_ImageTransferControlErr
8192 | LOST_MESSAGES | Error |  |  |  | Lost logmessages
8193 | SAVE_DIGITIZERSETTINGS | Error |  |  |  | Could not save digitizer settings
10752 | INCIDENCE_OF_LIGHT | Error |  |  |  | Incidence of light --> adj.PMM/PMT failed ; fatal! | Meaning: It is too bright to adjust the Photo Multiplier. Reason: Housing of digitizer is open. Cure: Close the digitizer housing.
10753 | PMM_ADJUST | Error |  |  |  | PMM/PMT adjust failed | Meaning: Photo multiplier malfunction. Reason: - Cure: Switch digitizer off/on and try a scan, if failed again replace PMT + LIGHT COLLECTOR.
10754 | GEO_CALIBRATION | Error |  |  |  | Geometric and jitter calibration failed | Meaning: Geometric and jitter calibration failed. Reason: Optic module, calibration board, sensor path misadjusted. Cure: Check proper seating of optic module, clean scan path with dust brush.
10755 | START_POLYGON | Error |  |  |  | Start polygon failed | Meaning: Polygon locked signal is not set. Reason: - Cable broken - Polygon(Optic Mudule) defective - PMI board defective - PMI POWER DISTRIBUTION board defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check the cabling of the following boards: - PMI BOARD - Optic Module - DISTRIBUTOR BOARD - PMI BOARD 3. Exchange DISTRIBUTOR BOARD, refer to the Service Manual for order and replacement 4. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement 5. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement
10756 | ADC_ADJUST | Error |  |  |  | ADC adjustment out of spec | Meaning: ADC offset adjust out of range. Reason: ADC defective or PMM/PMT offset voltage out of range. Cure: Switch digitizer off/on, replace PMI board, replace PMT + LIGHT COLLECTOR.
10757 | START_SCAN | Error |  |  |  | Start scan failed | Meaning: Start scan failed, because scan is already running Reason: Software error Cure: Update software or inform RDE
10758 | POLYGON_LOCKED | Error |  |  |  | Polygon get not locked | Meaning: Polygon can't reach the requested speed. Reason: Polygon defective. Cure: Replace optic unit.
10759 | PREPARE_SCAN | Error |  |  |  | Prepare Scan failed
10760 | TOO_LESS_DATA | Error |  |  |  | Too less data stored in SDRAM --> SDRAM filled with RAMP | Meaning: Too less data stored in SDRAM --> SDRAM filled with stripes. Reason: Electrostatic discharge, error in begin or end of line signal (BOL,EOL). Cure: Check grounding of calibration board and grounding of front cover.
10761 | TEST_SCAN | Error |  |  |  | Test scan failed --> BOL or EOL out of range | Meaning: One or both of the outermost sensors of the CALIBRATION BOARD, which detects the begin (BOL) and the end (EOL) of the scan line, does not detect the expected amount of incident laser beam from the OPTIC MODULE. The Digitizer is out of operation and requires service support. Reason: One or both of the outermost sensors of the CALIBRATION BOARD does not detect the expected amount of incident laser beam because of: - The optical path of the laser beam is obstructed (e.g. dust cover, particles on the sensors, broken light collector, foreign body, ...) - Misalignment of the OPTIC MODULE - Defective of any of the following cables or connectors: - CALIBRATION BOARD - DISTRIBUTOR BOARD - DISTRIBUTOR BOARD - PMI BOARD - Defective DISTRIBUTOR BOARD - Defective CALIBRATION BOARD - Defective PMI-BOARD - Defective OPTIC MODULE Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check, that the optical path of the laser beam is not obstructed. If applicable remove the dust cover, refer to Service Bulletin No.37 3. Check the position of the laser beam with the new CALIBRATION BOARD or the LASER BEAM DETECTION TOOLKIT, refer to the Service Manual, adjustments and Service Bulletin No.28 4. Check the cabling of the following boards: - CALIBRATION BOARD - DISTRIBUTOR BOARD - DISTRIBUTOR BOARD - PMI BOARD 5. Exchange DISTRIBUTOR BOARD, refer to the Service Manual for order and replacement 6. Exchange the CALIBRATION BOARD, refer to the Service Manual for order and replacement 7. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement 8. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
10762 | PMM_REFCAL | Error |  |  |  | Calculation of REFCAL from IPF out of range | Meaning: The IP depended calculation of sensitivity factor results in a value out of range. Reason: Wrong PMM/PMT adjustment values stored in PMT eLabel. Cure: Replace PMT + LIGHT COLLECTOR.
10763 | PMM_AGING | Error |  |  |  | PMT sensitivity become critical --> reaches 94% = 750V | Meaning: The PMTs end is coming soon. Cure: Replace PMT + LIGHT COLLECTOR
10764 | PMM_FINE_ADJUST | Error |  |  |  | PMM/PMT fine adjust failed --> possibly bad image quality | Meaning: PMT sensitivity adjustment failed Reason: - PMI POWER DISTRIBUTION board defective - PMT defective or end of life Cure: 1. Replace PMI POWER DISTRIBUTION board 2. If error still persists replace PMT + LIGHT COLLECTOR
10765 | START_GEO_CALIBRATION | Error |  |  |  | Start geometric and jitter calibration failed
10766 | IPF_OUT_OF_RANGE | Error |  |  |  | IP sensitivity factor out of range | Meaning: IP is not compatible with digitizer Reason: IP does nit fit to this digitizer Cure: Use an IP, which is compatible to digitizer
10767 | TEST_SCAN_CAL | Error |  |  |  | Test scan failed: Not all CAL sensors detected --> defect or IP in digitizer | Meaning: Can't detect the expected amount of calibration signals. Reason: Optic path interrupted, IP in digitizer, broken light collector, cable PMI to calibration board defective, CAL-board defective Cure: Clean the scan line, check optic path, check cable PMI to calibration board, replace calibration board.
10768 | READ_PMM_ELABEL | Error |  |  |  | Can't read PMM settings from PMM eLabel |  | COM_PMMConnectionErr | CIR221 | Connection between PMI Board and PMT failed. Meaning: Reason: Cure:
10769 | READY_TO_SCAN_CAL_SENSORS | Error |  |  |  | Calibration sensor defective or two IP's in scanner | Meaning: Can't detect the expected amount of calibration signals. Reason: CAL-board defective or optic path interrupted (IP or dust). Cure: Clean the scan line, replace calibration board.
10770 | PMM_BLUE_LED | Error |  |  |  | Photo Multiplier Module (PMM/PMT), blue LED or power supply board defective | Meaning: Can't detect PMT blue LED. Reason: - PMT blue LED defective or not connected - PMI POWER DISTRIBUTION board defective - PMT defective Cure: 1. Check connection and cable from blue LED to PMT 2. Replace PMI POWER DISTRIBUTION board 3. If error still persists replace PMT + LIGHT COLLECTOR
10771 | FPGA_BUS | Error |  |  |  | FPGA BUS error detected | Meaning: FPGA BUS error occurred. Reason: unknown Cure: Restart digitizer, if error occurred again replace PMI board.
10772 | AUTOEOL_ADDED | Error |  |  |  | FPGA has added an EOL at end of EOL valid | Meaning: No EOL detected, FPGA adds one at end of EOL valid window. Reason: polygon speed out of range. Cure: Restart digitizer, replace optic module.
11780 | SM_DRIVER_FAILURE | Error |  |  |  | Procedure selftest of SMx terminated with DIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11781 | SM_GROUND_CURRENT_FAILURE | Error |  |  |  | Procedure selftest of SMx terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11782 | ADC_MEAS_TOT_FAILURE | Error |  |  |  | Selftest of SMx ended due to ADC measurement timeout failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11876 | WARN_INVALID_PAR | Error |  |  |  | Warning: Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
11877 | WARN_INVALID_ACCESS_FUNC | Error |  |  |  | Warning: Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
11878 | WARN_INVALID_POINTER | Error |  |  |  | Warning: Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
11879 | WARN_LOW_VOLT_MEAS_POS | Error |  |  |  | Warning: Test pos current, measured voltage value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at driver voltage of an active pos current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11880 | WARN_LOW_VOLT_MEAS_NEG | Error |  |  |  | Warning: Test neg current, measured voltage value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at driver voltage of an active neg current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11881 | WARN_LOW_SIN_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement sin value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at sin current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11882 | WARN_LOW_COS_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement cos value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at cos current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11883 | WARN_LOW_SIN_NEG_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement sin neg value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at sin negative current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11884 | WARN_LOW_COS_NEG_CURRENT_MEAS | Error |  |  |  | Warning: Test measurement cos neg value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at cos negative current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
12289 | FIXER_MOTOR_UNPLUGGED | Error |  |  |  | Cassette fixing motor(M1) unplugged. | Meaning: Cassette fixing motor(M1) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P7. 2) Replace motor M1. | CHM_FixerUnpluggedError | CIR582 | Contact or cable to MOTOR M1 (cassette fixing) problem. Meaning: Selftest of MOTOR M1 (cassette fixing) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M1 (cassette fixing) problem 2) MOTOR M1 (cassette fixing) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M1 (cassette fixing), re-connect cable at HANDLING CONTROL BOARD (P7) 2) Exchange the MOTOR M1 (cassette fixing), refer to the Service Manual for order and replacement.
12290 | FIXER_MOTOR_DEFECTIVE | Error |  |  |  | Cassette fixing motor(M1) defective. | Meaning: Cassette fixing motor(M1) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M1. 2) Replace Handling board. | CHM_FixerMotorError | CIR581 | MOTOR M1 (cassette fixing) defective. Meaning: Selftest of MOTOR M1 (cassette fixing) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in MOTOR M1 (cassette fixing) or motor cable. 2) Electrical short in the control for MOTOR M1 (cassette fixing) on the HANDLING CONTROL BOARD. Cure: 1) First exchange the MOTOR M1 (cassette fixing), refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12292 | FIXER_DRIVER | Error |  |  |  | Cassette fixing motor(M1), driver defective. | Meaning: Cassette fixing motor(M1) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | CHM_FixerDriverError | CIR583 | Motor control for MOTOR M1 (cassette fixing) is defective . Meaning: Selftest of MOTOR M1 (cassette fixing) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M1 (cassette fixing) on the HANDLING CONTROL BOARD is defective. Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12293 | FIXER_MOVEMENT | Error |  |  |  | Cassette fixing motor(M1), movement failed. | Meaning:Cassette fixing motor(M1) movement failed. Reason: 1) Wrong stall parameters. 2) Drive is blocked by something or runs stiff. 3) Drive does not move at all. 4) Steering problem by software or hardware. Cure: 1) Do stall calibration of Cassette fixing motor M1. 2)Check mechanics and replace defective part(s). 3) Check electronis (motor, cables, handling board) and replace defective part(s). 4) Reinstall software or replace Handling board. | CHM_FixerClampError | CIR584 | Drive of MOTOR M1 (cassette finxing) moves with difficulty. Meaning: The drive (MOTOR M1) of the cassette fixing does not reach the end position 'cassette clamped'. The Digitizer is out of operation and requires service support. Reason: 1) Coupling on the motor shaft is not proper fixed or defect 2) Towing bracket on the spindel bush is not proper fixed or broken 3) Infrequent contact problems of cable or MOTOR M1 4) STALL-Calibration line out of range Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the coupling on the motor shaft for proper fixation and function 4) Check the towing bracket on the spindel bush for any defects like stress cracks 3) Perform the 'Diagnose Cycle' in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 4) Perform a STALL-Calibration of MOTOR M1 (cassette fixing), refer to the Service Manual 5) If the problem remains, exchange the CASSETTE FIXATION UNIT, refer to the Service Manual for order and replacement.
12295 | LOCKER_MOTOR_UNPLUGGED | Error |  |  |  | Cassette unlatch / locker motor(M17) unplugged. | Meaning: Cassette unlatch / locker motor(M17) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P8. 2) Replace motor M17. | CHM_LockerUnpluggedError | CIR562 | Contact or cable to MOTOR M17 (cassette unlatch unit) problem. Meaning: Selftest of MOTOR M17 (cassette unlatch assembly) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M17 (cassette unlatch assembly) problem 2) MOTOR M17 (cassette unlatch assembly) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M17 (cassette unlatch assembly), re-connect cable at HANDLING CONTROL BOARD (P8) 2) Exchange the CASSETTE UNLATCH ASSEMBLY, refer to the Service Manual for order and replacement.
12296 | LOCKER_MOTOR_DEFECTIVE | Error |  |  |  | Cassette unlatch / locker motor(M17) defective. | Meaning: Cassette unlatch / locker motor(M17) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M17. 2) Replace Handling board. | CHM_LockerMotorError | CIR561 | MOTOR M17 (cassette unlatch assembly) defective. Meaning: Selftest of MOTOR M17 (cassette unlatch assembly) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in MOTOR M17 (cassette unlatch assembly) or motor cable. 2) Electrical short in the control for MOTOR M17 (cassette unlatch assembly) on the HANDLING CONTROL BOARD. Cure: 1) First exchange the CASSETTE UNLATCH ASSEMBLY, refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12298 | LOCKER_DRIVER | Error |  |  |  | Cassette unlatch / locker motor(M17), driver defective. | Meaning: Cassette unlatch / locker motor(M17) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | CHM_LockerDriverError | CIR563 | Motor control for MOTOR M17 (cassette unlatch assembly) is defective. Meaning: Selftest of MOTOR M17 (cassette unlatch assembly) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M17 (cassette unlatch assembly) on the HANDLING CONTROL BOARD is defective. Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12299 | LOCKER_MOVEMENT | Error |  |  |  | Cassette unlatch / locker motor(M17), movement failed. | Meaning:Cassette unlatch / locker motor(M17) movement failed. Reason: 1) Wrong stall parameters. 2) Drive is blocked by something or runs stiff. 3) Drive does not move at all. 4) Steering problem by software or hardware. Cure: 1) Do stall calibration of Cassette unlatch / locker motor M17. 2)Check mechanics and replace defective part(s). 3) Check electronis (motor, cables, handling board) and replace defective part(s). 4) Reinstall software or replace Handling board. | CHM_LockerLockIPError | CIR564 | Drive of MOTOR M17 (cassette unlatch assembly) moves with difficulties. Meaning: The drive (MOTOR M17) of the cassette unlatch assembly does not reach the end position 'unlock pin pushed in'. The Digitizer is out of operation and requires service support. Reason: 1) Cassette unlock pin can not push in the cassette unlocker due to a defect at the cassette 2) Cam disk on the motor shaft or pin on the cam disk not proper fixed or defekt 3) Sluggishness of the cassette unlatch drive, e.g. guide rail get jammed 4) Infrequent contact problems of cable or MOTOR M17 Cure: 1) Restart the Digitizer and check, if the problem persists 2) If the cassette is still clamped, unclamp the cassette manually and remove it 3) Check the cassette unlocker for proper function and ease of movement 4) Check the cam shaft with the pin for proper fixation and function 5) Check the drive mechanism of the CASSETTE UNLATCH ASSEMBLY for proper function and ease of movement 6) Perform the 'Test Cycle' in the Service User Interface with different cassettes, refer to the Service Manual. If the error is related to specific cassettes, replace the affected cassettes. 7) Perform the 'Diagnose Cycle' in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 8) Exchange the CASSETTE UNLATCH ASSEMBLY, refer to the Service Manual for order and replacement.
12301 | OPENER_MOTOR_UNPLUGGED | Error |  |  |  | Cassette opener motor(M3) unplugged. | Meaning: Cassette opener motor(M3) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P10. 2) Replace motor M3. | CHM_OpenerUnpluggedError | CIR522 | Contact or cable to MOTOR M3 (cassette opener) problem. Meaning: Selftest of MOTOR M3 (cassette opener) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M3 (cassette opener) problem 2) MOTOR M3 (cassette opener) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M3 (cassette opener), re-connect cable at HANDLING CONTROL BOARD (P10) 2) Exchange the MOTOR M3 (cassette opener), refer to the Service Manual for order and replacement.
12302 | OPENER_MOTOR_DEFECTIVE | Error |  |  |  | Cassette opener motor(M3) defective. | Meaning: Cassette opener motor(M3) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M3. 2) Replace Handling board. | CHM_OpenerMotorError | CIR521 | MOTOR M3 (cassette opener) defective. Meaning: Selftest of MOTOR M3 (cassette opener) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in MOTOR M3 (cassette opener) or motor cable. 2) Electrical short in the control for MOTOR M3 (cassette opener) on the HANDLING CONTROL BOARD. Cure: 1) First exchange the MOTOR M3 (cassette opener), refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12304 | OPENER_DRIVER | Error |  |  |  | Cassette opener motor(M3), driver defective. | Meaning: Cassette opener motor(M3) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | CHM_OpenerDriverError | CIR523 | Motor control for MOTOR M3 (cassette opener) is defective. Meaning: Selftest of MOTOR M3 (cassette opener) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M3 (cassette opener) on the HANDLING CONTROL BOARD is defective Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12305 | OPENER_MOVEMENT | Error |  |  |  | Cassette opener motor(M3), movement failed. | Meaning: Cassette Opener motor(M3) movement failed. Reason: 1) Drive is blocked by something or runs stiff. 2) Drive does not move at all. 3) Steering problem by software or hardware. Cure: 1)Check mechanics and replace defective part(s). 2) Check electronis (motor, cables, handling board) and replace defective part(s). 3) Reinstall software or replace Handling board. | CHM_OpenerOpenError | CIR524 | MOTOR M3 (cassette opener) moves with difficulty. Meaning: The drive (MOTOR M3) of the cassette opener does not reach the end position 'cassette flap open'. The Digitizer is out of operation and requires service support. Reason: 1) Cassette flap can not be opened due to a defect at the cassette 2) Drive belt or the belt pulley is not proper fixed or defect 3) Any obstacle in the travel range of the cassette opener drive or defects at the gear rack 4) Sluggishness of the cassette opener drive Cure: 1) Restart the Digitizer and check, if the problem persists 2) If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual. 3) Check the cassette flap for any defects 4) Check the drive belt, the belt pulley on the motor shaft and the gear rack for any defects 5) Check the drive mechanism of Cassette Opener for proper function and ease of movement 6) Perform the 'Diagnose Cycle' in the Service User Interface, refer to the Service Manual 7) Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
12306 | OPENER_MOVEMENT_STEP_DEVIATION_NEG | Error |  |  |  | Cassette opener motor(M3), movement failed with less steps moved. | Meaning: Cassette Opener motor(M3) movement failed with less steps moved. Reason: 1) Opener drive or mechanics damaged or runs stiff 2) Upper stopper not mounted correctly or fixed at worng position. Cure: 1) Check opener mechanics for any damage or blockade (fixing screws of upper stopper ..). 2) Replace Opener Unit. | CHM_Opener_StepDeviationNeg | CIR526 | MOTOR M3 (cassette opener) detected less steps. Meaning: The drive (MOTOR M3) of the cassette opener detected less steps between the upper and the lower stop position. The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley is not proper fixed or defect 2) Defects at the gear rack or gearwheels 3) Any obstacle in the travel range of the cassette opener drive, e.g. upper stopper with fixing screws not proper mounted 4) Sluggishness of the cassette opener drive. Cure: 1) Check the drive belt, the belt pulley on the motor shaft and the gear rack for any defects 2) Check, if the rubber cushions at the upper and lower stopper are propper fixed 3) Check, if the two fixing screws of the upper stopper does not touch the cassette opener bracket 4) Check the whole drive mechanism of Cassette Opener for proper function and smooth running. 5) Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
12307 | OPENER_MOVEMENT_STEP_DEVIATION_POS | Error |  |  |  | Cassette opener motor(M3), movement failed with too much steps moved. | Meaning: Cassette Opener motor(M3) movement failed with too much steps moved. Reason: 1) Upper or lower rubber damper mounted not correctly 2) Upper stopper not fixed correctly or at worng position. 3) Lower stopper not fixed correctly. Cure: 1) Check both rubber damper. 2) Check upper stopper if correct mounted. 3) Check Lower stopper if corect mounted. 4) Insert washer (max 1mm) between upper stopper and rubber damper. 5) Replace Opener Unit. | CHM_Opener_StepDeviationPos | CIR527 | MOTOR M3 (cassette opener) detected too much steps. Meaning: The drive (MOTOR M3) of the cassette opener detected too much steps between the upper and the lower stop position. The Digitizer is out of operation and requires service support. Reason: 1) Drive belt or the belt pulley is not proper fixed or defect 2) Rubber cushions at the upper or lower stopper are not proper fixed 3) Upper or lower stopper not proper fixed Cure: 1. Check the drive belt, the belt pulley on the motor shaft for any defects 2. Check, if the rubber cushions at the upper and lower stopper are proper fixed 3. Check, if the upper and lower stopper are proper mounted 4. Exchange the Cassette Opener Unit, refer to the Service Manual for order and replacement.
12308 | BOOSTER_39V_SELFTEST | Error |  |  |  | Selftest of SM-Booster 24V to 39V failed.. | Meaning: Selftest of SM-Booster 24V to 39V on CHC-Board failed. Reason: Booster Hardware on CHC-Board defective. Cure: Exchange the Handling Control Board | CHM_39VSMSupplyError | CIR413 | Supply voltage +39V (Stepper motors) on Handling Control Board low or missing. Meaning: Test of the supply voltage (+39V) for motors failed on the Handling Control Board. The Digitizer is out of operation and requires service support. Reason: 1) The DC/DC converter on the Handling Control Board, which provides the +39V supply voltage for motors, is defective Cure: 1) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
12801 | ROLLERCLAMPER_MOTOR_UNPLUGGED | Error |  |  |  | Roller Clamping motor(M6) unplugged. | Meaning: Roller Clamping motor(M6) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P12. 2) Replace motor M6. | IPHM_RollClamperUnpluggedError | CIR602 | Contact or cable to MOTOR M6 (roller clamping) problem. Meaning: Selftest of MOTOR M6 (roller clamping) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M6 (roller clamping) problem 2) MOTOR M6 (roller clamping) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M6 (roller clamping), re-connect cable at HANDLING CONTROL BOARD (P12) 2) Exchange the MOTOR M6 (roller clamping), refer to the Service Manual for order and replacement.
12802 | ROLLERCLAMPER_MOTOR_DEFECTIVE | Error |  |  |  | Roller Clamping motor(M6) defective. | Meaning: Roller Clamping motor(M6) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M6. 2) Replace Handling board. | IPHM_RollClamperMotorError | CIR601 | MOTOR M6 (roller clamping) defective. Meaning: Selftest of MOTOR M6 (roller clamping) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in MOTOR M6 (roller clamping) or motor cable. 2) Electrical short in the control for MOTOR M6 (roller clamping) on the HANDLING CONTROL BOARD. Cure: 1) First exchange the MOTOR M6 (roller clamping), refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12804 | ROLLERCLAMPER_DRIVER | Error |  |  |  | Roller Clamping motor(M6), driver defective. | Meaning: Roller Clamping motor(M6) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | IPHM_RollClamperDriverError | CIR603 | Motor control for MOTOR M6 (roller clamping) is defective. Meaning: Selftest of MOTOR M6 (roller clamping) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M6 (roller clamping) on the HANDLING CONTROL BOARD is defective. Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12805 | ROLLERCLAMPER_MOVEMENT | Error |  |  |  | Roller Clamping motor(M6), movement failed. | Meaning: Roller Clamping motor(M6) movement failed. Reason: 1) Wrong stall parameters. 2) Drive is blocked by something or runs stiff. 3) Drive does not move at all. 4) Steering problem by software or hardware. Cure: 1) Do stall calibration of Roller Clamping motor M6. 2)Check mechanics and replace defective part(s). 3) Check electronis (motor, cables, handling board) and replace defective part(s). 4) Reinstall software or replace Handling board. | IPHM_RollClamperClampError | CIR604 | MOTOR M6 (roller clamping) moves with difficulty. Meaning: The drive (MOTOR M6) of the roller clamping does not reach the end position 'roller clamped'. The Digitizer is out of operation and requires service support. Reason: 1) Any obstacle in the travel range of the drive and lifting mechanism of the roller clamping 2) Drive shaft is not proper mounted: Gear wheels at the left and right side on the drive shaft are meshing the gear racks left and right with an offset (not in sync) 3) Infrequent contact problems of cable or MOTOR M6 Cure: 1) Restart the Digitizer and check, if the problem persists 2) Check the gear wheels and gear racks on the left and right side of the drive shaft for proper function and ease of movement 3) Check the lifting mechanism of the roller clamping for proper function and ease of movement 4) Check, if the gear wheel drive on the left and right side are in sync to lift the rollers on both sides in the same way 5) Perform the 'Endurance Run Cycle' (e.g. 20 cycles) in the Service User Interface to detect infrequent contact problems, refer to the Service Manual 6) For replacement of the Input Tray contact your Regional Support Network for further assistance
12807 | DRAWER_MOTOR_UNPLUGGED | Error |  |  |  | Drawer motor(M4) unplugged. | Meaning: Drawer motor(M4) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P9. 2) Replace motor M4. | IPHM_DrawerUnpluggedError | CIR672 | Contact or cable to MOTOR M4 (drawer unit) problem. Meaning: Selftest of MOTOR M4 (drawer unit) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M4 (drawer unit) problem 2) MOTOR M4 (drawer unit) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M4 (drawer unit), re-connect cable at HANDLING CONTROL BOARD (P9) 2) Exchange the MOTOR M4 (drawer unit), refer to the Service Manual for order and replacement.
12808 | DRAWER_MOTOR_DEFECTIVE | Error |  |  |  | Drawer motor(M4) defective. | Meaning: Drawer motor(M4) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M4. 2) Replace Handling board. | IPHM_DrawerMotorError | CIR671 | MOTOR M4 (drawer unit) defective. Meaning: Selftest of MOTOR M4 (drawer unit) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in MOTOR M4 (drawer unit) or motor cable. 2) Electrical short in the control for MOTOR M4 (drawer unit) on the HANDLING CONTROL BOARD. Cure: 1) First exchange the MOTOR M4 (drawer unit), refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12809 | DRAWER_MOVEMENT | Error |  |  |  | Drawer motor(M4), movement failed. | MOTOR M4 (drawer) movement drive parameter out of range. Meaning: Movement parameter to drive the drawer (MOTOR M4) are out of range. The Digitizer is out of operation and requires service support. . Reason: 1) Parameter, which defines the movement of the drawer (MOTOR M4) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance. | IPHM_DrawerMovementError
12810 | DRAWER_DRIVER | Error |  |  |  | Drawer motor(M4), driver defective | Meaning: Drawer motor(M4) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | IPHM_DrawerDriverError | CIR673 | Motor control for MOTOR M4 (drawer unit) is defective. Meaning: Selftest of MOTOR M4 (drawer unit) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M4 (drawer unit) on the HANDLING CONTROL BOARD is defective. Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12813 | ROLLER_MOTOR_UNPLUGGED | Error |  |  |  | Transport roller motor(M5) unplugged. | Meaning: Transport roller motor(M5) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P13. 2) Replace motor M5. | IPHM_RollerUnpluggedError | CIR452 | Contact or cable to MOTOR M5 (transport roller) problem. Meaning: Selftest of MOTOR M5 (transport roller) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M5 (transport roller) problem 2) MOTOR M5 (transport roller) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M5 (transport roller), re-connect cable at HANDLING CONTROL BOARD (P13) 2) Exchange the MOTOR M5 (transport roller), refer to the Service Manual for order and replacement.
12814 | ROLLER_MOTOR_DEFECTIVE | Error |  |  |  | Transport roller motor(M5) defective. | Meaning: Transport Roller motor(M5) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M5. 2) Replace Handling board. | IPHM_RollerMotorError | CIR451 | MOTOR M5 (transport roller) defective. Meaning: Selftest of MOTOR M5 (transport roller) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in M5 (transport roller) or motor cable. 2) Electrical short in the control for MOTOR M5 (transport roller) on the HANDLING CONTROL BOARD. Cure: 1) First exchange the MOTOR M5 (transport roller), refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12816 | ROLLER_DRIVER | Error |  |  |  | Transport roller motor(M5), driver defective. | Meaning: Transport Roller motor(M5) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | IPHM_RollerDriverError | CIR453 | Motor control for MOTOR M5 (transport roller) is defective. Meaning: Selftest of MOTOR M5 (transport roller) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M5 (transport roller) on the HANDLING CONTROL BOARD is defective. Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
12817 | ROLLER_MOVEMENT | Error |  |  |  | Transport roller motor(M5), movement failed. | Meaning: Transport Roller motor(M5) movement failed. Reason: 1) Drive is blocked by something or runs stiff. 2) Drive does not move at all. 3) Steering problem by software or hardware. Cure: 1)Check mechanics and replace defective part(s). 2) Check electronis (motor, cables, handling board) and replace defective part(s). 3) Reinstall software or replace Handling board. | IPHM_RollerMovementError | CIR454 | MOTOR M5 (transport roller) movement drive parameter out of range. Meaning: Movement parameter to drive the transport roller (MOTOR M5) are out of range. The Digitizer is out of operation and requires service support. . Reason: 1) Parameter, which defines the movement of the transport roller (MOTOR M5) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
13313 | SLOWSCAN_MOTOR_UNPLUGGED | Error |  |  |  | Slowscan motor(M2) unplugged. | Meaning: Slowscan motor(M2) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P6. 2) Replace motor M2. | SSU_SSUnpluggedError | CIR252 | Contact or cable to MOTOR M2 (slow scan) problem. Meaning: Selftest of MOTOR M2 (slow scan) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M2 (slow scan) problem 2) MOTOR M2 (slow scan) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M2 (slow scan), re-connect cable at HANDLING CONTROL BOARD (P6) 2) Exchange the MOTOR M2 (slow scan), refer to the Service Manual for order and replacement .
13314 | SLOWSCAN_MOTOR_DEFECTIVE | Error |  |  |  | Slowscan motor(M2) defective. | Meaning: Slowscan motor(M2) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M2. 2) Replace Handling board. | SSU_SSMotorError | CIR251 | MOTOR M2 (slowscan) defective. Meaning: Selftest of MOTOR M2 (slow scan) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in MOTOR M2 (slow scan) or motor cable. 2) Electrical short in the control for MOTOR M2 (slow scan) on the HANDLING CONTROL BOARD. Cure: 1) First exchange the MOTOR M2 (slow scan), refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
13316 | SLOWSCAN_DRIVER | Error |  |  |  | Slowscan motor(M2), driver defective. | Meaning:Slowscan motor(M2) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | SSU_SSDriverError | CIR253 | Motor control for MOTOR M2 (slow scan) is defective. Meaning: Selftest of MOTOR M2 (slow scan) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M2 (slow scan) on the HANDLING CONTROL BOARD is defective. Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement .
13317 | SLOWSCAN_MOVEMENT | Error |  |  |  | Slowscan motor(M2), movement failed. | Meaning: Slowscan motor(M2) movement failed. Reason: 1) Drive is blocked by something or runs stiff. 2) Drive does not move at all. 3) Steering problem by software or hardware. Cure: 1)Check mechanics and replace defective part(s). 2) Check electronis (motor, cables, handling board) and replace defective part(s). 3) Reinstall software or replace Handling board. | SSU_SSMovementError | CIR254 | MOTOR M2 (slow scan) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan (MOTOR M2) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan (MOTOR M2) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
13319 | SLOWSCAN_LIFT_MOTOR_UNPLUGGED | Error |  |  |  | Slowscan rolls lift motor(Mx) unplugged. | Meaning: Slowscan rolls lift motor(M7) selftest failed. Reason: 1) Motor is unplugged. 2) Motor or cable defective. Cure: 1) Plug in motor cable to handling board connector P11. 2) Replace motor M7. | SSU_SSLUnpluggedError | CIR262 | Contact or cable to MOTOR M7 (slow scan roller lift) problem. Meaning: Selftest of MOTOR M7 (slow scan roller lift) failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact or cable to MOTOR M7 (slow scan roller lift) problem 2) MOTOR M7 (slow scan roller lift) or motor cable is defective. Cure: 1) First check the cabling and connection of the MOTOR M7 (slow scan roller lift), re-connect cable at HANDLING CONTROL BOARD (P11) 2) Exchange the MOTOR M7 (slow scan roller lift), refer to the Service Manual for order and replacement
13320 | SLOWSCAN_LIFT_MOTOR_DEFECTIVE | Error |  |  |  | Slowscan rolls lift motor(Mx) defective. | Meaning: Slowscan rolls lift motor(M7) selftest failed. Reason: 1) Motor is defective. 2) Handling board is defective. Cure: 1) Replace motor M7. 2) Replace Handling board. | SSU_SSLMotorError | CIR261 | MOTOR M7 (slow scan roller lift) defective Meaning: Selftest of MOTOR M7 (slow scan roller lift) failed. The Digitizer is out of operation and requires service support. Reason: 1) Electrical short in MOTOR M7 (slow scan roller lift) or motor cable. 2) Electrical short in the control for MOTOR M7 (slow scan roller lift) on the HANDLING CONTROL BOARD Cure: 1) First exchange the MOTOR M7 (slow scan roller lift), refer to the Service Manual for order and replacement 2) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement.
13322 | SLOWSCAN_LIFT_DRIVER | Error |  |  |  | Slowscan rolls lift motor(Mx), driver defective. | Meaning:Slowscan rolls lift motor(M7) selftest failed. Reason: Motor driver on handling board is defective. Cure: Replace Handling board. | SSU_SSLDriverError | CIR263 | Motor control for MOTOR M7 (slow scan roller lift) is defective . Meaning: Selftest of MOTOR M7 (slow scan roller lift) failed. The Digitizer is out of operation and requires service support. Reason: 1) Motor control for MOTOR M7 (slow scan roller lift) on the HANDLING CONTROL BOARD is defective Cure: 1) Exchange the HANDLING CONTROL BOARD, refer to the Service Manual for order and replacement
13323 | SLOWSCAN_LIFT_MOVEMENT | Error |  |  |  | Slowscan rolls lift motor(Mx), movement failed. | Meaning: Slowscan rolls lift motor(M7) movement failed. Reason: 1) Drive is blocked by something or runs stiff. 2) Drive does not move at all. 3) Steering problem by software or hardware. Cure: 1)Check mechanics and replace defective part(s). 2) Check electronis (motor, cables, handling board) and replace defective part(s). 3) Reinstall software or replace Handling board. | SSU_SSLMovementError | CIR264 | MOTOR M7 (slow scan roller lift) movement drive parameter out of range. Meaning: Movement parameter to drive the slow scan roller lift (MOTOR M7) are out of range. The Digitizer is out of operation and requires service support. Reason: 1) Parameter, which defines the movement of the slow scan roller lift (MOTOR M7) are out of range. Cure: 1) Reinstall the digitizer software, refer to the relevant software installation instruction 2) If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
13824 | INIT_MODULE_FAILED_PLACEHOLDER | Error |  |  |  | Testlog sent during initialisation phase | Meaning: Serious error, function disturbed. Reason: The needed elements could not be intialised to valid values or function Cure: Restart Digitizer. If error persists please call Service.
13924 | WARN_CASS_INVALID_PLACEHOLDER | Error |  |  |  | Warning: Clamp Cassette Movement detected invalid Cassette Width | Meaning: Retry possible, function disturbed. Reason: The clamped Cassette has a invalid width. Cure: Re-clamp Cassette. If error persists please calibrate mechanical Cassette Unit.
13925 | FIXER_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Fixer failed. | Meaning: The calibration of the Fixer failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14854 | OPTICS_BOARD_CONNECTOR | Error |  |  |  | Optics board not detected | Meaning: Optic board not detected. Reason: Cable broken/not connected or optic board defective. Cure: 1) Check optic module cabling. 2) Replace optic module, if error persists. | COM_OpticsConnectionErr | CIR222 | Connection between PMI Board and Optic module failed. Meaning: Reason: Cure:
14855 | LASER_DIODE | Error |  |  |  | Check laser diode failed | Meaning: Check laser diode failed. Reason: Laser diode defective. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement | COM_NotNeeded_ForServiceId
14856 | OPTICS_TYPE | Error |  |  |  | Unknown optics type detected | Meaning: Unknown optics type detected Reason: Unknown optics type in elabel. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement | COM_NotNeeded_ForServiceId
14857 | INIT_ELABEL | Error |  |  |  | Error initializing optics eLabel | Meaning: Initialization of OPTIC MODULE e-label (module specific parameter storage) failed. The Digitizer is out of operation and requires service support. Reason: OPTIC MODULE e-label can't be initialized due to: - OPTIC MODULE not connected - Defective of the PMI-BOARD - OPTIC MODULE cable or connectors - Corrupted data structure on the e-label - Defective OPTIC MODULE Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between PMI-BOARD and OPTIC MODULE 3. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement 4. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement | COM_NotNeeded_ForServiceId
14858 | PMM_ELABEL | Error |  |  |  | Error getting data from PMM eLabel | Meaning: Can't read data from PMM eLabel. Reason: PMM eLabel defective or PMM not connected. Cure: 1) Check PMM cabling. 2) Replace PMM cabling. 3) Replace PMI board. 4) Replace PMM. | COM_NotNeeded_ForServiceId
14859 | PMM_BOARD_CONNECTOR | Error |  |  |  | PMM board not detected | Meaning: PMM board not detected Reason: PMM is not connected to PMI board. Cure: 1) Check PMM cabling. 2) Replace PMM cabling, if error persists. 3) Replace PMM, if error persists. | COM_PMMConnectionErr | CIR221 | Connection between PMI Board and PMT failed. Meaning: Reason: Cure:
14860 | START_POLYGON | Error |  |  |  | Start or Lock polygon failed | Meaning: Polygon does not reach its speed. Reason: Power supply or polygon defective. Cure: Check power supply, replace optics module. | COM_NotNeeded_ForServiceId
14861 | UPDATE_PMM | Error |  |  |  | Error: PMM firmware update failed - Switch digitizer Off/On for retry | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists. | COM_NotNeeded_ForServiceId
14862 | UPDATE_PMM_FAILED | Error |  |  |  | Error: PMM firmware update failed | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists. | COM_NotNeeded_ForServiceId
14863 | PMMADJ_LPV_TOO_HIGH | Error |  |  |  | Prepare PMM adjust: LPV is too high |  | COM_NotNeeded_ForServiceId
14864 | PMMADJ_LED_BRIGHTNESS | Error |  |  |  | Prepare PMM adjust: Error LED not bright enough |  | COM_NotNeeded_ForServiceId
14865 | PMMADJ_SEL_SENS | Error |  |  |  | Prepare PMM adjust: Error selecting sensitivity |  | COM_NotNeeded_ForServiceId
14866 | PMMADJ_INCIDENCE_OF_LIGHT | Error |  |  |  | Prepare PMM adjust: Error because incidence of light |  | COM_NotNeeded_ForServiceId
14867 | PMMADJ_WRITE_ROV | Error |  |  |  | Analyse PMM adjust: Error writing new REFCAL to PMM |  | COM_NotNeeded_ForServiceId
14868 | PMMADJ_PARAMETER | Error |  |  |  | Analyse PMM adjust: Error parameter real dose or target index out of range |  | COM_NotNeeded_ForServiceId
14869 | BOL_TIMEOUT | Error |  |  |  | Lost BOL signal | Meaning: ESD or optics problem. Reason: electric discharge or optics module defective. Cure: if permanent check optics module. | COM_NotNeeded_ForServiceId
14870 | PMI_ELABEL | Error |  |  |  | Error getting data from PMI eLabel | Meaning: Error getting data from PMI eLabel. Reason: PMI eLabel defective. Cure: Replace PMI board. | COM_NotNeeded_ForServiceId
14871 | DIST_BOL_EOL_OUT_OF_RANGE | Error |  |  |  | Distance BOL to EOL out of range | Meaning: Distance BOL to EOL out of range. Reason: Polygon speed out of range, Optics module out of range. Cure: Check polygon speed (image width), replace optics module. | COM_NotNeeded_ForServiceId
14872 | PMI_HW_VER | Error |  |  |  | Software needs PMI-Board hardware version 1 or later | Meaning: PMI software does not fit to PMI hardware. Reason: Software update to an incompatible PMI board. Cure: Replace PMI board to index 1 or higher. | COM_NotNeeded_ForServiceId
14873 | PMI_FPGA_VER | Error |  |  |  | FPGA version does not fit to software version | Meaning: PMI software does not fit to PMI FPGA hardware. Reason: Software update failed or not completed. Cure: Update PMI software. | COM_NotNeeded_ForServiceId
14874 | OPTICS_ELABEL_CORRUPTED | Error |  |  |  | Optics eLabel data corrupted | Meaning: Optics eLabel can't be read, need to be restored from PC. Reason: Cable from PMI to optics module too long or unknown. Cure: Restore optics eLabel from PC. | COM_NotNeeded_ForServiceId
14875 | LD_AGING | Error |  |  |  | Laser diode is aging | Meaning: Laser diode is supplied with additional electrical power (1.5 of the standard value) to compensate aging for a while. Reason: Laser diode is short before end of life. | COM_NotNeeded_ForServiceId
14876 | INIDOSEDET_IPTYPEORDOSE | Error |  |  |  | IP type or dose out of range | Meaning: IP type or dose out of range. Reason: Software error. Cure: Install correct software. | COM_SWError
14877 | INIDOSEDET_WRONGHARDWARE | Error |  |  |  | initOverExpComparatorsFromData called on PMI1/2 | Meaning: The function initOverExpComparatorsFromData was called on the wrong hardware PMI1 or 2. Reason: Software error. Cure: Install correct software. | COM_NotNeeded_ForServiceId
14878 | PMM_BLUE_LED | Error |  |  |  | Photo Multiplier Module (PMM/PMT), blue LED defective or not connected | Meaning: Can't detect PMT blue LED. Reason: - PMT blue LED defective or not connected - PMT defective Cure: 1. Check connection and cable from blue LED to PMT 2. If error still persists replace PMT + LIGHT COLLECTOR | COM_NotNeeded_ForServiceId
14879 | PMM_TUBE_BOARD | Error |  |  |  | Photo Multiplier Module (PMM/PMT)defective or tube board not connected | Meaning: Can't read set HV. Reason: - PMM defective or tube board not connected Cure: 1. Check connection and cable from PMM main- to PMM tube board 2. If error still persists replace PMT + LIGHT COLLECTOR | COM_NotNeeded_ForServiceId
14880 | OPTICS_ELABEL | Error |  |  |  | Error getting data from optics eLabel | Meaning: Can't read data from optics eLabel. Reason: Optic module eLabel defective or optic module not connected. Cure: 1) Check optic module cabling. 2) Replace optic module cabling. 3) Replace PMI board. 4) Replace optic module. | COM_NotNeeded_ForServiceId
14881 | COULD_NOT_READ_PMM_FLASH_CNT | Warning |  |  |  | Error reading counters from PMM flash | Meaning: Cannot read counters from PMM flash memory. Cycle counter and operating hours may not correspond to real values of this PMM module. | COM_NotNeeded_ForServiceId
14882 | COULD_NOT_WRITE_PMM_FLASH_CNT | Warning |  |  |  | Error writing counters into PMM flash | Meaning: Cannot write counters into PMM flash memory. Cycle counter and operating hours may not correspond to real values of this PMM module. | COM_NotNeeded_ForServiceId
14883 | COULD_NOT_WRITE_OPTICS_ELABEL_CNT | Warning |  |  |  | Error writing counters into optics eLabel | Meaning: Cannot write counters into optics eLabel. Cycle counter and operating hours may not correspond to real values of this optics module. | COM_NotNeeded_ForServiceId
14884 | COULD_NOT_READ_OPTICS_ELABEL_CNT | Warning |  |  |  | Error reading counters from optics eLabel | Meaning: Cannot read counters from optics eLabel. Cycle counter and operating hours may not correspond to real values of this optics module. | COM_NotNeeded_ForServiceId
14885 | COULD_NOT_READ_OPTICS_ELABEL_INFO | Error |  |  |  | Error reading serial number from optics eLabel | Meaning: Cannot read serialnumber from optics eLabel . Reason: The device eLabel is empty, was formatted, contains invalid data or is disconnected. Cure: 1. Format the Device E-Label via ESF, then write valid data to the Device E-Label via ESF. 2. Check cabling connection berween PMI board and optics eLabel. | COM_OpticModuleDataAccessErr
14886 | COULD_NOT_READ_PMM_INFO | Error |  |  |  | Error reading serial number from PMM flash memory | Meaning: Cannot read serialnumber from PMM flash memory . Reason: The flash Memory is empty, was formatted, contains invalid data or is disconnected. Cure: 1.write valid data to the PMM flash memory via ESF. 2. Check cabling connection berween PMI board and PMM module. | COM_PMMDataAccessErr
15360 | DIC_TO_IHC_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while DIC is sending to IHC | Meaning: No receiver found after sending message from DigitizerControl (DIC) on PMI BOARD to IP-HandlingControl (IHC) on HANDLING CONTROL BOARD. Reason: Software error, input queue of IHC may be full. Cure: Install new software; inform R&D team.
15361 | DIC_TO_IHC_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from DIC to IHC | Meaning: The CAN channel number for sending message from DigitizerControl (DIC) on PMI BOARD to IP-HandlingControl (IHC) on HANDLING CONTROL BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15362 | DIC_TO_IHC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | DIC could not read acknowledge from IHC | Meaning: Could not read acknowledge after sending message from DigitizerControl (DIC) on PMI BOARD to IP-HandlingControl (IHC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on HANDLING CONTROL BOARD; if connections are OK, install new software and inform R&D team.
15363 | DIC_TO_IHC_CAN_ERROR | Fatal error |  |  |  | sending from DIC to IHC failed with CAN fast channel error | Meaning: CAN error during sending message from DigitizerControl (DIC) on PMI BOARD to IP-HandlingControl (IHC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15364 | DIC_TO_CHC_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while DIC is sending to CHC | Meaning: No receiver found after sending message from DigitizerControl (DIC) on PMI BOARD to CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD. Reason: Software error, input queue of CHC may be full. Cure: Install new software; inform R&D team.
15365 | DIC_TO_CHC_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from DIC to CHC | Meaning: The CAN channel number for sending message from DigitizerControl (DIC) on PMI BOARD to CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15366 | DIC_TO_CHC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | DIC could not read acknowledge from CHC | Meaning: Could not read acknowledge after sending message from DigitizerControl (DIC) on PMI BOARD to CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on HANDLING CONTROL BOARD; if connections are OK, install new software and inform R&D team.
15367 | DIC_TO_CHC_CAN_ERROR | Fatal error |  |  |  | sending from DIC to CHC failed with CAN fast channel error | Meaning: CAN error during sending message from DigitizerControl (DIC) on PMI BOARD to CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15368 | DIC_TO_TRC_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while DIC is sending to TRC | Meaning: No receiver found after sending message from DigitizerControl (DIC) on PMI BOARD to TagReaderControl (TRC) on HANDLING CONTROL BOARD. Reason: Software error, input queue of TRC may be full. Cure: Install new software; inform R&D team.
15369 | DIC_TO_TRC_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from DIC to TRC | Meaning: The CAN channel number for sending message from DigitizerControl (DIC) on PMI BOARD to TagReaderControl (TRC) on HANDLING CONTROL BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15370 | DIC_TO_TRC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | DIC could not read acknowledge from TRC | Meaning: Could not read acknowledge after sending message from DigitizerControl (DIC) on PMI BOARD to TagReaderControl (TRC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on HANDLING CONTROL BOARD; if connections are OK, install new software and inform R&D team.
15371 | DIC_TO_TRC_CAN_ERROR | Fatal error |  |  |  | sending from DIC to TRC failed with CAN fast channel error | Meaning: CAN error during sending message from DigitizerControl (DIC) on PMI BOARD to TagReaderControl (TRC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15372 | UID_TO_UIC_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while UID is sending to UIC | Meaning: No receiver found after sending message from UserInterfacDispatcher(UID) on PMI BOARD to UserInterfaceControl (UIC) on HANDLING CONTROL BOARD. Reason: Software error, input queue of UIC may be full. Cure: Install new software; inform R&D team.
15373 | UID_TO_UIC_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from UID to UIC | Meaning: The CAN channel number for sending message from UserInterfacDispatcher(UID) on PMI BOARD to UserInterfaceControl (UIC) on HANDLING CONTROL BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15374 | UID_TO_UIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | UID could not read acknowledge from UIC | Meaning: Could not read acknowledge after sending message from UserInterfacDispatcher(UID) on PMI BOARD to UserInterfaceControl (UIC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on HANDLING CONTROL BOARD ; if connections are OK, install new software and inform R&D team.
15375 | UID_TO_UIC_CAN_ERROR | Fatal error |  |  |  | sending from UID to UIC failed with CAN fast channel error | Meaning: CAN error during sending message from UserInterfacDispatcher(UID) on PMI BOARD to UserInterfaceControl (UIC) on HANDLING CONTROL BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15376 | IHC_TO_DIC_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while IHC is sending to DIC | Meaning: No receiver found after sending message from IP-HandlingControl (IHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: Software error, input queue of DIC may be full. Cure: Install new software; inform R&D team.
15377 | IHC_TO_DIC_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from IHC to DIC | Meaning: The CAN channel number for sending message from IP-HandlingControl (IHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15378 | IHC_TO_DIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | IHC could not read acknowledge from DIC | Meaning: Could not read acknowledge after sending message from IP-HandlingControl (IHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15379 | IHC_TO_DIC_CAN_ERROR | Fatal error |  |  |  | sending from IHC to DIC failed with CAN fast channel error | Meaning: CAN error during sending message from IP-HandlingControl (IHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15380 | CHC_TO_DIC_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while CHC is sending to DIC | Meaning: No receiver found after sending message from CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: Software error, input queue of DIC may be full. Cure: Install new software; inform R&D team.
15381 | CHC_TO_DIC_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from CHC to DIC | Meaning: The CAN channel number for sending message from CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15382 | CHC_TO_DIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | CHC could not read acknowledge from DIC | Meaning: Could not read acknowledge after sending message from CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15383 | CHC_TO_DIC_CAN_ERROR | Fatal error |  |  |  | sending from CHC to DIC failed with CAN fast channel error | Meaning: CAN error during sending message from CassetteHandlingControl (CHC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15384 | TRC_TO_DIC_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while TRC is sending to DIC | Meaning: No receiver found after sending message from TagReaderControl (TRC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: Software error, input queue of DIC may be full. Cure: Install new software; inform R&D team.
15385 | TRC_TO_DIC_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from TRC to DIC | Meaning: The CAN channel number for sending message from TagReaderControl (TRC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15386 | TRC_TO_DIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | TRC could not read acknowledge from DIC | Meaning: Could not read acknowledge after sending message from TagReaderControl (TRC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15387 | TRC_TO_DIC_CAN_ERROR | Fatal error |  |  |  | sending from TRC to DIC failed with CAN fast channel error | Meaning: CAN error during sending message from TagReaderControl (TRC) on HANDLING CONTROL BOARD to DigitizerControl (DIC) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15388 | UIC_TO_UID_NO_RECEIVER_FOUND | Fatal error |  |  |  | no suitable receiver found while UIC is sending to UID | Meaning: No receiver found after sending message from UserInterfaceControl (UIC) on HANDLING CONTROL BOARD to UserInterfacDispatcher(UID) on PMI BOARD. Reason: Software error, input queue of UID may be full. Cure: Install new software; inform R&D team.
15389 | UIC_TO_UID_UNKNOWN_CHANNEL | Fatal error |  |  |  | unknown CAN channel for sending from UIC to UID | Meaning: The CAN channel number for sending message from UserInterfaceControl (UIC) on HANDLING CONTROL BOARD to UserInterfacDispatcher(UID) on PMI BOARD is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15390 | UIC_TO_UID_ERR_READING_ACKNOWLEDGE | Fatal error |  |  |  | UIC could not read acknowledge from UID | Meaning: Could not read acknowledge after sending message from UserInterfaceControl (UIC) on HANDLING CONTROL BOARD to UserInterfacDispatcher(UID) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15391 | UIC_TO_UID_CAN_ERROR | Fatal error |  |  |  | sending from UIC to UID failed with CAN fast channel error | Meaning: CAN error during sending message from UserInterfaceControl (UIC) on HANDLING CONTROL BOARD to UserInterfacDispatcher(UID) on PMI BOARD. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15392 | INTERFACE_WARNING | Warning |  |  |  | adapter received a warning from the interface | Meaning: The adapter received a warning from the interface. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15972 | MAXIMUM_NUMBER_OF_PLATEAUS_EXCEEDED | Error |  |  |  | The maximum number of plateaus was exceeded. | Meaning: The maximum number of plateaus was exceeded. Reason: The system (mechanics/stepper motor) is not stable enough for a calibration. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
15973 | NO_PLATEAU_FOUND_FOR_EVALUATION | Error |  |  |  | No plateau was found for determining the optimal run current level. | Meaning: No plateau was found for determining the optimal run current level. Reason: The system (mechanics/stepper motor) is not stable enough for a calibration. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
15974 | WRONG_RUN_CURRENT_VALUE_SET | Error |  |  |  | The run current value set for the calibration cycle is wrong. | Meaning: The run current value set for the calibration cycle is wrong. The required and the set values are logged. Reason: -- Cure: Check the corresponding SW routine
15975 | FUNCTION_POINTER_ZERO | Error |  |  |  | A function pointer was zero. | Meaning: A function pointer was zero. Reason: -- Cure: Check the corresponding SW routine
15976 | MOTOR_SIM_CALIBRATION_FAILED | Error |  |  |  | The calibration of the Motor Simulation failed. | Meaning: The calibration of the Motor Simulation failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding Software or Settings.
15977 | MOTOR_SIM_CALIBRATION_ABORTED | Error |  |  |  | The calibration procedure of the Motor Simulation was aborted. | Meaning: The calibration procedure of the Motor Simulation was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
15978 | MOTOR_SIM_CALIBRATION_INVALID_VALUE | Error |  |  |  | The determined optimal run current value of the Motor Simulation was invalid. | Meaning: The determined optimal run current value of the Motor Simulation was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding Software or Settings. Cure: Check the corresponding Software and Settings and retry the calibration procedure.
15979 | CB_FUNCTION_POINTER_ZERO | Error |  |  |  | A cb function pointer was zero. | Meaning: A cb function pointer was zero. Reason: -- Cure: Check the corresponding SW routine
15980 | WRONG_RUN_CURRENT_VALUE_STORED | Error |  |  |  | The run current value stored in the eLabel is wrong. | Meaning: The run current value stored in the eLabel is wrong. The required and the set values are logged. Reason: -- Cure: Check the corresponding SW routines and eventually the hardware.
19456 | TXFLDSENDMSG | Error |  |  |  | transmitFolder() - sendMsg failed |  | COM_DicomErr
19457 | TXFLDABORTASSOCIATION | Error |  |  |  | transmitFolder() - abortAssociation() failed |  | COM_DicomErr
19458 | CREATEOBJECT | Error |  |  |  | assembleDICOMObject() - createObject failed |  | COM_SWError
19459 | ADDPATNAME | Error |  |  |  | assembleDICOMObject() - adding attr PatientName failed |  | COM_SWError
19460 | ADDPATID | Error |  |  |  | assembleDICOMObject() - adding attr PatientBirthDatefailed |  | COM_SWError
19461 | ADDPATBIRTHDATE | Error |  |  |  | assembleDICOMObject() - adding attr PatientName failed |  | COM_SWError
19462 | ADDPATSEX | Error |  |  |  | assembleDICOMObject() - adding attr PatientSex failed |  | COM_SWError
19463 | ADDSTUDYINSTUID | Error |  |  |  | assembleDICOMObject() - adding attr StudyInstanceUID failed |  | COM_SWError
19464 | ADDSTUDYDATE | Error |  |  |  | assembleDICOMObject() - adding attr StudyDate failed |  | COM_SWError
19465 | ADDSTUDYTIME | Error |  |  |  | assembleDICOMObject() - adding attr StudyTime failed |  | COM_SWError
19466 | ADDACCESSIONNUMBER | Error |  |  |  | assembleDICOMObject() - adding attr AccessionNumber failed |  | COM_SWError
19467 | ADDREFPHYSICIAN | Error |  |  |  | assembleDICOMObject() - adding attr ReferringPhysiciansName failed |  | COM_SWError
19468 | ADDSTUDYID | Error |  |  |  | assembleDICOMObject() - adding attr StudyID failed |  | COM_SWError
19469 | ADDMODALITY | Error |  |  |  | assembleDICOMObject() - adding attr Modality failed |  | COM_SWError
19470 | OFFLINESERIESINSTUID | Error |  |  |  | assembleDICOMObject() - adding attr PatientName failed |  | COM_SWError
19471 | ADDSERIESNUMBER | Error |  |  |  | assembleDICOMObject() - adding attr SeriesNumber failed |  | COM_SWError
19472 | ADDLATERALITY | Error |  |  |  | assembleDICOMObject() - adding attr Laterality failed |  | COM_SWError
19473 | ADDBODYPART | Error |  |  |  | assembleDICOMObject() - adding attr BodyPart failed |  | COM_SWError
19474 | ADDPLATETYPE | Error |  |  |  | assembleDICOMObject() - adding attr PlateType failed |  | COM_SWError
19475 | ADDVIEWPOSITION | Error |  |  |  | assembleDICOMObject() - adding attr ViewPosition failed |  | COM_SWError
19476 | ADDMANUFACTURER | Error |  |  |  | assembleDICOMObject() - adding attr Manufacturer failed |  | COM_SWError
19477 | ADDMANUFACTURERMODELNAME | Error |  |  |  | assembleDICOMObject() - adding attr ManufacturerModelName failed |  | COM_SWError
19478 | ADDSERIALNUMBER | Error |  |  |  | assembleDICOMObject() - adding attr SerialNumber failed |  | COM_SWError
19479 | ADDSWVERSION | Error |  |  |  | assembleDICOMObject() - adding attr SoftwareVersion failed |  | COM_SWError
19480 | ADDIMAGETYPE | Error |  |  |  | assembleDICOMObject() - adding attr ImageType failed |  | COM_SWError
19481 | ADDACQDATE | Error |  |  |  | assembleDICOMObject() - adding attr AcquisitionDate failed |  | COM_SWError
19482 | ADDACQTIME | Error |  |  |  | assembleDICOMObject() - adding attr AcquisitionTime failed |  | COM_SWError
19483 | ADDINSTANCENUMBER | Error |  |  |  | assembleDICOMObject() - adding attr InstaceNumber failed |  | COM_SWError
19484 | ADDPATORIENT | Error |  |  |  | assembleDICOMObject() - adding attr PatientOrientation failed |  | COM_SWError
19485 | ADDIMAGECOMMENTS | Error |  |  |  | assembleDICOMObject() - adding attr ImageComments failed |  | COM_SWError
19486 | ADDPLATEID | Error |  |  |  | assembleDICOMObject() - adding attr PlateID failed |  | COM_SWError
19487 | ADDIMAGERPIXELSPACING | Error |  |  |  | assembleDICOMObject() - adding attr ImagerPixelSpacing failed |  | COM_SWError
19488 | ADDCASSETTESIZE | Error |  |  |  | assembleDICOMObject() - adding attr CassetteSize failed |  | COM_SWError
19489 | ADDEXPOSURESONPLATE | Error |  |  |  | assembleDICOMObject() - adding attr ExposuresOnPlate failed |  | COM_SWError
19490 | ADDRESCALEINTERC | Error |  |  |  | assembleDICOMObject() - adding attr RescaleIntercept failed |  | COM_SWError
19491 | ADDRESCALESLOPE | Error |  |  |  | assembleDICOMObject() - adding attr RescaleSlope failed |  | COM_SWError
19492 | ADDRESCALETYPE | Error |  |  |  | assembleDICOMObject() - adding attr RescaleType failed |  | COM_SWError
19493 | ADDPXLSAMPLES | Error |  |  |  | assembleDICOMObject() - adding attr PixelsPerSample failed |  | COM_SWError
19494 | ADDPHOTOMETRICINTERP | Error |  |  |  | assembleDICOMObject() - adding attr PhotometricInterpretation failed |  | COM_SWError
19495 | ADDROWS | Error |  |  |  | assembleDICOMObject() - adding attr Rows failed |  | COM_SWError
19496 | ADDCOLUMNS | Error |  |  |  | assembleDICOMObject() - adding attr Columns failed |  | COM_SWError
19497 | ADDALLOCEDBITS | Error |  |  |  | assembleDICOMObject() - adding attr AllocatedBits failed |  | COM_SWError
19498 | ADDSTOREDBITS | Error |  |  |  | assembleDICOMObject() - adding attr StoredBits failed |  | COM_SWError
19499 | ADDHIGHBIT | Error |  |  |  | assembleDICOMObject() - adding attr HighBit failed |  | COM_SWError
19500 | ADDPXLREPR | Error |  |  |  | assembleDICOMObject() - adding attr PixelRepresentation failed |  | COM_SWError
19501 | ADDSPECCHARSET | Error |  |  |  | assembleDICOMObject() - adding attr SpecificCharacterSet failed |  | COM_SWError
19502 | ADDINSTCREDATE | Error |  |  |  | assembleDICOMObject() - adding attr InstanceCreationDate failed |  | COM_SWError
19503 | ADDINSTCRETIME | Error |  |  |  | assembleDICOMObject() - adding attr InstanceCreationTime failed |  | COM_SWError
19504 | ADDSOPCLASSUID | Error |  |  |  | assembleDICOMObject() - adding attr SOPClassUID failed |  | COM_SWError
19505 | ADDSOPINSTUID | Error |  |  |  | assembleDICOMObject() - adding attr SOPInstanceUID failed |  | COM_SWError
19506 | ADDPRIVCREATOR | Error |  |  |  | assembleDICOMObject() - adding attr PrivateCreator failed |  | COM_SWError
19507 | ADDCASSETTEDATA | Error |  |  |  | assembleDICOMObject() - adding attr CassetteData failed |  | COM_SWError
19508 | ADDSCANDIRECTION | Error |  |  |  | assembleDICOMObject() - adding attr ScanDirection failed |  | COM_SWError
19509 | ADDIMG | Error |  |  |  | addImageFile() failed |  | COM_SWError
19510 | ADDCALLBACK | Error |  |  |  | addImageCallback() failed |  | COM_SWError
19511 | TXFLDRXMSG | Error |  |  |  | failed to receive a response from the SCP |  | COM_SWError
19512 | TXNDISRELEASE | Error |  |  |  | failed to release the association |  | COM_SWError
19513 | ADDIDSTUDYDESCRIPTION | Error |  |  |  | assembleDICOMObject() - adding attr StudyDescription failed |  | COM_SWError
19514 | ADDIDSERIESDESCRIPTION | Error |  |  |  | assembleDICOMObject() - adding attr SeriesDescription failed |  | COM_SWError
19515 | TXFINDPRESENTATIONCONTEXT | Error |  |  |  | transmitFolder - unable to find the presentation context ID |  | COM_DicomErr
19516 | DESTINATIONNOTFOUND | Error |  |  |  | destination not configured | the given destination name has not been found in the list of known destination. | COM_PSConfigurationErr
19517 | DESTINATIONNAMETOOLONG | Error |  |  |  | name of the destination is too long | the name of the destination exceeds 64 characters. | COM_PSConfigurationErr
19518 | DESTINATIONNOIPADDR | Error |  |  |  | This is not an ip address | The ipaddress exceeds the length of 15 characters. | COM_PSConfigurationErr
19519 | DESTINATIONCALLEDAETITLETOOLONG | Error |  |  |  | The called AE title of the destination is too long | The AE title of the destination exceeds 64 characters. | COM_PSConfigurationErr
19520 | NOUIDGENERATIONPOSSIBLE | Error |  |  |  | The generation of unique UIDs cannot be guaranteed | The system time used to generate unique identifiers as required by DICOM shows is not valid. Exchange the battery of the real-time clock or specify a valid timeserver (if applicable). | COM_SystemTimeErr | CIR218 | Battery of the CPMI Board(B410) is flat or discharged. Meaning: Reason: Cure: Exchange the CPMI Board (B410), refer to the Service Manual for order and replacement
19521 | DESTINATIONPRECAETITLETOOLONG | Error |  |  |  | The AE title of the fast preview destination is too long | The AE title of the destination exceeds 64 characters. | COM_PSConfigurationErr
19522 | TXOFFADDIMGFILE | Error |  |  |  | failed to add imgfile hook | cannot add file reference to the image file into the DICOM object - image file existing? | COM_SWError
19523 | TXONADDIMGFILE | Error |  |  |  | failed to add img memory hook | cannot add hook to the image data in memory into the DICOM object - software failure! | COM_SWError
19524 | BUILDASSOCRQ | Error |  |  |  | failed to build the association request | check networking configuration (calling AE title, called AE title) | COM_PSConfigurationErr
19525 | SENDASSOCRQ | Error |  |  |  | failed to send association request | DTF library was not able to send the association request. | COM_SWError
19526 | CONNBLDABORTASSOCIATION | Error |  |  |  | cannot abort association | cannot abort assoc after problem when assembling the DICOM object | COM_SWError
19527 | RECVANSWERTIMEOUT | Error |  |  |  | error while waiting for the response | Error reported while waiting for the response of the DIMSE_sendMsg call. The image is queued again, rely on automatic recovery. | COM_DicomErr
19528 | ADDIMAGESINACQUISITION | Error |  |  |  | assembleDICOMObject() - adding attr images in acquisition failed |  | COM_SWError
19968 | IPDATA_NOTACCESSIBLE | Error |  |  |  | IPData NULL reference | The reference to IPData is NULL for SpoolerInput. | COM_SWError
19969 | SCANPAR_NOTACCESSIBLE | Error |  |  |  | ScanPar NULL reference | The reference to ScanParameters is NULL for SpoolerInput. | COM_SWError
19970 | CREATEEMPTYQ | Error |  |  |  | sc_qecreate failed for 'empty' queue | out of ressources or software bug | COM_SWError
19971 | CREATEQUEUEDQ | Error |  |  |  | sc_qecreate failed for 'queued' queue | out of ressources or software bug | COM_SWError
19972 | CREATETXEDQ | Error |  |  |  | sc_qecreate failed for 'transmitted' queue | out of ressources or software bug | COM_SWError
19973 | CREATEWARNQ | Error |  |  |  | sc_qecreate failed for 'warning' queue | out of ressources or software bug | COM_SWError
19974 | MALLOCBUFFERS | Error |  |  |  | cannot allocate memory for all buffers | out of memory | COM_SWError
19975 | ENQUEUEBUFFERINIT | Error |  |  |  | enqueuing a buffer to the 'empty' queue failed | either queue is full or QID error | COM_SWError
19976 | MARKQUEUED | Error |  |  |  | cannot push to 'queued' queue | software bug | COM_SWError
19977 | GETNEWBUFFEREMPTY | Error |  |  |  | error accessing the 'empty' queue | software bug | COM_SWError
19978 | GETNEWBUFFERTXED | Error |  |  |  | error accessing the 'transmitted' queue | software bug | COM_SWError
19979 | MARKEMPTY | Error |  |  |  | cannot push to 'empty' queue | software bug | COM_SWError
19980 | MARKTRANSMITTED | Error |  |  |  | cannot push to 'transmitted' queue | software bug | COM_SWError
19981 | MARKWARNING | Error |  |  |  | cannot push to 'warning' queue | software bug | COM_SWError
19982 | GETNEXTQUEUED | Error |  |  |  | cannot access the 'queued' queue | software bug | COM_SWError
19983 | GETNEXTWARNING | Error |  |  |  | cannot access the 'warning' queue | software bug | COM_SWError
19984 | MALLOCIMGBLOCK | Error |  |  |  | cannot allocate 64 kBytes for image block buffer | out of resources?!?! | COM_SWError
19985 | OPENIMGFILE | Error |  |  |  | cannot open image file | open imagefile to continue from disk failed | COM_SWErrorNeedsFormatFlash
19986 | IMAGELARGERTHANDEFINED | Error |  |  |  | the file found on disk is larger than the control data indicates | software bug | COM_SWErrorNeedsFormatFlash
19987 | CIS_OPENIMGFILE | Error |  |  |  | cannot open the image file in correctImageSize | software bug | COM_SWErrorNeedsFormatFlash
19988 | CIS_WRITEIMGFILE | Error |  |  |  | cannot fill incomplete image to required size | check hardware (CF-Card) | COM_SWErrorNeedsFormatFlash
19989 | DESTNOTAVAILABLE | Info | 301 | Image cannot be transferred. Destination {0} is not available. Check the connection to the destination. Use the routing functionality to send the image to another destination and press the confirm button. |  | destination is not available or returned error | Either the transmission failed due to a network problem or the target of an image transmission did not react or did reject the image for any reason. Check the destination. Image can be rerouted. | COM_PSConfigurationErr
19990 | WRITEPIXELSTODISK | Error |  |  |  | write error in imagefile | Meaning: writing the pixels to disk failed. Reason: CF-card defective or not enough memory left in partion e:/ Cure: replace CF-card. | COM_SWErrorNeedsFormatFlash
19991 | WRITECTRLTODISK | Error |  |  |  | write error in controlfile | Meaning: writing the controlfile to disk failed. Reason: CF-card defective or not enough memory left in partion e:/ Cure: replace CF-card. | COM_SWErrorNeedsFormatFlash
19992 | NODISKSPACE | Error |  |  |  | no enough space for spooler | Meaning: The available diskspace on e:/ is not sufficient for the spooler. Reason: Either too many files stored on e:/ or disk capacity decreased below an acceptable value. Cure: Exchange CF-card. | COM_SWErrorNeedsFormatFlash
19993 | DESTNOTAVAILABLE1DEST | Info | 314 | Image cannot be transferred. Destination {0} is not available. Check the connection to the destination and press the confirm button. | Confirm | destination is not available or returned error | Either the transmission failed due to a network problem or the target of an image transmission did not react or did reject the image for any reason. Check the destination. No other destination available. | COM_PSConfigurationErr
19994 | DEL_TIMEOUT | Error |  |  |  | Timeout deleting image. | Meaning: Waiting for the end of a previous delete image request timed out. New request will not be executed. Cure: Try again later or delete the image folder manually and reset. | COM_SWError
19995 | DEL_TIMEOUT2 | Error |  |  |  | Timeout deleting image. | Meaning: Waiting for the end of a delete image request timed out. Request was cancelled. Cure: Try again later or delete the image folder manually and reset. | COM_SWError
20480 | GET_PSDATA | Error |  |  |  | getCPFData_ps failed | Retrieving the PS data from the CPF file failed | COM_NotNeeded_ForServiceId
20481 | GET_APPLDATA | Error |  |  |  | getCPFData_appl failed | Retrieving the APPL data from the CPF file failed | COM_NotNeeded_ForServiceId
20482 | GET_NETDATA | Error |  |  |  | getCPFData_net failed | Retrieving the NET data from the CPF file failed | COM_NotNeeded_ForServiceId
20483 | GET_DIGDATA | Error |  |  |  | getCPFData_dig failed | Retrieving the DIG data from the CPF file failed | COM_SWError
20484 | ADD_DEST_TO_DS | Error |  |  |  | AddDataToStore failed for DS_CPF_DESTCOORD | Adding the Destination Coordinates to the Data Store failed | COM_SWError
20485 | ADD_STATIONNAME_TO_DS | Error |  |  |  | AddDataToStore failed for DS_CPF_DIGCOORD | Adding the Station Names Coordinates to the Data Store failed | COM_SWError
20486 | NOMEM_DEST | Error |  |  |  | malloc returned NULL | No memory can be allocated for the structure holding all Destinations | COM_SWError
20487 | GET_PRECDATA | Error |  |  |  | getCPFData_prec_failed | Retrieving the PREC data from the CPF file failed | COM_NotNeeded_ForServiceId
20488 | NOMEM_PREC | Error |  |  |  | malloc returned NULL | No memory can be allocated for the structure holding all precheck destinations | COM_SWError
20489 | MAINTENANCE_EXPECTED | Warning | 120 | Maintenance interval will expire within a short time. Please contact service. |  | Maintenance Interval will expire within a short time. Please contact service. |  | COM_NotNeeded_ForServiceId
20490 | MAINTENANCE_REQUIRED | Warning | 258 | Maintenance interval expired. Please contact service. |  | Maintenance Interval is expired. Please contact service as soon as possible. |  | COM_NotNeeded_ForServiceId
20491 | GET_IDENDATA | Error |  |  |  | getCPFData_iden failed | Retrieving the IDEN data from the CPF file failed | COM_NotNeeded_ForServiceId
20992 | GROUPNAME_SMALLER_THAN_ZERO | Error |  |  |  | The group name is smaller than zero |  | COM_NotNeeded_ForServiceId
20993 | GROUPNAME_TOO_BIG | Error |  |  |  | The group name is too large |  | COM_NotNeeded_ForServiceId
20994 | NOBUFFER | Error |  |  |  | InitializeDataStore failed | Not enough memory could be allocated to hold the Data Store | COM_SWError
20995 | NOT_ENOUGH_MEMORY | Error |  |  |  | AddDataToStore failed | Not enough memory could be allocated | COM_SWError
20996 | NOT_INITIALIZED | Error |  |  |  | DataStore not initialized yet | Call 'InitializeDataStore' first. | COM_SWError
20997 | TIMEOUT | Error |  |  |  | lock already taken | A lock was taken and caused a waiting method to time out | COM_SWError
21504 | General | Error |  |  |  | General image processing error occurred | Meaning: A non-specific error occurred during image processing. Reason: Unknown. Cure: Unknown. | COM_SWError
21505 | MFS_FsLength | Error |  |  |  | Too few space for CRFileSystem. | Meaning: While trying to mount CRFileSystem, detected too few disk space. Reason: The flash memory on the photomultiplier unit is not existing or too small. Cure: Replace photomultiplier unit. | COM_PMMDataAccessErr
21506 | MFS_Mount | Error |  |  |  | Could not mount CRFileSystem. | Meaning: Could not mount CRFileSystem. Reason: . Cure: Perform a scan line calibration. | COM_PMMDataAccessErr
21507 | PRP_NoCRImage | Error |  |  |  | No valid CRImage given for preparing processing. | Meaning: The event to prepare processing contained no valid CRImage. Reason: Programming error. Either CRImage or one of its elements is a NULL pointer. Cure: Install another software version. | COM_SWError
21508 | PRP_NoDoseCalPIP | Error |  |  |  | Dose detection for PIP not calibrated. | Meaning: Dose detection for powder image plates not calibrated. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: NIP may be used. For PIP, replace photomultiplier module. | COM_PMMDataAccessErr
21509 | PRP_NoDoseCalNIP | Error |  |  |  | Dose detection for NIP not calibrated. | Meaning: Dose detection for needle image plates not calibrated. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: PIP may be used. For NIP, replace photomultiplier module. | COM_PMMDataAccessErr
21514 | PRO_NoCRImage | Error |  |  |  | No valid CRImage given for processing. | Meaning: The event to process an image contained no valid CRImage. Reason: Programming error. Either CRImage or one of its elements is a NULL pointer. Cure: Install another software version. | COM_SWError
21515 | PRO_WriteRaw | Error |  |  |  | Writing raw pixels to file failed. | Meaning: Writing raw pixels to file failed. Reason: Unknown. Cure: Ignore. | COM_NotNeeded_ForServiceId
21524 | INI_NoDoseCal | Error |  |  |  | Found no dose detection calibration at all. | Meaning: Found no dose detection calibration at all. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: Replace photomultiplier module. | COM_PMMDataAccessErr
21525 | INI_NoCRTools | Error |  |  |  | CRTools not started. | Meaning: At system start, the CRTools library was not found in memory. Reason: CRTools was not started. Cure: Install software again. | COM_SWError
21526 | INI_IncompatibleCRTools | Error |  |  |  | CRTools incompatible. | Meaning: The CRTools library is incompatible. Reason: The library CRTools does not fit to the rest of the application software. This is software configuration error and should never occur in the field. Cure: Try to install software again. If this doesn't help, install another software. | COM_SWError
21554 | IMB_INI_BufferCreation | Error |  |  |  | Creation of image buffer failed. | Meaning: Could not create a buffer for the image to scan. Reason: Memory for image buffer could not be allocated. Too few memory or problem with other parts of the software. Cure: Install the complete system again. If this fails, install another software version. Replace the CPU board. | COM_SWError
21555 | IMB_PRP_TooManyColumns | Error |  |  |  | Too many columns for preparing image buffer. | Meaning: At prepare operation, the number of coulumns and the block size have a wrong relation. Reason: Programming error. Cure: Install another software version. | COM_SWError
21556 | IMB_GNW_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21557 | IMB_GNP_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21558 | IMB_GNP_Timeout | Error |  |  |  | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21559 | IMB_GNR_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21560 | IMB_GNR_Timeout | Error |  |  |  | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21561 | IMB_GAR_Cancelled | Error |  |  |  | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it. | COM_NotNeeded_ForServiceId
21562 | IMB_GAR_Timeout | Error |  |  |  | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21563 | IMB_GAR_Occupied | Error |  |  |  | All read blocks occupied. | Meaning: Could not get a new image block to read. All occupied. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21564 | IMB_RAR_InvalidID | Error |  |  |  | Tried to release an image block with an invalid ID. | Meaning: Tried to release an image block, which has not been reserved. Reason: Programming error. Cure: Ignore this error, if it does not cause any problem. | COM_NotNeeded_ForServiceId
21565 | IMB_GNW_Overflow | Error |  |  |  | Image buffer overflow | Meaning: The image buffer has not been read out fast enough, so that it is not ready for further image data. Scanning has been stopped. Reason: Image data was not sent or stored fast enough. Probably too slow flash memory. Cure: Try if replacing the flash memory by a faster one helps. | COM_SWError
21566 | IMB_GAR_Released | Error |  |  |  | Image data no longer available. | Meaning: Image data is no longer available. The memory location is reused in circular buffer. Reason: Could be programming error. Cure: Install another software version. | COM_SWError
21567 | IMB_RBL_NoData | Error |  |  |  | Attempt to release more data than available. | Meaning: Attempt to release more data than available. Corrected to actual value. Reason: Programming error. Cure: Only warning, should be harmless. In case of problems install another software version. | COM_NotNeeded_ForServiceId
21604 | ISP_SignalRejected | Error |  |  |  | The incoming signal was rejected in this state. | Meaning: Image splitter got a signal in a state where it cannot be handled. Reason: Programming error. Cure: In case of problems install another software version. | COM_SWError
22016 | SYNCHRONIZING_FAILED | Error |  |  |  | Synchronizing failed.
22017 | DETECT_CONTROLLERS_FAILED | Error |  |  |  | detectControllers failed. | Meaning: detectControllers failed. Reason: Not all controllers were detected. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels.
22018 | DETECT_ELABELS_FAILED | Error |  |  |  | detectELabels failed. | Meaning: detectELabels failed. Reason: Not all eLabels were detected. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels.
22019 | ELABEL_MISSING | Error |  |  |  | Not all controllers and/or eLabels are connected. | Meaning: Not all controllers and/or eLabels are connected. Reason: Not all controllers and/or eLabels were detected. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels. | ELC_ELabelMissing
22020 | READ_ALL_ELABELS_FAILED | Error |  |  |  | readAllEntries failed. | Meaning: readAllEntries failed. Reason: Not all entries of all eLabels could be read. Cure: 1) Check all cables of the eLabels. 2) Replace eLabels.
22021 | DATASTOREDATA_FILE_NOT_AVAILABLE | Error |  |  |  | DataStoreData file does not exist. | Meaning: DataStoreData file does not exist. Reason: DataStoreData file not stored on CF card. Cure: Please restore this file via Restore.
22022 | DATASTOREDATA_FILE_CORRUPT | Error |  |  |  | DataStoreData file is corrupt. | Meaning: DataStoreData file is corrupt. Reason: DataStoreData file could not opened. Cure: Please restore this file via Restore.
22023 | OrderRepairToolForOptics | Error |  |  |  | Parameter set of optic module is corrupted. | Meaning: Parameter set of optic module is corrupted. Reason: The structure of the eLabel on the optic module is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted optic module.
22024 | OrderRepairToolForPMT | Error |  |  |  | Parameter set of PMT/light collector module is corrupted. | Meaning: Parameter set of PMT/light collector module is corrupted. Reason: The structure of the eLabel on PMT is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted PMT/light collector module.
22025 | OrderRepairToolForHandling | Error |  |  |  | Parameter set of handling board is corrupted. | Meaning: Parameter set of handling board is corrupted. Reason: The structure of the eLabel on the handling board is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted handling board.
22026 | OrderRepairToolForScanner | Error |  |  |  | Parameter set of scanner module is corrupted. | Meaning: Parameter set of scanner module is corrupted. Reason: The structure of the eLabel on the scanner module is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted scanner module.
22027 | OrderRepairToolForCassClamping | Error |  |  |  | Parameter set of cassette clamping module is corrupted. | Meaning: Parameter set of cassette clamping module is corrupted. Reason: The structure of the eLabel on the cassette clamping module is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted cassette clamping module.
22028 | AUTOMATIC_REGISTRATION_FAILED | Error |  |  |  | Automatic Registration failed. | Meaning: Automatic Registration failed.
22029 | REPAIR_ELABEL_FAILED | Error |  |  |  | Repair of corrupted eLabel failed. | Meaning: Repair of corrupted eLabel failed.
22030 | WrongOpticsModuleConfiguration | Error |  |  |  | Wrong Optics module configuration. | Meaning: Wrong Optics module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace Optics module.
22031 | WrongPMTModuleConfiguration | Error |  |  |  | Wrong PMT module configuration. | Meaning: Wrong PMT module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace PMT module.
22032 | WrongScannerModuleConfiguration | Error |  |  |  | Wrong Scanner module configuration. | Meaning: Wrong Scanner module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace Scanner module.
22033 | WrongCassClampingModuleConfiguration | Error |  |  |  | Wrong Cassette Clamping module configuration. | Meaning: Wrong Cassette Clamping module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace Cassette Clamping module.
22034 | GenradPMTInMammoDigitizer | Error |  |  |  | Genrad PMT replaced in mammo digitizer. | Meaning: Genrad PMT replaced in mammo digitizer. Reason: Module name is not PMT-M and not allowed for mammo digitizer. Only PMTs with module name PMT-M are allowed. Cure: Replace PMT. | ELC_GenradPMTInMammoDig | CIR318 | Incorrect PMT with light collector for GenRad only detected during start-up Meaning: PMT with light collector for Genrad only application detected during start-up, which is incompatible with this device version. The Digitizer is out of operation and requires service support. Reason: 1) PMT with light collector for Genrad only application has been mounted, which is incompatible with this device version. Cure: 1) Replace the detected PMT with light collector for Genrad only application by a compatible version for Mammography application, refer to the Service Manual for order and replacement
22035 | SparePartExistsNoDataStoreDataAvailable | Error |  |  |  | Sparepart(Module serialnumber 9999) exists but no data are available. | Meaning: Sparepart(Module serialnumber 9999) exists. Reason: Sparepart(Module serialnumber 9999) exists, but no registration was made. Cure: Replace sparepart with the old board and restart digitizer.
22036 | GenradGearboxInMammoDigitizer | Error |  |  |  | Genrad Gearbox replaced in mammo digitizer. | Meaning: Genrad Gearbox replaced in mammo digitizer. Reason: Module name is not Gearbox-M and not allowed for mammo digitizer. Only Gearboxes with module name Gearbox-M are allowed. Cure: Replace Gearbox. | ELC_GenradGearboxInMammoDig | CIR258 | Incorrect Gear Box with Scan Rollers for GenRad only detected during start-up Meaning: Gear Box with Scan Rollers for Genrad only application detected during start-up, which is incompatible with this device version. The Digitizer is out of operation and requires service support. Reason: 1) Gear Box with Scan Rollers for Genrad only application has been mounted, which is incompatible with this device version. Cure: 1) Replace the detected Gear Box with Scan Rollers for Genrad only application by a compatible version for Mammography application, refer to the Service Manual for order and replacement
22037 | MammoGearboxInGenradDigitizer | Error |  |  |  | Mammo Gearbox replaced in genrad digitizer. | Meaning: Mammo Gearbox replaced in genrad digitizer. Reason: Module name is not Gearbox-G and not allowed for genrad digitizer. Only Gearboxes with module name Gearbox-G are allowed. Cure: Replace Gearbox. | ELC_MammoGearboxInGenradDig | CIR257 | Incorrect Gear Box with Scan Rollers for Mammography detected during start-up Meaning: Gear Box with Scan Rollers for Mammography application detected during start-up, which is incompatible with this device version. The Digitizer is out of operation and requires service support. Reason: 1) Gear Box with Scan Rollers for Mammography application has been mounted, which is incompatible with this device version. Cure: 1) Replace the detected Gear Box with Scan Rollers for Mammography application by a compatible version for Genrad only application, refer to the Service Manual for order and replacement
22038 | GearboxFromOtherDigitizer | Error |  |  |  | Gearbox replaced from another digitizer. | Meaning: Gearbox replaced from another digitizer. Reason: Module name of Gearbox is not set and is from another digitizer. Only Gearboxes with module name Gearbox-G and/or Gearbox-M are allowed. Cure: Replace Gearbox. | ELC_GearboxFromOtherDig | CIR259 | Unknown module name in the E-Label Gear Box with Motor and Scan Rollers Meaning: Module name in the E-Label of this assembly detected, which is either not defined or invalid. The Digitizer is out of operation and requires service support. Reason: 1) The module name, which is stored in the E-Label attached to this Gear Box with Motor and Scan Rollers is either not defined or invalid. The software requires this module name to detect the version of this assembly. Cure: Case A): A new spare part Gear Box with Motor and Scan Rollers was installed: 1) Remove the installed spare part Gear Box with Motor and Scan Rollers and return it according to the return procedure of spare parts, see Return Label attached Case B): A used Gear Box with Motor and Scan Rollers e.g. from another device was installed: 1) Remove the installed Gear Box with Motor and Scan Rollers and mount it in the origin device 2) Install in the origin device the software CIR_13xx / CIRM13xx or higher 3) Swap again the used Gear Box with Motor and Scan Rollers from the origin device in the device to be repaired 4) If this swap fails again the message 'CIR259', order a new spare part Gear Box with Motor and Scan Rollers, refer to the Service Manual for order and replacement
22039 | MammoPMTInGenradDigitizer | Error |  |  |  | Mammo PMT replaced in genrad digitizer. | Meaning: Mammo PMT replaced in genrad digitizer. Reason: Module name is PMT-M and not allowed for genrad digitizer. Only PMTs with module name different from PMT-M are allowed. Cure: Replace PMT. | ELC_MammoPMTInGenradDig | CIR317 | Incorrect PMT with light collector for Mammography detected during start-up Meaning: PMT with light collector for Mammography application detected during start-up, which is incompatible with this device version. The Digitizer is out of operation and requires service support. Reason: 1) PMT with light collector for Mammography application has been mounted, which is incompatible with this device version. Cure: 1) Replace the detected PMT with light collector for Mammography application by a compatible version for Genrad only application, refer to the Service Manual for order and replacement
22040 | GenradOpticsInMammoDigitizer | Error |  |  |  | Genrad Optics replaced in mammo digitizer. | Meaning: Genrad Optics replaced in mammo digitizer. Reason: Optics type is not 2 and not allowed for mammo digitizer. Only Optics with Optics type 2 are allowed. Cure: Replace Optics. | ELC_GenradOpticsInMammoDig | CIR278 | Incorrect optic module for GenRad only detected during start-up. Meaning: Optic module for Genrad only application detected during start-up, which is incompatible with this device version. The Digitizer is out of operation and requires service support. Reason: 1) Optic module for Genrad only application has been mounted, which is incompatible with this device version. Cure: 1) Replace the detected Optic module for Genrad only application by a compatible version for Mammography application, refer to the Service Manual for order and replacement
22041 | MammoOpticsInGenradDigitizer | Error |  |  |  | Mammo Optics replaced in genrad digitizer. | Meaning: Mammo Optics replaced in genrad digitizer. Reason: Optics type is not 0 or 1 and not allowed for genrad digitizer. Only Optics with optics type 0 or 1 are allowed. Cure: Replace Optics. | ELC_MammoOpticsInGenradDig | CIR277 | Incorrect optic module for Mammography detected during start-up. Meaning: Optic module for Mammography application detected during start-up, which is incompatible with this device version. The Digitizer is out of operation and requires service support. Reason: 1) Optic module for Mammography application has been mounted, which is incompatible with this device version. Cure: 1) Replace the detected Optic module for Mammography application by a compatible version for Genrad only application, refer to the Service Manual for order and replacement
22042 | GearboxELabelMissing | Error |  |  |  | Gearbox eLabel is missing. | Meaning: Gearbox eLabel is missing. Reason: Gearbox eLabel is missing or defect. Cure: 1) Check eLabel on Gearbox. 2) Check cable for Gearbox eLabel to Handling board. | ELC_ELabelGearBoxMissing | CIR256 | Data access to E-Label Gear Box with Motor and Scan Rollers failed. Meaning: Data access of the Handling Control Board to the E-Label Gear Box with Motor and Scan Rollers failed. The Digitizer is out of operation and requires service support. . Reason: 1) Contact problem or defect at data cable between Handling Control Board and E-Label Gear Box with Motor and Scan Rollers (B425) 2) Contact problem or defect on the E-Label Gear Box with Motor and Scan Rollers (B425) 3) Contact problem (connector P3) or defect on the Handling Control Board Cure: 1) Check the data cable between Handling Control Board and E-Label Gear Box with Motor and Scan Rollers (B425). If required replace the data cable, contact your Regional Support Network for further assistance 2) Check connector on the E-Label Gear Box with Motor and Scan Rollers (B425) 3) Check connector P3 on the Handling Control Board 4) Exchange the Gear Box with Motor and Scan Rollers, refer to the Service Manual for order and replacement 5) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
22043 | CassetteFixationUnitELabelMissing | Error |  |  |  | Cassette Fixation Unit eLabel is missing. | Meaning: Cassette Fixation Unit eLabel is missing. Reason: Cassette Fixation Unit eLabel is missing or defect. Cure: 1) Check eLabel on Cassette Fixation Unit. 2) Check cable for Cassette Fixation Unit eLabel to Handling board. | ELC_ELabelCassFixUnitMissing | CIR586 | Data access to E-Label Cassette Fixation Unit failed Meaning: Data access of the Handling Control Board to the E-Label Cassette Fixation Unit failed. The Digitizer is out of operation and requires service support. Reason: 1) Contact problem or defect at data cable between Handling Control Board and E-Label Cassette Fixation Unit (B429) 2) Contact problem or defect on the E-Label Cassette Fixation Unit (B429) 3) Contact problem (connector P3) or defect on the Handling Control Board Cure: 1) Check the data cable between Handling Control Board and E-Label Cassette Fixation Unit (B429). If required replace the data cable, contact your Regional Support Network for further assistance 2) Check connector on the E-Label Cassette Fixation Unit (B429) 3) Check connector P3 on the Handling Control Board 4) Exchange the Cassette Fixation Unit, refer to the Service Manual for order and replacement 5) Exchange the Handling Control Board, refer to the Service Manual for order and replacement
22066 | HDL_ELABEL_INIT_FAILED | Error |  |  |  | Initialisation of the Elabel on the Handlingboard failed. | Meaning: Failed selftest and initialisation of data storage at the HdlBoard Elabel. Reason: No successful communication to the Elabel during initialisation Cure: 1) Check the HdlBoard for loose cables 2) Upgrade SW 3) Replace HdlBoad
22067 | HDL_ELABEL_INIT_CHECK_ENTRIES_FAILED | Error |  |  |  | Initialisation of the Elabel with check of entries on the Handlingboard failed. | Meaning: Failed initialisation with check of entries at the HdlBoard Elabel. Reason: Elabel entries missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22068 | SLOW_SCAN_GEARBOX_ELABEL_INIT_FAILED | Error |  |  |  | Initialisation of the Elabel at the Slow Scan Gearbox failed. | Meaning: Failed selftest and initialisation of data storage at the Slow Scan Gearbox Elabel. Reason: No successful communication to the Elabel during initialisation Cure: 1) Check cable from the HdlBoard to the Slow Scan Gearbox Elabel 2) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Slow Scan Gearbox with its Elabel
22069 | CASSETTE_SLOT_ELABEL_INIT_FAILED | Error |  |  |  | Initialisation of the Elabel at the Cassette Slot failed. | Meaning: Failed selftest and initialisation of data storage at the Cassette Slot Elabel. Reason: No successful communication to the Elabel during initialisation Cure: 1) Check cable from the HdlBoard to the Cassette Slot Elabel 2) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Cassette Slot unit with its Elabel
22070 | CASSETTE_SLOT_ELABEL_INIT_CHECK_ENTRIES_FAILED | Error |  |  |  | Initialisation of the Elabel with check of entries at the Cassette Slot failed. | Meaning: Failed initialisation with check of entries at the Cassette Slot Elabel. Reason: Elabel entries missing or defect. Cure: 1) Restart machine 2) Check Elabel at Cassette Slot connected to Hdl board 3) Configure Elabel at Cassette Slot 4) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Cassette Slot unit with its Elabel
22071 | SLOW_SCAN_GEARBOX_ELABEL_INIT_CHECK_ENTRIES_FAILED | Error |  |  |  | Initialisation of the Elabel with check of entries at the Slow Scan Gearbox failed. | Meaning: Failed initialisation with check of entries at the Slow Scan Gearbox. Reason: Elabel entries missing or defect. Cure: 1) Restart machine 2) Check Elabel at Slow Scan Gearbox connected to Hdl board 3) Configure Elabel at Slow Scan Gearbox 4) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Slow Scan Gearbox unit with its Elabel
22076 | FIXER_IRUN_ELABEL_INIT_TO_DEFAULT | Error |  |  |  | Initialisation of the Elabel entry Fixer_PreClamp_iRun required creation with default value. | Meaning: Elabel entry Fixer_PreClamp_iRun is missing at initialisation and the error handling required creation with default value. Reason: Elabel entry missing or defect. Cure: 1) Perform stall calibration of SM Fixer
22077 | LOCKER_IRUN_ELABEL_INIT_TO_DEFAULT | Error |  |  |  | Initialisation of the Elabel entry Locker_UnLockIp_iRun required creation with default value. | Meaning: Elabel entry Locker_UnLockIp_iRun is missing at initialisation and the error handling required creation with default value. Reason: Elabel entry missing or defect. Cure: 1) Perform stall calibration of SM Locker
22078 | ROLLER_CLAMPING_IRUN_ELABEL_INIT_TO_DEFAULT | Error |  |  |  | Initialisation of the Elabel entry RollerClamping_iRun_Adjustable required creation with default value. | Meaning: Elabel entry RollerClamping_iRun_Adjustable is missing at initialisation and the error handling required creation with default value. Reason: Elabel entry missing or defect. Cure: 1) Perform stall calibration of SM Roller Clamping
22079 | DRAWER_IRUN_ELABEL_INIT_TO_DEFAULT | Error |  |  |  | Initialisation of the Elabel entry Drawer_iRun_Adjustable required creation with default value. | Meaning: Elabel entry Drawer_iRun_Adjustable is missing at initialisation and the error handling required creation with default value. Reason: Elabel entry missing or defect. Cure: 1) Perform stall calibration of SM Drawer
22080 | LOCKER_GET_IRUN_FAILED | Error |  |  |  | Get Elabel entry Locker_UnLockIp_iRun failed. | Meaning: Access to get Elabel entry Locker_UnLockIp_iRun value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22081 | LOCKER_SET_IRUN_FAILED | Error |  |  |  | Set Elabel entry Locker_UnLockIp_iRun failed. | Meaning: Access to set Elabel entry Locker_UnLockIp_iRun value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22082 | FIXER_GET_IRUN_FAILED | Error |  |  |  | Get Elabel entry Fixer_PreClamp_iRun failed. | Meaning: Access to get Elabel entry Fixer_PreClamp_iRun value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22083 | FIXER_SET_IRUN_FAILED | Error |  |  |  | Set Elabel entry Fixer_PreClamp_iRun failed. | Meaning: Access to set Elabel entry Fixer_PreClamp_iRun value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22084 | DRAWER_GET_IRUN_FAILED | Error |  |  |  | Get Elabel entry Drawer_iRun_Adjustable failed. | Meaning: Access to get Elabel entry Drawer_iRun_Adjustable value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22085 | DRAWER_SET_IRUN_FAILED | Error |  |  |  | Set Elabel entry Drawer_iRun_Adjustable failed. | Meaning: Access to set Elabel entry Drawer_iRun_Adjustable value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22086 | ROLLER_CLAMPING_GET_IRUN_FAILED | Error |  |  |  | Get Elabel entry RollerClamping_iRun_Adjustable failed. | Meaning: Access to get Elabel entry RollerClamping_iRun_Adjustable value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22087 | ROLLER_CLAMPING_SET_IRUN_FAILED | Error |  |  |  | Set Elabel entry RollerClamping_iRun_Adjustable failed. | Meaning: Access to set Elabel entry RollerClamping_iRun_Adjustable value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22088 | DRAWER_INC_RETRIES_INTO_CASSETTE_FAILED | Error |  |  |  | Increment Elabel entry Drawer_NumRetries_IntoCassette failed. | Meaning: Access to increment Elabel entry Drawer_NumRetries_IntoCassette value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22089 | DRAWER_INC_RETRIES_OUT_OF_CASSETTE_FAILED | Error |  |  |  | Increment Elabel entry Drawer_NumRetries_OutOfCassette failed. | Meaning: Access to increment Elabel entry Drawer_NumRetries_OutOfCassette value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22090 | DRAWER_INC_RETRIES_TO_REFERENCE_FAILED | Error |  |  |  | Increment Elabel entry Drawer_NumRetries_ToReference failed. | Meaning: Access to increment Elabel entry Drawer_NumRetries_ToReference value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22091 | ROLLER_CLAMPING_INC_RETRIES_CLAMP_IP_FAILED | Error |  |  |  | Increment Elabel entry RollerClamping_NumRetries_ClampIP failed. | Meaning: Access to increment Elabel entry RollerClamping_NumRetries_ClampIP value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22092 | ROLLER_CLAMPING_INC_RETRIES_FREE_IP_FAILED | Error |  |  |  | Increment Elabel entry RollerClamping_NumRetries_FreeIP failed. | Meaning: Access to increment Elabel entry RollerClamping_NumRetries_FreeIP value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22093 | FIXER_INC_RETRIES_UNCLAMP_FAILED | Error |  |  |  | Increment Elabel entry Fixer_NumRetries_Unclamp failed. | Meaning: Access to increment Elabel entry Fixer_NumRetries_Unclamp value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22094 | FIXER_INC_RETRIES_NUM_WRONG_FORMAT_CLAMPED_FAILED | Error |  |  |  | Increment Elabel entry Fixer_NumWrongFormat_Clamped failed. | Meaning: Access to increment Elabel entry Fixer_NumWrongFormat_Clamped value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22095 | LOCKER_INC_RETRIES_LOCK_IP_FAILED | Error |  |  |  | Increment Elabel entry Locker_NumRetries_LockIp failed. | Meaning: Access to increment Elabel entry Locker_NumRetries_LockIp value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22096 | LOCKER_INC_RETRIES_UNLOCK_IP_FAILED | Error |  |  |  | Increment Elabel entry Locker_NumRetries_UnLockIp failed. | Meaning: Access to increment Elabel entry Locker_NumRetries_UnLockIp value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22097 | OPENER_INC_RETRIES_OPEN_COVER_FAILED | Error |  |  |  | Increment Elabel entry Opener_NumRetries_OpenCover failed. | Meaning: Access to increment Elabel entry Opener_NumRetries_OpenCover value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22098 | OPENER_INC_RETRIES_CLOSE_COVER_FAILED | Error |  |  |  | Increment Elabel entry Opener_NumRetries_CloseCover failed. | Meaning: Access to increment Elabel entry Opener_NumRetries_CloseCover value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22099 | SEQUENCE_INC_RETRIES_ERASE_IP_FAILED | Error |  |  |  | Increment Elabel entry Sequence_NumRetries_EraseIP failed. | Meaning: Access to increment Elabel entry Sequence_NumRetries_EraseIP value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22100 | SEQUENCE_INC_RETRIES_FEED_IP_FAILED | Error |  |  |  | Increment Elabel entry Sequence_NumRetries_FeedIP failed. | Meaning: Access to increment Elabel entry Sequence_NumRetries_FeedIP value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22101 | CASS_SLOT_ELABEL_INIT_CASS_SLOT_WIDTH_ENTRY_NOT_FOUND | Error |  |  |  | Initialisation of the Elabel entry CassSlotWidth is missing. | Meaning: Failed initialisation with check of entries at the Cassette Slot Elabel. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel at Cassette Slot connected to Hdl board 3) Configure Elabel at Cassette Slot 4) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Cassette Slot unit with its Elabel
22102 | CASS_SLOT_GET_CASS_SLOT_WIDTH_FAILED | Error |  |  |  | Get Elabel entry CassSlotWidth failed. | Meaning: Access to get Elabel entry CassSlotWidth value failed. Reason: Elabel entry missing or defect. Cure: 1) Check cable from the HdlBoard to the Cassette Slot Elabel 2) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Cassette Slot unit with its Elabel
22103 | CASS_SLOT_SET_CASS_SLOT_WIDTH_FAILED | Error |  |  |  | Set Elabel entry CassSlotWidth failed. | Meaning: Access to set Elabel entry CassSlotWidth value failed. Reason: Elabel entry missing or defect. Cure: 1) Check cable from the HdlBoard to the Cassette Slot Elabel 2) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Cassette Slot unit with its Elabel
22104 | SSC_GEARBOX_ELABEL_INIT_SPEED_FACTOR_ENTRY_NOT_FOUND | Error |  |  |  | Initialisation of the Elabel entry SpeedFactor is missing. | Meaning: Failed initialisation with check of entries at the Slow Scan Gearbox Elabel. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel at Slow Scan Gearbox connected to Hdl board 3) Configure Elabel at Slow Scan Gearbox 4) Upgrade SW 3) Replace cable 4) Replace HdlBoad 5) Replace Slow Scan Gearbox unit with its Elabel
22105 | SEQUENCE_INC_RETRIES_RETURN_IP_FAILED | Error |  |  |  | Increment Elabel entry Sequence_NumRetries_ReturnIP failed. | Meaning: Access to increment Elabel entry Sequence_NumRetries_ReturnIP value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22106 | FIXER_GET_REPETITION_FAILED | Error |  |  |  | Retrieval of theElabel entry Fixer_PreClamp_Repetition failed. | Meaning: Access to read Elabel entry Fixer_PreClamp_Repetition value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22107 | FIXER_SET_REPETITION_FAILED | Error |  |  |  | Setting of theElabel entry Fixer_PreClamp_Repetition failed. | Meaning: Access to set Elabel entry Fixer_PreClamp_Repetition value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22108 | FIXER_REPETITION_ENTRY_AT_ELABEL_INIT_TO_DEFAULT | Error |  |  |  | Initialisation of the Elabel entry Fixer_PreClamp_Repetition required creation with default value. | Meaning: Elabel entry Fixer_PreClamp_Repetition is missing at initialisation and the error handling required creation with default value. Reason: Elabel entry missing or defect. Cure: 1) Perform stall calibration of SM Fixer
22109 | SLOWSCAN_INC_RETRIES_Bring_BACK_IP_FAILED | Error |  |  |  | Increment Elabel entry Slowscan_NumRetries_BringBackIP failed. | Meaning: Access to increment Elabel entry Slowscan_NumRetries_BringBackIP value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22110 | OPENER_GET_OPENSENSORTYPE_FAILED | Error |  |  |  | Get Elabel entry Opener_OpenSensorType_IsLB failed. | Meaning: Access to get Elabel entry Opener_OpenSensorType_IsLB value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22111 | OPENER_SET_OPENSENSORTYPE_FAILED | Error |  |  |  | Set Elabel entry Opener_OpenSensorType_IsLB failed. | Meaning: Access to set Elabel entry Opener_OpenSensorType_IsLB value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22112 | OPENER_GET_OPENSTEPSFROMLB_FAILED | Error |  |  |  | Get Elabel entry Opener_OpenSteps_FromLB failed. | Meaning: Access to get Elabel entry Opener_OpenSteps_FromLB value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22113 | OPENER_SET_OPENSTEPSFROMLB_FAILED | Error |  |  |  | Set Elabel entry Opener_OpenSteps_FromLB failed. | Meaning: Access to get Elabel entry Opener_OpenSteps_FromLB value failed. Reason: Elabel entry missing or defect. Cure: 1) Restart machine 2) Check Elabel on Hdl board 3) Configure Elabel on Hdl board
22528 | CRT_Base | Error |  |  |  | General error occurred in CRTools library | Meaning: A non-specific error occurred in the CRTools library. Reason: This error should normally not occur. Its value is used as base for all other CRTools errors. Cure: Ignore this error. | COM_SWError
22529 | PXC_InsufficientCalWidth | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Calibration line does not cover the complete width of the image to be corrected. Parts of the image will not be corrected. | Meaning: Only parts of the image will be corrected. The rest might have banding artefacts. Reason: Calibration line does not cover the complete width of the image to be corrected. Cure: Calibrate with the biggest used cassette. | COM_SWShadingCalibration
22530 | PXC_CalcCorrLines | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Calculating the correction lines failed. Using constant correction line. | Meaning: The image is uncorrected. Reason: An exception occurred during preparation of the correction line. Might be software bug. Cure: Check the log file to find out more information about the error. | COM_SWShadingCalibration
22531 | PXC_NoCalibrationLine | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | No calibration line loaded to correct the image. Using constant correction line. | Meaning: The image is uncorrected. Reason: Calibration line of this image plate type not available, invalid or could not be loaded. Cure: Calibrate the image plate type, which has generated the error. | COM_SWShadingCalibration
22532 | PXC_InvalidIPType | Error |  |  |  | An invalid IPType was given when clearing the calibration line. |  | COM_SWError
22533 | PXC_ExtrapolateCorrLine | Error | 257 | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Extrapolating the correction lines failed. Using constant correction line. | Meaning: The image is uncorrected. Reason: An exception occurred during preparation of the correction line. Might be software bug. Cure: Check the log file to find out more information about the error. | COM_SWShadingCalibration
22534 | PXC_UnsupportedSamplesOrBPP | Error |  |  |  | Unsupported samples or bytes per pixel. |  | COM_SWError
22535 | PXC_UnsupportedSamplesOrBPP2 | Error |  |  |  | Unsupported samples or bytes per pixel. |  | COM_SWError
22536 | PXC_UnsupportedColumns | Error |  |  |  | Unsupported number of pixels per line. | Meaning: The requested number of pixels per line is negative or too small. Reason: Software bug. Cure: Install another software. | COM_SWError
22537 | PXC_UnsupportedLines | Error |  |  |  | Unsupported number of lines. | Meaning: The requested number lines is negative or too small. Reason: Software bug. Cure: Install another software. | COM_SWError
22538 | LCL_CreateCalibrationLine | Error |  |  |  | Creating a new CalibrationLine failed. |  | COM_SWError
22539 | LCL_WriteToCRFile | Error |  |  |  | Writing CalibrationLine to PMM failed. |  | COM_SWError
22540 | DCL_CreateCalibrationLine | Error |  |  |  | Creating a new dose CalibrationLine failed. |  | COM_SWError
22541 | DCL_WriteToCRFile | Error |  |  |  | Writing dose CalibrationLine to PMM failed. |  | COM_SWError
22542 | PXC_LoadIPCalNullResult | Error |  |  |  | PixelCorrector_LoadIPCal called with result null pointer. | Meaning: Null pointer in fuction call. Cannot execute. Reason: Software error. Cure: Install new software. | COM_SWError
22543 | PXC_CorrectLinesParameter | Error |  |  |  | PixelCorrector_CorrectLines called with invalid parameter. | Meaning: PixelCorrector_CorrectLines called with invalid parameter. Reason: Software error. Cure: Install new software. | COM_SWError
22544 | PXC_EOIBlocksPending | Error |  |  |  | There are still image blocks pending at end of image. Software error. |  | COM_SWError
22548 | CIM_UnknownImageFormat | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Reading IP calibration image failed because of unknown image file format. |  | COM_Image2DCalibrationErr
22549 | CIM_ImageNotLoaded | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Reading IP calibration image failed. |  | COM_Image2DCalibrationErr
22550 | CIM_ReadImageMemoryException | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Exception at reading IP calibration image from ID tag. |  | COM_Image2DCalibrationErr
22551 | CIM_ReadImageFileNameMismatch | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Wrong IPSerialNumber in IP calibration file. | Meaning: The name of the IP calibration image is different than the IPSerialNumber, stored in the file. Reason: Probably manual renaming of the calibration file. Cure: Never use an IP calibration file for another image plate. | COM_Image2DCalibrationErr
22552 | CIM_ReadImageFileException | Error |  |  |  | Exception at reading IP calibration image from file. |  | COM_Image2DCalibrationErr
22553 | CIM_WriteImageMemoryException | Error |  |  |  | Exception at writing IP calibration image to memory for storing on ID tag. |  | COM_Image2DCalibrationErr
22554 | CIM_WriteImageFileException | Error |  |  |  | Exception at writing IP calibration image to file. |  | COM_Image2DCalibrationErr
22555 | CIM_WriteNoImage | Error |  |  |  | Writing IP calibration image failed because no image present. Software error. |  | COM_Image2DCalibrationErr
22558 | CIM_RescaleNoImage | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | No image to rescale. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration or IP correction was disabled for this scan. Cure: Return the image plate to production for calibration. | COM_Image2DCalibrationErr
22559 | CIM_RescaleIPCalDotsPerMeter | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | The IP calibration image (TIFF file) contains wrong values for DotsPerMeter. | Meaning: Parameter error in the image plate calibration image. Reason: This could be prototype plate from DX-S development. But usually, this problem should not occur. Cure: Delete the calibration image from the hard disk and return the image plate to production for calibration. | COM_Image2DCalibrationErr
22560 | CIM_RescaleIPCalFailed | Error |  |  |  | Rescaling the given IP calibration image failed. |  | COM_Image2DCalibrationErr
22561 | CIM_RescaleException | Error |  |  |  | Exception at rescaling IP calibration image. |  | COM_Image2DCalibrationErr
22568 | CIM_ShiftAddCutNoCal | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | No IP calibration image available. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration. Cure: Return the image plate to production for calibration. | COM_Image2DCalibrationErr
22569 | CIM_ShiftAddCutException | Error |  |  |  | Exception at shifting / extending / cutting IP calibration image. |  | COM_Image2DCalibrationErr
22570 | CIM_ShiftAddCutTopExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many lines on top of the calibration image had to be extrapolated. | Meaning: Too many lines on top of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22571 | CIM_ShiftAddCutBottomExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many lines at the bottom of the calibration image had to be extrapolated. | Meaning: Too many lines at the bottom of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22572 | CIM_ShiftAddCutLeftExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many columns at the left side of the calibration image had to be extrapolated. | Meaning: Too many columns at the left side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22573 | CIM_ShiftAddCutRightExtra | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many columns at the right side of the calibration image had to be extrapolated. | Meaning: Too many columns at the right side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22574 | CIM_AddIPCalFile | Error |  |  |  | Could not load IPCalFile. |  | COM_Image2DCalibrationErr
22575 | CIM_AddInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22576 | CIM_WriteMemoryReadFile | Error |  |  |  | Could not read IPCalFile to write to memory. |  | COM_Image2DCalibrationErr
22577 | CIM_WriteMemoryInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22578 | CIM_WriteFileInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22579 | CIM_WriteFileRescaledInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22580 | CIM_RescaleAndFlipException | Error |  |  |  | Error occurred while rescaling a calibration image. |  | COM_Image2DCalibrationErr
22581 | CIM_WriteFileTrailerRead | Error |  |  |  | Could not read IPCalFile trailer. |  | COM_Image2DCalibrationErr
22582 | CIM_WriteFileTrailerWrite | Error |  |  |  | Could not write IPCalFile trailer. |  | COM_Image2DCalibrationErr
22583 | CIM_ReadFileTrailerInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22584 | CIM_HeaderVersion | Error |  |  |  | Wrong IPCalibration header version. |  | COM_Image2DCalibrationErr
22585 | CIM_HeaderTruncated | Error |  |  |  | IPCal header truncated. |  | COM_Image2DCalibrationErr
22586 | CIM_DataTruncated | Error |  |  |  | IPCalibration data truncated. |  | COM_Image2DCalibrationErr
22587 | CIM_HeaderSize | Error |  |  |  | Wrong IPCal header size. |  | COM_Image2DCalibrationErr
22588 | CIM_HeaderType | Error |  |  |  | Wrong IPCal header type. |  | COM_Image2DCalibrationErr
22589 | CIM_IPCalHeaderVersion | Error |  |  |  | Wrong IPCal header version. |  | COM_Image2DCalibrationErr
22590 | CIM_IPCalHeaderTruncated | Error |  |  |  | IPCalFile header truncated. |  | COM_Image2DCalibrationErr
22591 | CIM_IPCalDataTruncated | Error |  |  |  | Not enough IPCalFile data. |  | COM_Image2DCalibrationErr
22592 | CIM_IPCalHeaderSize | Error |  |  |  | Wrong IPCalFile header size. |  | COM_Image2DCalibrationErr
22593 | CIM_Checksum | Error |  |  |  | IPCalibrationData checksum error. |  | COM_Image2DCalibrationErr
22594 | CIM_FileFormat | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Unknown image file format. |  | COM_Image2DCalibrationErr
22595 | CIM_SerialNumberMismatch | Error |  |  |  | The image file name does not match the IP serial number. |  | COM_Image2DCalibrationErr
22596 | CIM_WriteMNotValid | Error |  |  |  | The IPCalibrationData to be written to tag is not valid. | Meaning: The IP calibration data to be written to the tag is not valid. Reason: Probably software error. Cure: Install new software. | COM_Image2DCalibrationErr
22597 | CIM_WriteMTooBig | Error |  |  |  | The IPCalData to be written to tag is too big. | Meaning: Detected that the given IP calibration data does not fit onto the tag. Reason: Probably software error. Cure: Install new software. | COM_Image2DCalibrationErr
22598 | CIM_WriteMNotValid2 | Error |  |  |  | The IPCalibrationData to be written to tag is not valid. |  | COM_Image2DCalibrationErr
22599 | CIM_WriteNotValid | Error |  |  |  | The IPCalFile data to be written to file is not valid. |  | COM_Image2DCalibrationErr
22600 | CIM_NoFile | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Could not open CalibrationImage file. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: . Cure: Return the image plate to production for new calibration and check of the serial number. | COM_Image2DCalibrationErr
22601 | CIM_IPCalDataVersion | Error |  |  |  | Wrong IPCalData version. |  | COM_Image2DCalibrationErr
22602 | CIM_IPCalDataVersion2 | Error |  |  |  | Wrong IPCalData version. |  | COM_Image2DCalibrationErr
22603 | CIM_IPCalDataType | Error |  |  |  | Wrong IPCalData type. |  | COM_Image2DCalibrationErr
22604 | CIM_IPCalDataTrailerLen | Error |  |  |  | IPCalFile trailer longer than data. |  | COM_Image2DCalibrationErr
22605 | CIM_IPCalFileTooSmall | Error |  |  |  | IPCalFile too small. |  | COM_Image2DCalibrationErr
22606 | CIM_IPCalFileChecksum | Error |  |  |  | IPCalFile checksum error. |  | COM_Image2DCalibrationErr
22607 | CIM_WriteThumbnailInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22608 | CIM_WriteThumbnailNoData | Error |  |  |  | CalibrationImage object contains no image data. |  | COM_Image2DCalibrationErr
22609 | CIM_WriteThumbnailRescale | Error |  |  |  | Rescaling the CalibrationImage failed. |  | COM_Image2DCalibrationErr
22610 | CIM_WriteThumbnailConvert | Error |  |  |  | Converting the CalibrationImage to 8 bits failed. |  | COM_Image2DCalibrationErr
22611 | CIM_AddNewAllocImage | Error |  |  |  | Could not allocate new image, when adding a new (empty) IP calibration image. |  | COM_Image2DCalibrationErr
22612 | CIM_AddNewIPCalFile | Error |  |  |  | Adding a new (empty) IP calibration image failed. |  | COM_Image2DCalibrationErr
22613 | CIM_AddNewInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22614 | CIM_CalibrateInvalid | Error |  |  |  | Invalid CalibrationImage object. |  | COM_Image2DCalibrationErr
22615 | CIM_CalibrateNoData | Error |  |  |  | CalibrationImage object contains no image data. |  | COM_Image2DCalibrationErr
22616 | CIM_NormalizeInvalid | Error |  |  |  | Invalid image for normalization. |  | COM_Image2DCalibrationErr
22617 | CIM_Normalize0 | Error |  |  |  | Could not normalize image. All pixels are 0. |  | COM_Image2DCalibrationErr
22618 | CIM_NoBinaryData | Error |  |  |  | No data given to create a CalibrationImage from. |  | COM_Image2DCalibrationErr
22619 | CIM_NoSingleBlock | Error |  |  |  | ImageArea does not consist of one single block. | Meaning: The ImageArea for line artefact correction does not consist of one single block. No correction done. Reason: Software error. Cure: No cure. The image is ok, besides the fact that a possible line artefact is not corrected. | COM_Image2DCalibrationErr
22628 | CLN_CreateType | Error |  |  |  | Invalid CalibrationLine file type. |  | COM_ImageCalibrationLineErr
22629 | CLN_CreatePixelWidth | Error |  |  |  | Invalid pixelWidth in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22630 | CLN_CreateNLines | Error |  |  |  | Invalid number of calibration lines in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22631 | CLN_CreateNColumns | Error |  |  |  | Invalid number of columns in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22632 | CLN_ROILeft | Error |  |  |  | Invalid ROILeft in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22633 | CLN_CreateBitsPP | Error |  |  |  | Unsupported number of bits per pixel in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22634 | CLN_CreateOpenFile | Error |  |  |  | Opening the CalibrationLine file failed. |  | COM_ImageCalibrationLineErr
22635 | CLN_CreateOpenCRFile | Error |  |  |  | Opening the CalibrationLine file in CRFileSystem failed. |  | COM_ImageCalibrationLineErr
22636 | CLN_DeserializeInvalid | Error |  |  |  | Invalid CalibrationLine object. |  | COM_ImageCalibrationLineErr
22637 | CLN_DeserializeLength | Error |  |  |  | CalibrationLine file too small. |  | COM_ImageCalibrationLineErr
22638 | CLN_DeserializeFileType | Error |  |  |  | Invalid CalibrationLine file type. |  | COM_ImageCalibrationLineErr
22639 | CLN_DeserializeHeaderLength | Error |  |  |  | Invalid CalibrationLine file header length. |  | COM_ImageCalibrationLineErr
22640 | CLN_DeserializeHeaderLength2 | Error |  |  |  | Invalid CalibrationLine file header length. |  | COM_ImageCalibrationLineErr
22641 | CLN_DeserializeHeaderLength1 | Error |  |  |  | Invalid CalibrationLine file header length. |  | COM_ImageCalibrationLineErr
22642 | CLN_DeserializeVersion | Error |  |  |  | Invalid version of CalibrationLine file. |  | COM_ImageCalibrationLineErr
22643 | CLN_DeserializeLength2 | Error |  |  |  | CalibrationLine file too small. |  | COM_ImageCalibrationLineErr
22644 | CLN_DeserializePixelWidth1 | Error |  |  |  | Invalid pixelWidth in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22645 | CLN_DeserializePixelWidth | Error |  |  |  | Invalid pixelWidth in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22646 | CLN_DeserializeNColumns | Error |  |  |  | Invalid number of columns in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22647 | CLN_DeserializeBitsPP | Error |  |  |  | Unsupported number of bits per pixel in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22648 | CLN_DeserializeLength3 | Error |  |  |  | CalibrationLine file too small. |  | COM_ImageCalibrationLineErr
22649 | CLN_DeserializeNLines | Error |  |  |  | Invalid number of calibration lines in CalibrationLine file. |  | COM_ImageCalibrationLineErr
22650 | CLN_DeserializeROILeft | Error |  |  |  | Invalid ROILeft in calibration file. |  | COM_ImageCalibrationLineErr
22651 | CLN_DeserializeLength4 | Error |  |  |  | Too few data to deserialize calibration line. |  | COM_ImageCalibrationLineErr
22652 | CLN_SerializeInvalid | Error |  |  |  | CalibrationLine_Serialize: Empty calibration line. |  | COM_ImageCalibrationLineErr
22653 | CLN_SerializeHeaderBytes | Error |  |  |  | CalibrationLine_Serialize: Wrong HeaderBytes (software bug). |  | COM_ImageCalibrationLineErr
22654 | CLN_NormalizeInvalid | Error |  |  |  | CalibrationLine_Normalize: Empty calibration line. |  | COM_ImageCalibrationLineErr
22655 | CLN_ClipInvalid | Error |  |  |  | Clip: Empty calibration line. |  | COM_ImageCalibrationLineErr
22656 | CLN_MaximumSlopeInvalid | Error |  |  |  | CalibrationLine_MaximumSlope: Empty calibration line. |  | COM_ImageCalibrationLineErr
22657 | CLN_MaximumSlopeNegative | Error |  |  |  | CalibrationLine_MaximumSlope: Unexpected negative value of maxDifference. |  | COM_ImageCalibrationLineErr
22658 | CLN_InterpolateInvalid | Error |  |  |  | CalibrationLine_Interpolate: Empty calibration line. |  | COM_ImageCalibrationLineErr
22659 | CLN_InterpolatePixelWidth | Error |  |  |  | CalibrationLine_Interpolate: Invalid new pixel width. |  | COM_ImageCalibrationLineErr
22660 | CLN_RescaleNNMatch | Error |  |  |  | CalibrationLine_RescaleNearestNeighbor: Length of input line and number of sections do not match. |  | COM_ImageCalibrationLineErr
22661 | CLN_RescaleCubicMatch | Error |  |  |  | CalibrationLine_RescaleCubic: Length of input line and number of sections do not match. |  | COM_ImageCalibrationLineErr
22662 | CLN_ExtrapolateInvalid | Error |  |  |  | CalibrationLine_Extrapolate: Empty calibration line. |  | COM_ImageCalibrationLineErr
22663 | CLN_ExtrapolateCut | Error |  |  |  | CalibrationLine_Extrapolate: Cannot cut more left columns than available. |  | COM_ImageCalibrationLineErr
22664 | CLN_ExtrapolateExtend | Error |  |  |  | CalibrationLine_Extrapolate: Cannot extend beyond begin of line. |  | COM_ImageCalibrationLineErr
22665 | CLN_ExtrapolateColumns | Error |  |  |  | CalibrationLine_Extrapolate: Invalid number of new columns. |  | COM_ImageCalibrationLineErr
22666 | CLN_MultiplyInvalid | Error |  |  |  | CalibrationLine_Multiply: Empty calibration line. |  | COM_ImageCalibrationLineErr
22667 | CLN_MultiplyLinesInvalid | Error |  |  |  | CalibrationLine_Multiply: Empty calibration line. |  | COM_ImageCalibrationLineErr
22668 | CLN_MultiplyLinesInvalid2 | Error |  |  |  | CalibrationLine_Multiply: Empty calibration line. |  | COM_ImageCalibrationLineErr
22669 | CLN_MultiplyLinesBitsPP | Error |  |  |  | CalibrationLine_Multiply: PixelBits do not match. |  | COM_ImageCalibrationLineErr
22670 | CLN_MultiplyLinesNCalLines | Error |  |  |  | CalibrationLine_Multiply: NumberOfCalLines do not match. |  | COM_ImageCalibrationLineErr
22671 | CLN_DivideInvalid | Error |  |  |  | CalibrationLine_Divide: Empty calibration line. |  | COM_ImageCalibrationLineErr
22672 | CLN_DivideZero | Error |  |  |  | CalibrationLine_Divide: Division by zero. |  | COM_ImageCalibrationLineErr
22673 | CLN_SubtractInvalid | Error |  |  |  | CalibrationLine_Subtract: Empty minuend line. |  | COM_ImageCalibrationLineErr
22674 | CLN_SubtractInvalid2 | Error |  |  |  | CalibrationLine_Subtract: Empty subtrahend line. |  | COM_ImageCalibrationLineErr
22675 | CLN_ExLineInvalid | Error |  |  |  | CalibrationLine_ExtrapolateLine: Line to be extrapolated is NULL! |  | COM_ImageCalibrationLineErr
22676 | CLN_ExLineStart | Error |  |  |  | CalibrationLine_ExtrapolateLine: startIndex not valid for line. |  | COM_ImageCalibrationLineErr
22677 | CLN_ExLineLength | Error |  |  |  | CalibrationLine_ExtrapolateLine: lineLength not valid for line. |  | COM_ImageCalibrationLineErr
22678 | CLN_ExLineFirstvalid | Error |  |  |  | CalibrationLine_ExtrapolateLine: firstValid is out of range. |  | COM_ImageCalibrationLineErr
22679 | CLN_ExLineLastValid | Error |  |  |  | CalibrationLine_ExtrapolateLine: lastValid is out of range. |  | COM_ImageCalibrationLineErr
22680 | CLN_MinMaxScanValInvalidLine | Error |  |  |  | CalibrationLine_CalculateMinMaxScanVal: Empty calibration line. |  | COM_ImageCalibrationLineErr
22681 | CLN_SplineDivByZero | Error |  |  |  | Bad input to routine splint leads to division by 0. |  | COM_ImageCalibrationLineErr
22682 | CLN_SetCalLinesLength | Error |  |  |  | Could not set CalLines due to wrong length. |  | COM_ImageCalibrationLineErr
22683 | CLN_SetCalLinesMalloc | Error |  |  |  | CalibrationLine_set_CalLines failed. malloc error. |  | COM_ImageCalibrationLineErr
22684 | CLN_ConsistencyInvalid | Error |  |  |  | CalibrationLine_CheckConsistency: Empty calibration line. |  | COM_ImageCalibrationLineErr
22685 | CLN_ConsistencyNotNormalized | Error |  |  |  | CalibrationLine_CheckConsistency: Gain line is not normalized. |  | COM_ImageCalibrationLineErr
22686 | CLN_ConsistencyRange | Error |  |  |  | CalibrationLine_CheckConsistency: Gain line contains too low pixel values. |  | COM_ImageCalibrationLineErr
22708 | MED_USNoData | Error |  |  |  | SlidingMedian_UShort: No data given to calculate median. |  | COM_SWError
22709 | MED_USRangeTooSmall | Error |  |  |  | SlidingMedian_UShort: The median range is too small. |  | COM_SWError
22710 | MED_USRangeTooBig | Error |  |  |  | SlidingMedian_UShort: The median range is bigger than the data length. |  | COM_SWError
22711 | MED_FNoData | Error |  |  |  | SlidingMedian_Float: No data given to calculate median. |  | COM_SWError
22712 | MED_FRangeTooSmall | Error |  |  |  | SlidingMedian_Float: The median range is too small. |  | COM_SWError
22713 | MED_FRangeTooBig | Error |  |  |  | SlidingMedian_Float: The median range is bigger than the data length. |  | COM_SWError
22714 | ART_DetNoData | Error |  |  |  | LineArtefact_Detect: No image data given. |  | COM_SWError
22715 | ART_DetAreaSize | Error |  |  |  | LineArtefact_Detect: Columns or line is less or equal 0 |  | COM_SWError
22751 | SYN_CRFSNoData | Error |  |  |  | SyncCalLines: Calibration file on CRFileSystem contains no data. |  | COM_SWError
22752 | SYN_Malloc | Error |  |  |  | SyncCalLines: Could not allocate buffer for CRFile. |  | COM_SWError
22753 | SYN_FOpen | Error |  |  |  | SyncCalLines: Could not open calibration file for writing. |  | COM_SWError
22754 | SYN_Dir | Error |  |  |  | SyncCalLines: Could not get directory. |  | COM_SWError
22755 | SYN_CreateCL | Error |  |  |  | SyncCalLines: Could not create CalibrationLine from file. |  | COM_SWError
22756 | SYN_CRFSWrite | Error |  |  |  | SyncCalLines: Could not write CalibrationLine to CRFileSystem. |  | COM_SWError
22757 | SYN_ExchangedAcq | Error | 159 | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. |  | SyncCalLines: Exchanged PMM detected. | Meaning: Detected an exchanged photomultiplier module. Shading calibration not good enough any more. Reason: Replacement by spare part. Cure: Perform a shading calibration. | COM_SWShadingCalibration
22758 | SYN_ExchangedStim | Error | 159 | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. |  | SyncCalLines: Exchanged optics module detected. | Meaning: Detected an exchanged optics module. Shading calibration not good enough any more. Reason: Replacement by spare part. Cure: Perform a shading calibration. | COM_SWShadingCalibration
22759 | SYN_AcqSerNo | Error | 159 | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. |  | SyncCalLines: Wrong PMM serial number. | Meaning: Detected wrong photomultiplier module serial number in calibration line. Shading calibration not valid. Reason: PMM serial number was changed or wrong calibration line was manually copied. Cure: Perform a shading calibration. | COM_SWShadingCalibration
22760 | SYN_GetLastWriteTime | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | SyncCalLines: Could not get the CalibrationLine write time (disk). |  | COM_SWError
22761 | SYN_CRFSSetLastWriteTime | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | SyncCalLines: Could not update the CalibrationLine write time in CRFileSystem. |  | COM_SWError
22762 | SYN_FWrite | Error | 232 | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | SyncCalLines: Could not write CalibrationLine to disk. |  | COM_SWError
23040 | SetError | Error |  |  |  | Setting an error failed.
23041 | RemoveError | Error |  |  |  | Removing an error failed.
23042 | StartDiagnoseCycle | Info | 283 | Start diagnose cycle. | Confirm | Start Diagnose Cycle
23043 | StartEnduranceRunCycle | Info | 284 | Start endurance test cycle. | Confirm | Start Endurancerun Cycles
23044 | StartShadingCalibration | Info | 285 | Start shading calibration. | Confirm | Start shading clibration
23045 | MoveIPCarrierToMaintenance | Info | 296 | Device ready to move IP carrier to maintenance position. | Confirm | Device ready to move IP Carrier to maintenance position
23046 | MoveIPCarrierToPark | Info | 290 | Device ready to move IP carrier to park position | Confirm | Device ready to move IP Carrier to park position
23049 | StartStallCalibration | Info | 297 | Device ready to start stall calibration. | Confirm | Start Stall Calibration
23050 | NOT_IMPLEMENTED | Error |  |  |  | This function has no implementation
23552 | START_SERVER_FAILED | Error |  |  |  | Start Server failed | Starting of DirectID Server failed | COM_SWError
26112 | SW_VERSION_MISMATCH | Error |  |  |  | The software of the digitizer and the ErrorViewer on PC do not match. | Meaning: The software of the digitizer and the ErrorViewer on NX workstation do not match. The ErrorViewer is not able to display messages of the digitizer. Reason: 1. An old version of ErrorViewer is installed on NX workstation. Cure: Please install the ErrorViewer on NX wordkstation released for actual version of digitizer software. | COM_SWErrorViewerVersionMismatch
26113 | CREATE_SENDER_SOCKET | Error |  |  |  | Could not create sender socket. | Meaning: Internal software error connecting to NX workstation. Reason: TCP/IP communication problem with NX workstation Cure: Please restart the digitizer. If the problem occurs after restart again, please call service. | COM_NotNeeded_ForServiceId
26114 | CONNECT_FAILED | Error |  |  |  | Could not connect to NX workstation. | Meaning: Connection of ErrorViewer to NX failed Reason: 1. The NX workstation is not available in the network. 2. The NX workstation is not started or user is not logged in at NX workstation 3. The ErrorViewer on NX workstation is not started. 4. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 5. ErrorViewer version on NX workstation is not supported 6. The digitizer is configured for another NX workstation than the one actually used. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. 3. Start the ErrorViewer on NX workstation. 4. Exit the ErrorViewer and restart it. 5. Install an ErrorViewer with version 2.0 or higher and start the ErrorViewer on NX workstation. 5. Reconfigure NX workstation and digitizer. | COM_SWErrorViewerCommunicationProblem
26115 | TIMEOUT_SEND | Error |  |  |  | No response from NX workstation | Meaning:NX workstation does not respond. Reason: 1. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 2. ErrorViewer version on NX workstation is not supported Cure: 1. Exit the ErrorViewer and restart it. 2. Install an ErrorViewer with version 2.0 or higher | COM_SWErrorViewerCommunicationProblem
26116 | FATAL_ERROR | Error |  |  |  | Fatal error building message string. |  | COM_SWError
26117 | ACK_NOT_SYNC | Error |  |  |  | Acknoledge asynchron | Meaning: Communication problem between between digitizer and Erroviewer on NX wordkstation Reason: Communication between digitizer and Erroviewer on NX wordkstation is not synchron. Cure: Restart digitizer and ErrorViewer on NX workstation. If the problem occures again reinstall SW. | COM_SWErrorViewerCommunicationProblem
26118 | ERROR_SENDING_MESSAGE | Error |  |  |  | Error sending a message to NX workstation | Meaning: An error message could not be sent to ErrorViewer on NX wrokstation. Reason: 1. The NX workstation has actually been removed from the network. 2. The NX workstation has actually been shut dwon or user has logged out from NX workstation 3. The ErrorViewer has actually been shut down on NX workstation. 4. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 6. The digitizer configuration has actually been modified over WEB insterface , the digitizer now is configured for another NX workstation than the one actually used. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. 3. Start the ErrorViewer on NX workstation. 4. Exit the ErrorViewer and restart it. 5. Install an ErrorViewer with version 2.0 or higher and start the ErrorViewer on NX workstation. 5. Check configuration of NX workstation and digitizer. | COM_SWErrorViewerCommunicationProblem
26119 | PING_FAILED | Error |  |  |  | Could not connect to NX workstation. | Meaning: Ping of NX workstation failed Reason: 1. The NX workstation is not available in the network. 2. The NX workstation is not started. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. | COM_SWErrorViewerCommunicationProblem