## CR30FirstErrors
ErrorCode | Identifier | Priority | TextToDisplay | ConfirmRequired | Description | ExtendedDescription
512 | NoErrorDefinitionsFound | Error |  |  | During the initialization of the error manager the files containing the error definitions have not been found. | Meaning: During the initialization of the error manager the files containing the error definitions (digitizerGlobal.xml / SMEL files) have not been found. Reason: Incomplete installation of PC software or certain files have been deleted. Cure: Install PC software again. If error remains, check for correct software version.
513 | ReadingConfigFile | Error |  |  | Error reading application's configuration file.
514 | InitializationEventDispatcher | Error |  |  | Initialization of EventDispatcher failed.
515 | EventLoggingInitialization | Error |  |  | Initialization of EventLogging failed.
516 | TakeStandardBmp | Error |  |  | Access to standard bitmap failed.
517 | RemoveError | Error |  |  | Removing an error failed.
518 | Reset | Error |  |  | Resetting the errortable failed.
519 | CreateTicket | Error |  |  | Creation of an ticket failed.
520 | RemoveTicket | Error |  |  | Removing a ticket failed.
521 | WriteToSysLog | Error |  |  | Writing to system log failed.
522 | AcknowledgeTicket | Error |  |  | Acknowledging of an ticket failed.
523 | GenerationDigitizerGlobal | Error |  |  | Generation of DigitizerGlobal failed.
524 | CannotCopyLogfile | Error |  |  | Cannot copy the logfile in case of a fatal error.
1024 | DTC_PERFTX_RDSYNC | Error |  |  | Error while waiting for sync frame | Timeout or error occured while initiating the transmission of identification data from the PMI Board to the TagReader component at the IP Handling Board. Either the CAN bus is defective or at least the TagReader component at the IP Handling Board (or the complete IP Handling Board) does not respond anymore.
1025 | DTC_PERFTX_BKLWRITE | Error |  |  | Error during block transfer of ipdata | Either the CAN bus is defective or at least the TagReader component at the IP Handling Board (or the complete IP Handling Board) does not respond anymore.
1026 | DTC_INITX_FCHOPEN | Error |  |  | Cannot open CAN channel for data transmission | Unable to open a channel to transmit the identification data from the PMI Board to the TagReader component at the IP Handling Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1027 | DTC_PERFRX_FCHOPEN | Error |  |  | Cannot open CAN channel for data retrieval | Unable to open a channel to transmit the identification data from the IP Handling Board to the PMI Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1028 | DTC_PERFRX_INITBLKRX | Error |  |  | Cannot initialize CAN channel for block retrieval | Timeout or error occured while initializing a block transfer over the CAN bus from the IP Handling Board to the PMI Board. This error is related to the PMI Board itself or to the software running on this PCB.
1029 | DTC_PERFRX_WRSYNC | Error |  |  | Error while writing sync frame | An error occured when trying to start the transmission of identification data from the IP Handling Board to the PMI Board over the CAN bus. This error is related to the PMI Board itself or to the software running on this PCB.
1030 | DTC_PERFRX_BLKRX | Error |  |  | Error while performing the block read | An error occured when reading the identifiaction data that are sent from the IP Handling Board to the PMI Board. This error is related to one or more of the following reasons: - The CAN bus is defective. - The IP Handling Board aborted the transmission (e.g. software crash, hardware fault, ...) - The PMI Board itself reported an error condition in the CAN logic.
1031 | VHLDLL_INITUART | Error |  |  | DLL: Initialization of the UART failed | Initializing the serial line that connects the IP Handling Board with the Chip Reader failed. This is an internal error of the IP Handling Board and a string indication that this PCB is defective and should be replaced.
1032 | VHLDLL_RCF_RCT1 | Error |  |  | DLL: Timeout while reading a character from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1033 | VHLDLL_RCF_RCT2 | Error |  |  | DLL: Timeout while reading a character from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1034 | VHLDLL_SCF_RCT | Error |  |  | DLL: Timeout while waiting for Sync ACK from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1035 | VHLDLL_RA_RCT | Error |  |  | DLL: Timeout while waiting for Sync ACK from the TagReader component | The TagReader component at the IP Handling Board waits for one or more characters from the Chip Reader and encountered a timeout condition. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1036 | VHLDLL_RX_UNKNOWN | Error |  |  | DLL: Unknown event in method receive() | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1037 | VHLDLL_TX_UNKNOWN | Error |  |  | DLL: Unknown event in method send() | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1038 | VHLDLL_STT_RXINITIAL_UNKNOWN | Error |  |  | DLL: Unknown event in initial transition in state OK::receiving | An unexpected character has been received from the Chip Reader. One or more of the following is likely to be the reason: - The cable between the IP Handling Board and the Chip Reader is defective or not connected properly. - The Chip Reader hung up. - The IP Handling Board is defective.
1039 | VHLAL_POWERON_TX | Error |  |  | AL: Error sending command powerON() | Sending the command 'antenna on' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1040 | VHLAL_POWERON_RX | Error |  |  | AL: Error receiving answer for powerON() | Receiving the answer to the command 'antenna on' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1041 | VHLAL_POWERON_CHECKSUM | Error |  |  | AL: Checksum error in answer to powerON() | The answer to the command 'antenna on' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1042 | VHLAL_POWERON_SEQID | Error |  |  | AL: Sequence error in answer for powerON() | The answer to the command 'antenna on' from the Chip Reader carries a wrong sequence ID.
1043 | VHLAL_SELECT_TX | Error |  |  | AL: Error sending command SELECT() | Sending the command 'select tag' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1044 | VHLAL_SELECT_RX | Error |  |  | AL: Error receiving answer for SELECT() | Receiving the answer to the command 'select tag' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1045 | VHLAL_SELECT_CHECKSUM | Error |  |  | AL: Checksum error in answer to SELECT() | The answer to the command 'select tag' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1046 | VHLAL_SELECT_SEQID | Error |  |  | AL: Sequence error in answer for SELECT() | The answer to the command 'select tag' from the Chip Reader carries a wrong sequence ID.
1047 | VHLAL_GETINFO_TX | Error |  |  | AL: Error sending command GETINFO() | Sending the command 'get info' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1048 | VHLAL_GETINFO_RX | Error |  |  | AL: Error receiving answer for GETINFO() | Receiving the answer to the command 'get info' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1049 | VHLAL_GETINFO_CHECKSUM | Error |  |  | AL: Checksum error in answer to GETINFO() | The answer to the command 'get info' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1050 | VHLAL_GETINFO_SEQID | Error |  |  | AL: Sequence error in answer for GETINFO() | The answer to the command 'get info' from the Chip Reader carries a wrong sequence ID.
1051 | VHLAL_READ_TX | Error |  |  | AL: Error sending command READ() | Sending the command 'read' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1052 | VHLAL_READ_RX | Error |  |  | AL: Error receiving answer for READ() | Receiving the answer to the command 'read' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1053 | VHLAL_READ_CHECKSUM | Error |  |  | AL: Checksum error in answer to READ() | The answer to the command 'read' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1054 | VHLAL_READ_SEQID | Error |  |  | AL: Sequence error in answer for READ() | The answer to the command 'read' from the Chip Reader carries a wrong sequence ID.
1055 | VHLAL_WRITE_TX | Error |  |  | AL: Error sending command WRITE() | Sending the command 'write' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1056 | VHLAL_WRITE_RX | Error |  |  | AL: Error receiving answer for WRITE() | Receiving the answer to the command 'write' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1057 | VHLAL_WRITE_CHECKSUM | Error |  |  | AL: Checksum error in answer to WRITE() | The answer to the command 'write' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1058 | VHLAL_WRITE_SEQID | Error |  |  | AL: Sequence error in answer for WRITE() | The answer to the command 'write' from the Chip Reader carries a wrong sequence ID.
1059 | VHLAL_POWEROFF_TX | Error |  |  | AL: Error sending command powerOFF() | Sending the command 'antenna off' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1060 | VHLAL_POWEROFF_RX | Error |  |  | AL: Error receiving answer for powerOFF() | Receiving the answer to the command 'antenna off' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1061 | VHLAL_POWEROFF_CHECKSUM | Error |  |  | AL: Checksum error in answer to powerOFF() | The answer to the command 'antenna off' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1062 | VHLAL_POWEROFF_SEQID | Error |  |  | AL: Sequence error in answer for powerOFF() | The answer to the command 'antenna off' from the Chip Reader carries a wrong sequence ID.
1063 | VHLAL_SETMIF_TX | Error |  |  | AL: Error sending command setupMIFARE() | Sending the command 'select Mifare access mode' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1064 | VHLAL_SETMIF_RX | Error |  |  | AL: Error receiving answer for setupMIFARE() | Receiving the answer to the command 'select Mifare access mode' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1065 | VHLAL_SETMIF_CHECKSUM | Error |  |  | AL: Checksum error in answer to setupMIFARE() | The answer to the command 'select Mifare access mode' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1066 | VHLAL_SETMIF_SEQID | Error |  |  | AL: Sequence error in answer for setupMIFARE() | The answer to the command 'select Mifare access mode' from the Chip Reader carries a wrong sequence ID.
1067 | VHLAL_GETSNR_TX | Error |  |  | AL: Error sending command getSNR() | Sending the command 'getSerialNumberOfTag' to the Chip Reader failed. Either the Chip Reader is defective or the cable between the IP Handling Board and the Chip Reader is broken or not connected properly.
1068 | VHLAL_GETSNR_RX | Error |  |  | AL: Error receiving answer for getSNR() | Receiving the answer to the command 'getSerialNumberOfTag' from the Chip Reader failed. Either a timeout or some other sort of protocol violation. A broken cable is unlikely, as the command itself had been sent successfully and had been acknowledged.
1069 | VHLAL_GETSNR_CHECKSUM | Error |  |  | AL: Checksum error in answer to getSNR() | The answer to the command 'getSerialNumberOfTag' from the Chip Reader is corrupted (wrong checksum). Please check the cable between the Chip Reader and the IP Handling Board.
1070 | VHLAL_GETSNR_SEQID | Error |  |  | AL: Sequence error in answer for getSNR() | The answer to the command 'getSerialNumberOfTag' from the Chip Reader carries a wrong sequence ID.
1071 | APP_TEST_ALINIT | Error |  |  | APP: Initializing the application layer failed | Cannot initialize the communication layer between the IP Handling Board and the Chip Reader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1072 | APP_TEST_GETINFO | Error |  |  | APP: Cannot retrieve firmware information | Cannot execute command 'getFirmwareVersion' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1073 | APP_GETINFO | Error |  |  | APP: Cannot retrieve firmware information | Cannot execute command 'getFirmwareVersion' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1074 | APP_TESTTAG_PWRON | Error |  |  | APP: Cannot switch on RF-power in testTAG() | Cannot execute command 'testForTag/AntennaOn' on the TagReader. Possible reasons are: - The cable between the IP Handling Board and the Chip Reader is broken. - The Chip Reader might be defective. - The IP Handling Board might be defective.
1086 | APP_WR_GETPAR | Error |  |  | APP: Cannot retrieve signal parameters in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to retrieve the parameters from the message. Internal software(?) error at the IP Handling Board.
1087 | APP_WR_BLKRD | Error |  |  | APP: Cannot perform block read in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to perform the data transfer from the PMI Board. Error on CAN bus or PMI Board or IP Handling Board.
1088 | APP_WR_SYNCFRAME | Error |  |  | APP: Cannot send SYNC frame in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to initiate the data transfer from the PMI Board. Error on CAN bus or PMI Board or IP Handling Board.
1089 | APP_WR_BRXINIT | Error |  |  | APP: Cannot initialize block transfer in method WRITE | Cannot execute command 'Write' on the TagReader. Unable to initialize the data transmission from the PMI Board. Error on IP Handling Board.
1090 | RD_FAILED | Error |  |  | command read failed. | Cure: Check cabling of the TagReader (broken?)
1536 | INFO_VERS_TERM | Error |  |  | IPInfo::Version: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::Version attribute read.
1537 | INFO_VERS_CORRUPTED | Error |  |  | IPInfo::Version: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Version.
1538 | INFO_IDNUM_TERM | Error |  |  | IPInfo::IDNumber: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IDNumber attribute read.
1539 | INFO_IDNUM_CORRUPTED | Error |  |  | IPInfo::IDNumber: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IDNumber.
1540 | INFO_INITDATE_TERM | Error |  |  | IPInfo::InitDate: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::InitDate attribute read.
1541 | INFO_INITDATE_CORRUPTED | Error |  |  | IPInfo::InitDate: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::InitDate.
1542 | INFO_IPSIZE_TERM | Error |  |  | IPInfo::IPSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSize attribute read.
1543 | INFO_IPSIZE_CORRUPTED | Error |  |  | IPInfo::IPSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSize.
1544 | INFO_SCANSIZE_TERM | Error |  |  | IPInfo::ScanSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::ScanSize attribute read.
1545 | INFO_SCANSIZE_CORRUPTED | Error |  |  | IPInfo::ScanSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::ScanSize.
1546 | INFO_IPTYPE_TERM | Error |  |  | IPInfo::IPType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPType attribute read.
1547 | INFO_IPTYPE_CORRUPTED | Error |  |  | IPInfo::IPType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPType.
1548 | INFO_IPSENS_TERM | Error |  |  | IPInfo::IPSensitivity: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSensitivity attribute read.
1549 | INFO_IPSENS_CORRUPTED | Error |  |  | IPInfo::IPSensitivity: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSensitivity.
1550 | INFO_IPERA_TERM | Error |  |  | IPInfo::IPErasibility: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPErasibility attribute read.
1551 | INFO_IPERA_CORRUPTED | Error |  |  | IPInfo::IPErasibility: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPErasibility.
1552 | INFO_IPTHICK_TERM | Error |  |  | IPInfo::IPThickness: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPThickness attribute read.
1553 | INFO_IPTHICK_CORRUPTED | Error |  |  | IPInfo::IPThickness: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPThickness.
1554 | INFO_CHECKSUM_TERM | Error |  |  | IPInfo::Checksum: unexpected attribute terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: No section terminator after IPInfo::Checksum attribute read.
1555 | INFO_CHECKSUM_CORRUPTED | Error |  |  | IPInfo::Checksum: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Checksum.
1556 | INFO_CHECKSUM_WRONG | Error |  |  | IPInfo::Checksum: checksum error | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Checksum error for section IPInfo.
1557 | STATUS_VERS_TERM | Error |  |  | IPStatus::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Version attribute read.
1558 | STATUS_VERS_NOTSUPPORTED | Error |  |  | IPStatus::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPStatus::Version not supported.
1559 | STATUS_VERS_CORRUPTED | Error |  |  | IPStatus::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Version.
1560 | STATUS_STATUS_TERM | Error |  |  | IPStatus::Status: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Status attribute read.
1561 | STATUS_STATUS_CORRUPTED | Error |  |  | IPStatus::Status: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Status.
1562 | STATUS_CYCLES_TERM | Error |  |  | IPStatus::Cycles: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Cycles attribute read.
1563 | STATUS_CYCLES_CORRUPTED | Error |  |  | IPStatus::Cycles: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Cycles.
1564 | STATUS_CHECKSUM_TERM | Error |  |  | IPStatus::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPStatus::Checksum attribute read.
1565 | STATUS_CHECKSUM_CORRUPTED | Error |  |  | IPStatus::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Checksum.
1566 | STATUS_CHECKSUM_WRONG | Error |  |  | IPStatus::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPStatus.
1567 | CTRL_VERS_TERM | Error |  |  | IPCtrl::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Version attribute read.
1568 | CTRL_VERS_NOTSUPPORTED | Error |  |  | IPCtrl::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPCtrl::Version not supported.
1569 | CTRL_VERS_CORRUPTED | Error |  |  | IPCtrl::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Version.
1570 | CTRL_SENSITIVITY_TERM | Error |  |  | IPCtrl::Sensitivity: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Sensitivity attribute read.
1584 | IPPSDATA_SOPINSTUID_TERM | Error |  |  | IPPSData::SOPInstanceUID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::SOPInstanceUID attribute read.
1585 | IPPSDATA_SOPINSTUID_CORRUPTED | Error |  |  | IPPSData::SOPInstanceUID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::SOPInstanceUID.
1586 | IPPSDATA_PROCSTAT_TERM | Error |  |  | IPPSData::ProcessingStation: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ProcessingStation attribute read.
1587 | IPPSDATA_PROCSTAT_CORRUPTED | Error |  |  | IPPSData::ProcessingStation: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ProcessingStation.
1588 | IPPSDATA_STIMIDENTID_TERM | Error |  |  | IPPSData::StimIdentID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::StimIdentID attribute read.
1589 | IPPSDATA_STIMIDENTID_CORRUPTED | Error |  |  | IPPSData::StimIdentID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::StimIdentID.
1590 | IPPSDATA_EXPROOMID_TERM | Error |  |  | IPPSData::ExposureRoomID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ExposureRoom attribute read.
1591 | IPPSDATA_EXPROOMID_CORRUPTED | Error |  |  | IPPSData::ExposureRoomID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ExposureRoom.
1592 | IPPSDATA_CHECKSUM_TERM | Error |  |  | IPPSData::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPPSData::Checksum attribute read.
1593 | IPPSDATA_CHECKSUM_CORRUPTED | Error |  |  | IPPSData::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::Checksum.
1594 | IPPSDATA_CHECKSUM_WRONG | Error |  |  | IPPSData::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPPSData.
1595 | INFO_VERSION_NOTSUPPORTED | Error |  |  | IPInfo::Version: this version is not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPInfo::Version not supported.
1596 | INFO_APPTYPE_TERM | Error |  |  | IPInfo::ApplicationType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::AppType attribute read.
1597 | INFO_APPTYPE_CORRUPTED | Error |  |  | IPInfo::ApplicationType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::AppType.
2048 | Cmn_GeneralError | Error |  |  | General error in ASAPCommon.
2049 | Sig_ToByteArrayConversionFailure | Error |  |  | The signal could not be converted to byte array.
2050 | Sig_FromByteArrayConversionFailure | Error |  |  | The signal could not be converted from byte array.
2051 | PTB_SendSignalCreate | Error |  |  | Creation of dignal to be sent failed.
2560 | GeneralError | Error |  |  | General error in ASAPServer.
2561 | SRV_AlreadyOpen | Error |  |  | Server has already been opened. Close first.
2562 | SRV_AlreadyClosed | Error |  |  | Server has already been closed or has not yet successfully been opened.
2563 | SRV_NewAgfaLogger | Error |  |  | Exception getting AgfaLogger.
2564 | SRV_ConfigAgfaLogger | Error |  |  | Exception configuring AgfaLogger.
2565 | SRV_Config | Error |  |  | Exception configuring the server.
2566 | SRV_NewErrorManager | Error |  |  | Exception creating and starting the ErrorManager.
2567 | SRV_CloseErrorManager | Error |  |  | Exception closing the ErrorManager.
2568 | SRV_ConnectDevice | Error |  |  | Exception connecting to specific device.
2569 | SRV_DisconnectDevice | Error |  |  | Exception disconnecting from the scanner device.
2570 | SRV_NoDevice | Error |  |  | No scanner device found.
2571 | SRV_NoBoardInterface | Error |  |  | Not connected to BoardInterface.
2572 | SRV_CreateBoardInterface | Error |  |  | Exception creating BoardInterface.
2573 | SRV_NoDataContainer | Error |  |  | Not connected to DataContainer.
2574 | SRV_CreateDataContainer | Error |  |  | Exception creating DataContainer.
2575 | SRV_NoImageData | Error |  |  | No DIFO object for ImageData.
2576 | SRV_CreateImageData | Error |  |  | Exception creating DIFO object for ImageData.
2577 | SRV_NewIddlHal | Error |  |  | Exception creating IDDL_HAL.
2578 | SRV_NewIddlHalAbort | Error |  |  | ThreadAbortException creating IDDL_HAL.
2579 | SRV_NewRemoting | Error |  |  | Exception creating and starting remoting.
2580 | SRV_CloseRemoting | Error |  |  | Exception closing remoting channel.
2581 | SRV_NewEsfDeviceServer | Error |  |  | Exception creating and starting ESFDeviceServer.
2582 | SRV_CloseEsfDeviceServer | Error |  |  | Exception closing ESFDeviceServer.
2583 | SRV_NewFileMapping | Error |  |  | Exception creating memory file for image transfer or initializing memory blocks.
2584 | SRV_CloseFileMapping | Error |  |  | Exception closing memory file for image transfer.
2585 | SRV_ProtocolWriteDeserialize | Error |  |  | Exception at deserialization of signal, before writing protocol data to device. | This is a software error and should not occur.
2586 | SRV_ProtocolWrite | Error |  |  | Exception writing protocol data to device. | Meaning: Communication problem between PC and digitizer. Reason: Cure: Upgrade software. Check Firewire cable.
2587 | SRV_ProtocolWriteCount | Error |  |  | Written less protocol bytes than requested to device. | Meaning: The request to send a signal via the Communicator protocol server was not completed successfully. Sent less bytes than requested. Reason: Transmission problems or software error. Cure:
2588 | SRV_GetIPData | Error |  |  | Exception at getting IPData. Using default.
2589 | SRV_SetIPData | Error |  |  | Exception at setting IPData.
2590 | SRV_GetReceivedCalLineData | Error |  |  | Exception at getting data from the DIFO object 'CalibrationLine'.
2591 | SRV_SetCalLineDataForTransmission | Error |  |  | Exception at setting data for the DIFO object 'CalibrationLine'.
2592 | SRV_GetIPCalibrationData | Error |  |  | Exception at getting IPCalibrationData.
2593 | SRV_SetIPCalibrationData | Error |  |  | Exception at setting IPCalibrationData.
2594 | SRV_OpenImageData | Error |  |  | Exception at opening image data.
2595 | SRV_ReadImageData | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Exception at reading image data from device. | Meaning: Communication problem between PC and digitizer. Reason: PC workload (especially during NX startup) too high to transmit image or digitizer PMI board problem. Cure: Check if image is transferred. If error persists, check Firewire cable, Firewire card in PC, PMI interface.
2596 | SRV_ReadImageDataCount | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Read less image block bytes from device than requested.
2597 | SRV_ReadImageDataNoServer | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Tried to read image data, but no server available. | Device probably off-line or there is any connection problem.
2598 | SRV_ReadImageDataWriteFile | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Exception at writing image data to file.
2599 | SRV_SendResponseToClient | Error |  |  | Exception at sending signal to clients.
2600 | SRV_SendResponseToLog | Error |  |  | Exception at sending signal to logging.
2601 | SRV_DataContainerRead | Error |  |  | Exception at treating outgoing data for DataContainer.
2602 | SRV_DataContainerWrite | Error |  |  | Exception at treating incoming data for DataContainer.
2603 | SRV_OpenSSoFManager | Error |  |  | Exception opening SSoFManager.
2604 | SRV_CloseSSoFManager | Error |  |  | Exception closing SSoFManager.
2605 | SRV_InitializeOutputFiles | Error |  |  | Exception at initializing output files during OpenImageData.
2606 | SRV_SetupCorrection | Error |  |  | Exception at SetupCorrection.
2607 | SRV_SendResponseToErrorManager | Error |  |  | Exception at sending signal to error manager.
2608 | SRV_ProtocolWriteOffLine | Error | The communication has been interrupted. The system tries to reconnect. Check whether the digitizer is powered on and the FireWire cable is connected. |  | Protocol data cannot be written to device because device is off-line.
2610 | SRV_NewEmsDifo | Error |  |  | Exception creating EMS_DIFO.
2611 | SRV_NewEmsDifoAbort | Error |  |  | ThreadAbortException creating EMS_DIFO.
2612 | SRV_NewEsfBridge | Error |  |  | Exception creating and starting ESFBridge.
2613 | SRV_CloseEsfBridge | Error |  |  | Exception closing ESFBridge.
2615 | SRV_DisconnectCommunicator | Error |  |  | Exception disconnecting Communicator servers.
2616 | SRV_InitIddlHal | Error |  |  | Error initializing IDDL_HAL.
2617 | SRV_InitEmsDifo | Error |  |  | Error initializing EMS_DIFO.
2620 | SRV_ClearCalibrationLineValid | Error |  |  | Exception at deleting the temporary local copy of the CalibrationLine.
2621 | SRV_SetCalibrationLine | Error | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. | Confirm | Setting CalibrationLine with data from scanner failed. Retry with data from PC hard disk may be successful. | Meaning: Calibration not available or not valid any more. Reason: Parts replacement. Cure: Perform shading calibration.
2622 | SRV_SetCalibrationLineFromFile | Error | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. | Confirm | Setting CalibrationLine with data from the PC hard disk failed. | Meaning: Calibration not available or not valid any more. Reason: Parts replacement. Cure: Perform shading calibration.
2623 | SRV_ReadImageDataChecksum | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Checksum error in read image data block. | Meaning: Communication problem between PC and digitizer. Reason: PC workload (especially during NX startup) too high to transmit image or digitizer PMI board problem. Cure: Check, if image is transferred; if error persists, check firewire cable, firewire card in PC, PMI interface.
2624 | SRV_LineOffset | Warning | Possible problem with image quality (stripes). |  | Invalid line offset (facet number of first image line). Using 0 instead. | One possible reason for a wrong facet number is a PMI FPGA bus error. In this case, the digitizer should be switched off and on.
2625 | SRV_LoadIPCalFromDisk | Error |  |  | Loading the IP calibration image failed. One reason may be, that is was not found on disk. It will then be loaded from the device and stored to disk.
2626 | SRV_LoadIPCalFromTag | Error |  |  | Loading the IP calibration image from the ID tag failed. Probably not present or corrupted..
2627 | SRV_ReadImageDataDifoException | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Exception from Communicator or Difo, while reading image data block. | Meaning: Exception while reading image data. Reason: There is a severe transmission problem for image data. The automatic retries did not help. Cure: First try to recover the image by restarting the PC. If that does not help, the image is lost and you have to switch off and on the digitizer.
2628 | SRV_GetDoseDetectionLevels | Error |  |  | Exception at getting data from the DIFO object 'DoseDetectionLevels'.
2629 | SRV_SetDoseDetectionLevels | Error |  |  | Exception at setting data for the DIFO object 'DoseDetectionLevels'.
2640 | SRV_NewDeviceErrorViewer | Error |  |  | Exception creating DeviceErrorViewer.
2641 | SRV_OpenDeviceErrorViewer | Error |  |  | Exception opening DeviceErrorViewer.
2642 | SRV_CloseDeviceErrorViewer | Error |  |  | Exception closing DeviceErrorViewer.
2660 | SRV_ExaSendSignalDenied | Error |  |  | Sender of signal has not exclusive access. | Meaning: Sender of signal on PC has not exclusive access. Reason: More than one client (e.g. TWAIN, ServiceClient) are active and conflicting. Cure: Stop NX when using ServiceClient.
2661 | SRV_ExaSendSignalToDevice | Error |  |  | Exception at checking exclusive access. | Meaning: Sending a signal (e.g. command) to the digitizer failed due to an exception when trying to get access. Reason: Software error. Cure: Report this error, if it occurs repeatedly.
2662 | SRV_ExaRegisterExclusiveAccess | Error |  |  | Exception during registration for exclusive access.
2663 | SRV_ExaUnRegisterExclusiveAccess | Error |  |  | Exception during unregistration of exclusive access.
2664 | SRV_ExaExclusiveAccessStopped | Error |  |  | Exception during polling the exclusive access.
2710 | ETP_SetError | Error |  |  | Setting an error failed.
2711 | ETP_RemoveError | Error |  |  | Removing an error failed.
2712 | SRV_VersionMismatch | Error | The software version of the digitizer and that one installed on PC do not match. Please contact service to finish installation. |  | ASAP Software versions loaded onto machine and installed on PC are not identical. | Meaning: ASAP software versions loaded onto machine and installed on PC are not identical. Reason: Incomplete software installation, upload failed or hardware replacement (PC or board digitizer) . Cure: Start Service Client and perform software upload to digitizer.
2713 | SRV_VersionCheckException | Error |  |  | Exception when checking ASAP Software versions.
2714 | SRV_SerialNumberNotRegistered | Error | The digitizer is not registered correctly. Serial number mismatch. Please contact service to fix the problem. |  | The device is not registered correctly: the serial number read from device and the one read from infocounter are not identical.
2715 | SRV_SyncDeviceDatabase | Info | Synchronizing digitizer database. |  | SSoF is busy with synchronizing the device database.
2716 | SRV_StopIddlEmsCommandWrapper | Error |  |  | Stopping IddlEmsCommandWrapper failed.
2717 | SRV_OpenIddlEmsCommandWrapper | Error |  |  | Opening IddlEmsCommandWrapper failed.
2718 | SRV_ChangedModuleNumberDetected | Error |  |  | At least one module/ board has been exchanged.
3072 | GeneralError | Error |  |  | General error in DigitizerAccessor.
3073 | PDS_Connect | Error |  |  | PcToDicSync.Connect failed.
3074 | PDS_GetIPData | Error |  |  | Synchronous call of GetIPData failed due to exception.
3075 | PDS_SetIPData | Error |  |  | Synchronous call of SetIPData failed due to exception.
3076 | PDS_OpenImageData | Error |  |  | Synchronous call of OpenImageData failed.
3077 | PDS_OpenImageDataEx | Error |  |  | Synchronous call of OpenImageData failed due to exception.
3078 | PDS_CloseImageData | Error |  |  | Synchronous call of CloseImageData failed.
3079 | PDS_CloseImageDataEx | Error |  |  | Synchronous call of CloseImageData failed due to exception.
3080 | PDS_ReadImageData | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Synchronous call of ReadImageData failed.
3081 | PDS_ReadImageDataEx | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Synchronous call of ReadImageData failed due to exception.
3082 | PDS_ReadImageDataIPtr | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Synchronous call of ReadImageData failed.
3083 | PDS_ReadImageDataIPtrEx | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | Synchronous call of ReadImageData failed due to exception.
3084 | PDS_ReadImageDataToFile | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | ReadImageData to file failed. Received not enough data.
3085 | PDS_ReadImageDataToFileEx | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | ReadImageData to file failed due to exception.
3086 | PDS_ReadIPCalData | Error |  |  | Synchronous call of ReadIPCalData failed due to exception.
3087 | PDS_StoreIPCalData | Error |  |  | Synchronous call of StoreIPCalData failed due to exception.
3088 | PDS_ReadIPCalDataNoData | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | ReadIPCalData returned no data.
3092 | PDS_FeedIPWrongState | Error |  |  | FeedIP was called at a wrong state.
3093 | PDS_FeedIPPending | Error |  |  | FeedIP response is already pending.
3094 | PDS_WaitFeedIPTimeout | Error |  |  | FeedIP timed out. | Meaning: Digitizer doesn't reply on PC request. Reason : - Cure: Power off and on digitizer.
3095 | PDS_WaitFeedIPOffLine | Error |  |  | Off-line during FeedIP.
3096 | PDS_WaitFeedIPWrongState | Error |  |  | Wrong state for waiting for end of FeedIP.
3097 | PDS_ScanWrongState | Error |  |  | Scan was called at a wrong state.
3098 | PDS_ScanPending | Error |  |  | Scan response is already pending.
3099 | PDS_WaitScanStartedTimeout | Error |  |  | Scan started timed out. | Meaning: IP did not reach 'begin of scan' detector. Reason: Wrong speedfactor, scan roller not fixed. Cure: Check mechanics of IP transport, clean IP, reboot NX and digitzer.
3100 | PDS_WaitScanStartedOffLine | Error |  |  | Off-line during preparing Scan. | Meaning: Off-line during preparing scan. Reason: Sporadic disconnect. Cure: Check Firewire cable.
3101 | PDS_WaitScanStartedWrongState | Error |  |  | Wrong state for waiting Scan started.
3104 | PDS_WaitScanTimeout | Error |  |  | Scan timed out.
3105 | PDS_WaitScanOffLine | Error |  |  | Off-line during Scan. | Meaning: Off-line during scan. Reason: Sporadic disonnect. Cure: Check Firewire cable.
3106 | PDS_WaitScanWrongState | Error |  |  | Wrong state for waiting for end of Scan.
3107 | PDS_EraseWrongState | Error |  |  | Erase was called at a wrong state.
3108 | PDS_ErasePending | Error |  |  | Erase response is already pending.
3109 | PDS_WaitEraseTimeout | Error |  |  | Erase timed out.
3110 | PDS_WaitEraseOffLine | Error |  |  | Off-line during Erase.
3111 | PDS_WaitEraseWrongState | Error |  |  | Wrong state for waiting for end of Erase.
3112 | PDS_ReturnIPWrongState | Error |  |  | ReturnIP was called at a wrong state.
3113 | PDS_ReturnIPPending | Error |  |  | ReturnIP response is already pending.
3114 | PDS_WaitReturnIPTimeout | Error |  |  | ReturnIP timed out.
3115 | PDS_WaitReturnIPOffLine | Error |  |  | Off-line during ReturnIP.
3116 | PDS_WaitReturnIPWrongState | Error |  |  | Wrong state for waiting for end of ReturnIP.
3117 | PDS_EjectWrongState | Error |  |  | Eject was called at a wrong state.
3118 | PDS_EjectPending | Error |  |  | Eject response is already pending.
3119 | PDS_WaitEjectTimeout | Error |  |  | Eject timed out.
3120 | PDS_WaitEjectOffLine | Error |  |  | Off-line during Eject.
3121 | PDS_WaitEjectWrongState | Error |  |  | Wrong state for waiting for end of Eject.
3122 | PDS_StartSelftestWrongState | Error |  |  | StartSelftest was called at a wrong state.
3123 | PDS_StartSelftestPending | Error |  |  | StartSelftest response is already pending.
3124 | PDS_WaitStartSelftestTimeout | Error |  |  | StartSelftest timed out.
3125 | PDS_WaitStartSelftestOffLine | Error |  |  | Off-line during StartSelftest.
3126 | PDS_WaitStartSelftestWrongState | Error |  |  | Wrong state for waiting for end of StartSelftest.
3127 | PDS_CleanUpWrongState | Error |  |  | CleanUp was called at a wrong state.
3132 | PDS_EnableFeederRspState | Error |  |  | Call of enable feeder returned a response state different from ok.
3137 | PDS_DisableFeederRspState | Error |  |  | Call of disable feeder returned a response state different from ok.
3138 | PDS_DirectIDTimeout | Error | ID button not pressed within the timeout period. Image plate was not scanned. Cassette stays clamped until ID data are supplied by the operator. |  | Timeout while waiting for identification. Scan was cancelled. Please try again. | Meaning: User did not confirm identification at NX in time. Reason: 5 minutes time span elapsed without identification. Cure: Instruct operator.
3139 | PDS_ScanNotStarted | Error | Scan failed. Image plate was not erased. Please confirm to eject the cassette. Try to scan the cassette again after restarting the digitizer. | Confirm | Error or timeout while waiting for end of scan. | Meaning: Scanning the image plate did not start. Reason: Error in digitizer or timeout. Cure: Reset digitizer.
3140 | PDS_DisconnectThreadNotTerminated | Error |  |  | StateMachineThread could not be terminated.
3142 | PDS_CmdBestFit | Error |  |  | A command returned with status best fit. | A command could not be executed exactly as required, but performed as good as possible. No error code was given. Refer to logging to see which command under which conditions had the problem.
3143 | PDS_CmdFailed | Error |  |  | A command returned with status failed. | Execution of a certain command failed. No error code was given. Refer to logging to see which command under which conditions had the problem.
3144 | PDS_CmdFatal | Error |  |  | A command returned with status fatal. | Execution of a certain command failed fatally. No error code was given. Refer to logging to see which command under which conditions had the problem.
3145 | PDS_CmdNotStarted | Error |  |  | A command returned with status not started. | A certain command could not be executed. No error code was given. Refer to logging to see which command under which conditions had the problem.
3146 | PDS_CmdTimeout | Error |  |  | A command returned with status timeout. | Execution of a certain command timed out. No error code was given. Refer to logging to see which command under which conditions had the problem.
3147 | PDS_PrepareImageCorrection | Error |  |  | Exception while preparing 2 dimensional image correction. | Execution of a certain command timed out. No error code was given. Refer to logging to see which command under which conditions had the problem.
3148 | PDS_IPCalNameMismatch | Error |  |  | The name of the IP calibration file and the contained serial number do not match. | Execution of a certain command timed out. No error code was given. Refer to logging to see which command under which conditions had the problem.
3149 | PDS_LoadIPCalFromDisk | Error |  |  | Exception loading the IP calibration file from disk. Probably not present. File will be created from ID tag data. | Execution of a certain command timed out. No error code was given. Refer to logging to see which command under which conditions had the problem.
3150 | PDS_LoadIPCalFromTag | Error |  |  | Exception loading the IP calibration file from the ID tag. Probably not present or corrupted. | Execution of a certain command timed out. No error code was given. Refer to logging to see which command under which conditions had the problem.
3151 | PDS_OldIPCal | Error |  |  | Old type IP calibration detected. Should never happen. No conversion foreseen. | Execution of a certain command timed out. No error code was given. Refer to logging to see which command under which conditions had the problem.
3172 | Rsp_ToByteArrayConversionFailure | Error |  |  | The response could not be converted to byte array.
3173 | Rsp_FromByteArrayConversionFailure | Error |  |  | The response could not be converted from byte array.
3174 | RspBool_ToByteArrayConversionFailure | Error |  |  | The boolean response could not be converted to byte array.
3175 | RspBool_FromByteArrayConversionFailure | Error |  |  | The boolean response could not be converted from byte array.
3176 | RspUInt16_ToByteArrayConversionFailure | Error |  |  | The UInt16 response could not be converted to byte array.
3177 | RspUInt16_FromByteArrayConversionFailure | Error |  |  | The UInt16 response could not be converted from byte array.
3180 | RspUInt32_ToByteArrayConversionFailure | Error |  |  | The UInt32 response could not be converted to byte array.
3181 | RspUInt32_FromByteArrayConversionFailure | Error |  |  | The UInt32 response could not be converted from byte array.
3182 | RspInt32_ToByteArrayConversionFailure | Error |  |  | The Int32 response could not be converted to byte array.
3183 | RspInt32_FromByteArrayConversionFailure | Error |  |  | The Int32 response could not be converted from byte array.
3184 | RspByteArr_ToByteArrayConversionFailure | Error |  |  | The Byte array response could not be converted to byte array.
3185 | RspByteArr_FromByteArrayConversionFailure | Error |  |  | The Byte array response could not be converted from byte array.
3186 | RspUInt16Arr_ToByteArrayConversionFailure | Error |  |  | The UInt16 array response could not be converted to byte array.
3187 | RspUInt16Arr_FromByteArrayConversionFailure | Error |  |  | The UInt16 array response could not be converted from byte array.
3188 | RspString_ToByteArrayConversionFailure | Error |  |  | The String response could not be converted to byte array.
3189 | RspString_FromByteArrayConversionFailure | Error |  |  | The String response could not be converted from byte array.
3192 | CmdBase_NotConnectedToServer | Error |  |  | Could not send signal to device, because we are not connected to the server. | The client software tried to send a command while there is no connection to the device server software. Refer to logging to see which command under which conditions had the problem.
3193 | CmdBase_BroadcastFailure | Error |  |  | Could not send signal to device. | The client software could not send a command to the device server software. Refer to logging to see which command under which conditions had the problem.
3194 | CmdBase_ResponseTimeout | Error | Error occurred. Restart the PC. If not successful, restart PC and digitizer. If error occurs again, please contact service. |  | The response did not arrive within the required time. | Meaning: Execution of a certain command timed out. Refer to logging to see which command under which conditions had the problem. Reason: Communication problem within PC. Cure: Restart PC.
3195 | CmdBase_OffLineNoResponse | Error | The communication has been interrupted. The system tries to reconnect. Check whether the digitizer is powered on and the FireWire cable is connected. |  | The device went off-line while waiting for a response. | Meaning: Execution of a certain command timed out. Refer to logging to see which command under which conditions had the problem. Reason: Digitizer went off-line. Cure: Check FireWire connection. Restart PC.
3222 | PIS_ReadCalibrationLine | Error |  |  | Synchronous call of ReadCalibrationLine failed due to exception.
3223 | PIS_WriteCalibrationLine | Error |  |  | Synchronous call of WriteCalibrationLine failed due to exception.
3224 | PIS_UpdateCalibrationLine | Error | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. | Confirm | UpdateCalibrationLine failed due to exception. | Meaning: Calibration not available. Reason: Exception when reading calibration data from digitizer. Cure: Restart PC.
3225 | PIS_UpdateCalExclAccess | Error | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. | Confirm | UpdateCalibrationLine failed: Could not get exclusive access. | Meaning: Calibration not available or not valid any more. Reason: Parts replacement. Cure: Perform shading calibration.
3226 | PIS_UpdateCalReadFromDevice | Error | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. | Confirm | UpdateCalibrationLine failed: Could not read data from scanner. | Meaning: Calibration not available. Reason: Parts replacements. Cure: Perform shading calibration.
3227 | PIS_UpdateCalReadFromFile | Error | Loading the digitizer calibration data failed. Possible bad image quality. Please contact service to calibrate the digitizer. | Confirm | UpdateCalibrationLine failed: Could neither read data from scanner nor from file. | Meaning: Calibration not available. Reason: Parts replacements. Cure: Perform shading calibration.
3228 | PIS_UpdateCalSetDoseDetectionLevels | Error |  |  | SetDoseDetectionLevels failed due to exception. | Meaning: Dose detection calibration not available. Reason: Exception when writing calibration data to the digitizer. Cure: Restart PC.
3230 | PIS_SetDoseDetectionLevels | Error |  |  | SetDoseDetectionLevels failed due to exception. | Meaning: Dose detection calibration not available. Reason: Exception when writing calibration data to the digitizer. Cure: Restart PC.
3584 | Tw_General_Error | Error |  |  | This error is used by Twain, if no other error is specified explicitely.
3585 | Tw_FXfer_SaveBitMap | Error |  |  | Saving the image bitmap for image file transfer failed.
3586 | Tw_MXfer_NoImageData | Error | Image transmission failed. System is retrying. If image does not appear, please restart the PC. |  | ReadImageData returned 0 bytes at image memory transfer. | Meaning: Communication problem between PC and digitizer. Reason: There is a severe transmission problem for image data. The automatic retries did not help. Cure: First try to recover the image by restarting the PC. If that does not help, the image is lost and you have to switch off and on the digitizer. Check Firewire cable, Firewire card in PC, PMI interface.
3587 | Tw_MXfer_ReadImageData | Error |  |  | ReadImageData failed at image memory transfer.
3588 | Tw_MXfer_ReadImageData2 | Error |  |  | ReadImageData failed with unexpected exception at image memory transfer.
3589 | Tw_InitInst_Configuration | Error |  |  | Error reading TWAIN configuration data.
3590 | Tw_CommandFailed | Error |  |  | A command did not succeed, but no specific error code was returned.
3591 | Tw_NotIdentified | Error | Cassette not identified. Remove and re-identify the cassette. | Confirm | Cassette is not identified. | Meaning: Cassette is not identified and direct ID not applicable. Reason: Cassette put into digitizer without prior identification. Cure: Identify cassette.
3592 | Tw_RejectedByPS | Error | The processing station refused to accept an image with the ID data read. Re-identify the cassette or process it where it had been identified. | Confirm | Cassette was rejected by processing station. | Meaning: Cassette was refused by processing station. Reason: Any wrong IPData. For details see display of processing station. Cure: Reidentification of cassette.
3604 | Tw_DIB_NoImageData | Error |  |  | ReadImageData returned 0 bytes while storing the image to a bitmap.
3605 | Tw_DIB_ReadImageData | Error |  |  | ReadImageData failed while storing the image to a bitmap.
3606 | Tw_DIB_ReadImageData2 | Error |  |  | ReadImageData failed with unexpected exception while storing the image to a bitmap.
3624 | Tw_Init_Caps | Error |  |  | Initialising the TWAIN capabilities failed.
3625 | Tw_Init_Caps2 | Error |  |  | Initialising the TWAIN capabilities failed with unexpected exception.
3626 | Tw_Update_Caps | Error |  |  | Updating the TWAIN capabilities failed.
3627 | Tw_Update_Caps2 | Error |  |  | Updating the TWAIN capabilities failed with unexpected exception.
3628 | Tw_NoSupportDoseDetection | Error |  |  | Digitizer does not support dose detection, although required. | Meaning: Dose detection calibration not available or dose detection not supported, although required. Reason: Dose detection was not calibrated at production time or wrong PMI board. Cure: Replace Photomultiplier. Spare part photomultiplier assemblies contain calibration data for dose detection. If PM ok, relace PMI board.
3634 | Tw_Cap_GetIPData | Error |  |  | Getting the image plate tag data failed.
3635 | Tw_Cap_GetIPData2 | Error |  |  | Getting the image plate tag data failed with unexpected exception.
3636 | Tw_Cap_ImportIPData | Error |  |  | Importing the image plate tag data failed.
3637 | Tw_Cap_SetIPData | Error |  |  | Setting the image plate tag data failed.
3638 | Tw_Cap_SetIPData2 | Error |  |  | Setting the image plate tag data failed with unexpected exception.
3639 | Tw_Cap_ImportIPData2 | Error |  |  | Importing the image plate tag data failed after setting the data.
3640 | Tw_Cap_GetXResolution | Error |  |  | Getting the resolution of the scanner failed.
3641 | Tw_Cap_GetXResolution2 | Error |  |  | Getting the resolution of the scanner failed with unexpected exception.
3642 | Tw_Cap_SetXResolution | Error |  |  | Setting the resolution of the scanner failed.
3643 | Tw_Cap_SetXResolution2 | Error |  |  | Setting the resolution of the scanner failed with unexpected exception.
3644 | Tw_Cap_GetFrame | Error |  |  | Getting the scanned frame from the scanner failed.
3645 | Tw_Cap_GetFrame2 | Error |  |  | Getting the scanned frame from the scanner failed with unexpected exception.
3646 | Tw_Cap_SetIPData3 | Error |  |  | Setting the image plate tag data returned with unexpected status.
3647 | Tw_Cap_GetSerialNumberTypeCode | Error |  |  | Suspicious scanner type code detected in GetSerialNumber. Probably not set correctly. | Meaning: Digitizer type identification invalid. Reason: Invalid system configuration. Cure: Upgrade software.
3648 | Tw_Cap_SetIPDataFailed | Error | Wrong cassette type (wrong application). Remove cassette. |  | Setting the image plate tag data failed. | Meaning: Setting the image plate tag data failed. Reason: IPData may contain parameters, which are not supported. Cure: Check the processing station menu settings.
3649 | Tw_Cap_SetIPDataNoReturnedData | Error |  |  | After setting the IPData, no current data was returned. Using data of set command.
3664 | Tw_SetArtefactSuppressionType | Error |  |  | Setting the artefact suppression type failed.
3665 | Tw_SetArtefactSuppressionType2 | Error |  |  | Setting the artefact suppression type failed with unexpected exception.
3684 | Tw_EnableFeeder_State | Error |  |  | The cassette detected event arrived at a wrong state.
3690 | Tw_RespEnableFeeder_Failed | Error |  |  | Enable feeder failed.
3691 | Tw_RespEnableFeeder_Warning | Error |  |  | Warning condition at enable feeder.
3692 | Tw_RespEnableFeeder_Status | Error |  |  | Enable feeder returned an unexpected status response.
3694 | Tw_DisableFeeder_State | Error |  |  | Disable feeder was called at wrong state.
3704 | Tw_FeedIP_State | Error |  |  | FeedIP was called at wrong state.
3709 | Tw_RespFeedIP_State | Error |  |  | The response to feedIP arrived at a wrong state.
3710 | Tw_RespFeedIP_Failed | Error |  |  | Feeding the image plate to the scan position failed.
3711 | Tw_RespFeedIP_Warning | Error |  |  | Feeding the image plate to the scan position returned with warning.
3712 | Tw_RespFeedIP_Status | Error |  |  | FeedIP returned an unexpected status response.
3714 | Tw_Scan_State | Error |  |  | Scan was called at wrong state.
3719 | Tw_RespScan_State | Error |  |  | The scan finished response arrived at a wrong state.
3720 | Tw_RespScan_Failed | Error |  |  | Scanning failed.
3721 | Tw_RespScan_Warning | Error | Problem during scan. The image may not have the required quality. |  | Warning condition during scan. | Meaning: Problem during scan. The image may not have the required quality. Reason: Not specific. Any warning condition during scan. Cure: Check other displayed errors or logging.
3722 | Tw_RespScan_Status | Error |  |  | Scan returned an unexpected status response.
3723 | Tw_RespScan_State2 | Error |  |  | The scan started response arrived at a wrong state.
3729 | Tw_RespErase_State | Error |  |  | The erasure finished response arrived at a wrong state.
3730 | Tw_RespErase_Failed | Error | Failure at erasure module. Please restart digitizer. If error remains, please contact service. Image plate must be erased after repair again. |  | Erasing the image plate failed. | Meaning: Failure at erasure module. Image plate must be erased after repair again. Reason: Problem in erasure module. Cure: Check and repair or exchange erasure module.
3731 | Tw_RespErase_Warning | Error |  |  | Erasing the image plate returned with warning.
3732 | Tw_RespErase_Status | Error |  |  | Erasing the image plate returned an unexpected status response.
3739 | Tw_RespReturnIP_State | Error |  |  | The response to ReturnIP arrived at a wrong state.
3740 | Tw_RespReturnIP_Failed | Error |  |  | Returning the image plate into the cassette failed.
3741 | Tw_RespReturnIP_Warning | Error |  |  | Warning condition, when returning the image plate into the cassette.
3742 | Tw_RespReturnIP_Status | Error |  |  | ReturnIP returned an unexpected status response.
3744 | Tw_RespEject_State | Error |  |  | The response to eject arrived at a wrong state.
3745 | Tw_RespEject_Failed | Error |  |  | Ejecting the cassette failed.
3746 | Tw_RespEject_Warning | Error |  |  | Warning condition, when ejecting the cassette.
3747 | Tw_RespEject_Status | Error |  |  | Eject returned an unexpected status response.
3754 | Tw_OpenImageData_State | Error |  |  | Opening the image data was requested in a state, where the scanner cannot handle it.
3755 | Tw_OpenImageData | Error |  |  | Opening the image data for read within the scanner failed.
3756 | Tw_OpenImageData2 | Error |  |  | Opening the image data for read within the scanner failed with unexpected exception.
3770 | Tw_CloseImageData | Error |  |  | Closing the image data within the scanner failed.
3771 | Tw_CloseImageData2 | Error |  |  | Closing the image data within the scanner failed with unexpected exception.
3772 | Tw_CloseImageDataNotTx | Error | Image transmission failed. Please restart the PC in order to retry getting the image. |  | Cannot close image data, because not yet (completely) transmitted. | Meaning: Image transmission failed. The image stays available in digitizer and is transmitted again after reconnection of NX. Reason: FireWire cable, PMI or PC FireWire board. Cure: Restart PC to try again getting the image. Check FireWire cable, PMI and PC FireWire board.
3775 | Tw_StartNotification | Error |  |  | Starting notifications by the scanner failed.
3776 | Tw_StartNotification2 | Error |  |  | Starting notifications by the scanner failed with unexpected exception.
3884 | Tw_AsyncCall_Offline | Error |  |  | An asynchronous call of the scanner cannot be executed, because the scanner is not accessible.
3885 | Tw_AsyncCall | Error |  |  | An asynchronous call of the scanner failed.
3886 | Tw_AsyncCall2 | Error |  |  | An asynchronous call of the scanner failed with unexpected exception.
3889 | Tw_SyncCall_Offline | Error |  |  | A synchronous call of the scanner cannot be executed, because the scanner is not accessible.
3890 | Tw_SyncCall | Error |  |  | A synchronous call of the scanner failed.
3891 | Tw_SyncCall2 | Error |  |  | A synchronous call of the scanner failed with unexpected exception.
3894 | Tw_GetProperty_Offline | Error |  |  | Getting a property from the scanner not possible, because the scanner is not accessible.
3895 | Tw_GetProperty | Error |  |  | Getting a property from the scanner failed.
3896 | Tw_GetProperty2 | Error |  |  | Getting a property from the scanner failed with unexpected exception.
3899 | Tw_SetProperty_Offline | Error |  |  | Setting a property in the scanner not possible, because the scanner is not accessible.
3900 | Tw_SetProperty | Error |  |  | Setting a property in the scanner failed.
3901 | Tw_SetProperty2 | Error |  |  | Setting a property in the scanner failed with unexpected exception.
3905 | Tw_RegisterAccessExclusive | Error |  |  | RegisterAccessExclusive failed.
3906 | Tw_RegisterAccessExclusive2 | Error |  |  | RegisterAccessExclusive failed with unexpected exception.
3910 | Tw_UnregisterAccessExclusive | Error |  |  | UnregisterAccessExclusive failed.
3911 | Tw_UnregisterAccessExclusive2 | Error |  |  | UnregisterAccessExclusive failed with unexpected exception.
3912 | Tw_NoExclusiveAccess | Error |  |  | We did not get exclusive access to the scanner, as we expected.
3915 | Tw_PrepareImageCorrection | Error |  |  | Preparing the image correction failed.
3916 | Tw_PrepareImageCorrection2 | Error |  |  | Preparing the image correction failed with unexpected exception.
3964 | Tw_Wait_FeedIP_Timeout | Error |  |  | Timeout while waiting for the IP to be at the scanner stage.
3965 | Tw_Wait_FeedIP_State | Error |  |  | Unexpected state while waiting for the IP to be at the scanner stage.
3966 | Tw_Wait_ScanStart_Timeout | Error |  |  | Timeout while waiting for scan started.
3967 | Tw_Wait_ScanStart_State | Error |  |  | Unexpected state while waiting for scan started.
3968 | Tw_Wait_ScanEnd_Timeout | Error |  |  | Timeout while waiting for end of scan.
3969 | Tw_DirectIDTimeout | Error | ID button not pressed within the timeout period. Image plate was not scanned. Cassette stays clamped until ID data are supplied by the operator. |  | Timeout while waiting for identification. Scan was cancelled. Please try again. | Meaning: IP was moved back into cassette. Reason: ID button not pressed within the timeout period. Image plate was not scanned. Cassette stays clamped until ID data are supplied by the operator. Cure: Identify this cassette with NX.
3984 | Tw_ConnectionTimer | Error |  |  | Starting the remote connection timer failed.
3985 | Tw_ConnectionTimer2 | Error |  |  | Starting the remote connection timer failed with unexpected exception.
3986 | Tw_RegisterChannel | Error |  |  | First try to register the remoting channel failed.
3987 | Tw_RegisterChannel2 | Error |  |  | First try to register the remoting channel failed with unexpected exception.
3988 | Tw_UnregisterChannel | Error |  |  | Unregistering the remoting channel failed.
3989 | Tw_UnregisterChannel2 | Error |  |  | Unregistering the remoting channel failed with unexpected exception.
3990 | Tw_ConnectErrorManager | Error |  |  | First try to connect to the error manager failed.
3991 | Tw_ConnectErrorManager2 | Error |  |  | First try to connect to the error manager failed with unexpected exception.
3992 | Tw_DisconnectErrorManager | Error |  |  | Unregistering the error manager event handler failed.
3993 | Tw_DisconnectErrorManager2 | Error |  |  | Unregistering the error manager event handler failed with unexpected exception.
3994 | Tw_ConnectServer | Error |  |  | First try to connect to the server failed.
3995 | Tw_ConnectServer2 | Error |  |  | First try to connect to the server failed with unexpected exception.
3996 | Tw_DisconnectServer | Error |  |  | Unregistering the server event handler failed.
3997 | Tw_DisconnectServer2 | Error |  |  | Unregistering the server event handler failed with unexpected exception.
4004 | Tw_CreateFileMapping | Error |  |  | CreateFileMapping for image transmission failed.
4005 | Tw_MapViewOfFile | Error |  |  | MapViewOfFile for image transmission failed.
4006 | Tw_AttachFileMapping | Error |  |  | Attaching the file mapping to server failed.
4007 | Tw_AttachFileMapping2 | Error |  |  | Attaching the file mapping to server failed with unexpected exception.
4008 | Tw_DetachFileMapping | Error |  |  | Detaching the file mapping from server failed.
4009 | Tw_DetachFileMapping2 | Error |  |  | Detaching the file mapping from server failed with unexpected exception.
4014 | Tw_SetTimeout | Error |  |  | Setting the timeout for a command failed.
4105 | CONFIRM_CASSETTE_WITH_IP_FOUND_AT_START_UP | Error | Image plate during self-test in the digitizer detected. The image plate was moved back into the cassette. Please check whether the image is present at the PC and re-scan the cassette, if necessary. If the image is present, the cassette was not erased possibly. Erase the cassette again in any case. | Confirm | Start-Up: Cassette with IP found, no IP is in slowscan or transport unit | Meaning: Cassette with an IP found during start-up. Reason: Digitizer was switched off during cycle. Cure: Confirm message, remove cassette; check if image was sent to the NX; if not, scan cassette again; erase cassette again in any case.
4106 | CONFIRM_EMPTY_CASSETTE_FOUND_AT_START_UP | Error | Empty cassette found at start-up. Remove cassette. | Confirm | Start-Up: Cassette without an IP found, no IP is in slowscan or transport unit | Meaning: Cassette found during start-up, but no IP. Reason: Cassette without IP was inserted and digitizer was switched off. Cure: Confirm message, remove cassette; put a tray and an IP into the cassette.
4107 | IP_WITHOUT_CASSETTE_FOUND_AT_START_UP_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4108 | WRONG_IHC_PROTVER | Error |  |  | Protocol version of IHC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of IHC (IP-HandlingControl) is wrong. Cure: Install new software; inform R&D team.
4109 | WRONG_CHC_PROTVER | Error |  |  | Protocol version of CHC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of CHC (CassetteHandlingControl) is wrong. Cure: Install new software; inform R&D team.
4110 | WRONG_TRC_PROTVER | Error |  |  | Protocol version of TRC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of TRC (TagReaderControl) is wrong. Cure: Install new software; inform R&D team.
4111 | WRONG_ICC_PROTVER | Error |  |  | Protocol version of ICC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ICC (ImageCaptureControl) is wrong. Cure: Install new software; inform R&D team.
4112 | PREPARE_SCAN_ERROR | Error |  |  | prepareScan returned an error | Meaning: The preparation of the scan failed. Reason: Scan unit (PMT, lens assembly) error. Cure: Check underlying scan unit error description.
4113 | ANALYZE_PREPARE_SCAN_ERROR | Error |  |  | unknown signal for analyzePrepareScanResult | Meaning: Incompatible SW components. Reason: SW component DigitizerControl receives an unknown event from SW component ImageCaptureControl. Cure: Install new software; inform R&D team.
4114 | SCAN_REJECTED_PREPARE_SCAN_FAILED | Error |  |  | scan rejected because prepare scan failed | Meaning: Scan rejected because preparation of scan failed before. Reason: Scan unit (PMIT, lens assembly) detected an error, thus scaning cannot start and is rejected. Cure: Check underlying scan unit error description.
4115 | SCAN_REJECTED_SCAN_NOT_PREPARED | Error |  |  | scan rejected because prepare scan was not executed before | Meaning: Scan rejected because scan was not prepared before (SW error). Reason: Scan unit (PMIT, lens assembly) was not prepared for scan, thus scaning cannot start and is rejected. Cure: Install new software; inform R&D team.
4116 | WRITING_TAG_FAILED_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4117 | ICC_SELFTEST_FAILED | Error |  |  | selftest of ICC failed | Meaning: Selftest failed. Reason: ICC (ImageCaptureControl) detected an error during selftest. Cure: Check underlying error description.
4118 | CHC_SELFTEST_FAILED | Error |  |  | selftest of CHC failed | Meaning: Selftest failed. Reason: CHC (CassetteHandlingControl) detected an error during selftest. Cure: Check underlying error description.
4119 | IHC_SELFTEST_FAILED | Error |  |  | selftest of IHC failed | Meaning: Selftest failed. Reason: IHC (IP-HandlingControl) detected an error during selftest. Cure: Check underlying error description.
4120 | TRC_SELFTEST_FAILED | Error |  |  | selftest of TRC failed | Meaning: Selftest failed. Reason: TRC (TagReaderControl) detected an error during selftest. Cure: Check underlying error description.
4121 | PREPARE_SCAN_WARNING | Error |  |  | prepareScan returned a warning | Meaning: The preparation of the scan finished with a warning, but IP can be scanned. Reason: Scan unit (PMT, lens assembly) warning. Cure: Check description of underlying scan unit warning.
4122 | SCAN_WARNING_ICC | Error |  |  | ICC finished scan with warning | Meaning: The scan finished with a warning. Reason: ICC (ImageCaptureControl) checked the scan and reported a scan unit (PMT, lens assembly) warning. Cure: Check description of underlying scan unit warning.
4123 | CONFIRM_WRONG_DATAFLOW_SPEC | Error | Wrong cassette type (wrong cassette chip initialization). Remove cassette. | Confirm | cassette is initialized with wrong data flow spec; spec version must be >= 1.3 | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong data flow specification (< version 1.3) Cure: Check if cassette is designated to be used in CR30; check initialization of the cassette.
4124 | CONFIRM_WRONG_IP_TYPE | Error | Wrong cassette type (wrong image plate type). Remove cassette. | Confirm | wrong image plate type; must be a flexible image plate | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong setting of the image plate type; the image plate type must be set to 'flexible'. Cure: Check if cassette is designated to be used in CR30; check initialization of the cassette.
4125 | CONFIRM_WRONG_APP_TYPE | Error | Wrong cassette type (wrong application). Remove cassette. | Confirm | wrong application type; must be GenRad HR | Meaning: Cassette is rejected because cassette is initialized wrongly. Reason: Cassette is rejected due to a wrong setting of the application type; the application type must be set to 'GenRad HR'. Cure: Check if cassette is designated to be used in CR30; check initialization of the cassette.
4126 | CONFIRM_UNSUPPORTED_IP_SIZE | Error | Wrong cassette type (unsupported size). Remove cassette. | Confirm | unsupported IP size | Meaning: Cassette is rejected because IP size is not supported or cassette is initialized with unsupported size. Note: The cassette format 35x35 cm is not supported if the product serial number is smaller than 30000 for standard digitizer (subtype 100) resp. 40000 for white label digitizer (subtype 110). Reason: Cassette is rejected due to a wrong setting of the IP size or digitizer does not support 35x35 cm because digitizer is too old (see above). Cure: Check if cassette is designated to be used in CR30; check initialization of the cassette. If this error occurs with format 35x35 cm check the product serial number and the subtype.
4127 | CONFIRM_WRONG_IP_SIZE | Error | Wrong cassette size initialization. Remove cassette. | Confirm | cassette initialized with wrong IP size | Meaning: Cassette is rejected because IP size is wrong or cassette is initialized with wrong size. Reason: Cassette is rejected due to a wrong setting of the IP size. Cure: Check if cassette is designated to be used in CR30; check initialization of the cassette.
4128 | CONFIRM_EMPTY_CASSETTE_AND_IP_FOUND_AT_START_UP | Error | Image plate during self-test in the digitizer detected. The image plate was moved back into the cassette. Please check whether the image is present at the PC and re-scan the cassette, if necessary. If the image is present, the cassette was not erased possibly. Erase the cassette again in any case. | Confirm | Start-Up: Cassette without IP found, IP is in slowscan or transport unit | Meaning: Cassette and IP was found in digitizer during start-up. Reason: Cassette was inserted and digitizer was switched off during the cycle. Cure: Confirm message, remove cassette; check if image was sent to the NX; if not, scan cassette again; erase cassette again in any case.
4129 | CONFIRM_ENABLE_TAGREADER_FAILED | Error | Cassette chip not readable or cassette tray is missing. Remove cassette. | Confirm | Enabling of tag reader failed | Meaning: Enabling the RF-tag reader failed. Reason: RF-tag on cassette tray is defective or chipreader is defective or not connected. Cure: Try it again with another cassette; check cable between IP-handling board and chipreader; check if chipreader is mounted correctly; if still NOK, replace chipreader, then IP-handling board.
4130 | WRONG_UID_PROTVER | Error |  |  | Protocol version of UID is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of UID (UserInterfaceDispatcher) is wrong. Cure: Install new software; inform R&D team.
4131 | TIMEOUT_CHC_ENABLE_CASS_DETECTION | Error |  |  | response timeout for CHC command enableCassDetection | Meaning: Communication timeout between PMI board and IP-handling board. Reason: The SW component CHC (CassetteHandlingControl) on IP-handling board didn't reply to the command 'enableCassDetection' in time. Cure: Check CAN connection between PMI board and IP-handling board.
4132 | TIMEOUT_CHC_DISABLE_CASS_DETECTION | Error |  |  | response timeout for CHC command disableCassDetection | Meaning: Communication timeout between PMI board and IP-handling board. Reason: The SW component CHC (CassetteHandlingControl) IP-handling board didn't reply to the command 'disableCassDetection' in time. Cure: Check CAN connection between PMI board and IP-handling board.
4133 | RELEASE_CASSETTE_FAILED_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4134 | FATAL_RELEASE_CASSETTE_FAILED | Error |  |  | releaseCassette failed | Meaning: Releasing the cassette failed. Reason: CHC (CassetteHandlingControl) on IP-handling board detected an error during releasing of the cassette. Cure: Check underlying error description.
4135 | CONFIRM_CASSETTE_WITH_IP_AND_IP_FOUND_AT_START_UP_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4136 | NOBODY_WAITED_ON_ERROR_CONFIRMATION | Error |  |  | Nobody waited on the confirmation of an error, just remove the displayed error on PC | Meaning: An error message was confirmed, but no SW module waited on the confirmation of the specific error code. Reason: Digitizer was disconnected while an error message was displayed on the PC. Cure: None. This is only an internal warning for the R&D team.
4137 | UNEXPECTED_SIGNAL | Error |  |  | received RoseRT signal cannot be handled in the current state | Meaning: A SW component received a command or response which cannot be handled in the current state. Reason: SW error. Cure: Install new software; inform R&D team.
4138 | FATAL_ERROR | Error |  |  | fatal error, going into error state | Meaning: Fatal error occurred, digitizer is in the error state and terminated current cycle. Reason: A SW component detected a fatal error. Cure: Check underlying error description.
4139 | FATAL_INITIALIZATION_FAILED | Error |  |  | fatal error during initialization | Meaning: Fatal error occurred during initialization, digitizer is in the error state. Reason: A SW component detected a fatal error. Cure: Check underlying error description.
4140 | FATAL_SCAN_MODE_FAILED | Error |  |  | fatal error in scan mode | Meaning: Fatal error occurred during a normal scan cycle, digitizer is in the error state and terminated current cycle. Reason: A SW component detected a fatal error. Cure: Check underlying error description.
4141 | FATAL_ERASE_MODE_FAILED | Error |  |  | fatal error in erase mode | Meaning: Fatal error occurred during an erasure cycle, digitizer is in the error state and terminated current cycle. Reason: A SW component detected a fatal error. Cure: Check underlying error description.
4142 | CMD_CURRENTLY_NOT_ALLOWED | Error |  |  | DIC received the signal CmdCurrentlyNotAllowed | Meaning: DIC (DigitizerControl) received an reply from another SW component, that the command is currently not allowed. Reason: SW error. Cure: Install new software; inform R&D team.
4143 | SEND_FATAL_INVALID_PORT | Error |  |  | method sendFatalResponsesToPC received invalid port | Meaning: DIC's (DigitizerControl) method sendFatalResponsesToPC() received an invalid port. Reason: SW error. Cure: Install new software; inform R&D team.
4144 | FATAL_CMD_ENABLE_FEEDER | Error |  |  | enableFeeder stopped with FATAL response | Meaning: PC command 'enableFeeder' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4145 | FATAL_CMD_DISABLE_FEEDER | Error |  |  | PC command 'disableFeeder' stopped with FATAL response | Meaning: PC command 'disableFeeder' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4146 | FATAL_CMD_EJECT | Error |  |  | PC command 'eject' stopped with FATAL response | Meaning: PC command 'eject' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4147 | FATAL_CMD_ERASE | Error |  |  | PC command 'erase' stopped with FATAL response | Meaning: PC command 'erase' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4148 | FATAL_CMD_FEED_IP | Error |  |  | PC command 'feedIP' stopped with FATAL response | Meaning: PC command 'feedIP' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4149 | FATAL_CMD_RETURN_IP | Error |  |  | PC command 'returnIP' stopped with FATAL response | Meaning: PC command 'returnIP' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4150 | FATAL_CMD_SCAN | Error |  |  | PC command 'scan' stopped with FATAL response | Meaning: PC command 'scan failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4151 | FATAL_CMD_START_SELFTEST | Error |  |  | PC command 'startSelftest' stopped with FATAL response | Meaning: PC command 'startSelftest' failed with fatal error. Reason: Any fatal error while command was executed. Cure: Restart the digitizer to execute selftest.
4152 | CONFIRM_UNKNOWN_FORMAT_DETECTED | Error | Unknown cassette format detected (cassette inserted correctly?). Remove cassette and insert it again. | Confirm | cassette fixer detected unknown format | Meaning: The MOTOR CASSETTE FIXING M1 detected an unknown cassette format. The cassette is still clamped and a user message is displayed to press confirm to release the cassette. The digitizer is still working, but if this error occurs frequently, it requires service support. Reason: The MOTOR CASSETTE FIXING M1 measured a cassette width which doesn't match with a known cassette format. The determined cassette width is too small or too big due to: - The cassette was inserted wrongly by the user (skewed, upside down or not with the flap first) - The CASSETTE FIXATION STEEL BAND is not moving smoothly - The CASSETTE DETECTION HOOK is not properly fixed or not the improved version - Load detection for MOTOR CASSETTE FIXING M1 is imperfect, STALL CALIBRATION was not successfully - Drive mechanism of MOTOR CASSETTE FIXING M1 is blocked or defect Cure: 1. Remove the cassette and insert it correctly again. 2. Check if cassette can be pushed smoothly into the right edge of the input slot and that the wedge of the cassette fixer presses correctly into the left notch of the cassette. 3. Insert another cassette of the same (preferred) or any other size. 4. Check if the CASSETTE FIXATION STEEL BAND and the cassette fixing motor is moving smoothly. 5. Check if the CASSETTE DETECTION HOOK is properly fixed. If applicable mount the improved CASSETTE DETECTION HOOK, refer to Service Bulletin No.35 6. Perform a calibration of the MOTOR CASSETTE FIXING M1, refer to the Service Manual 7. Replace the MOTOR CASSETTE FIXING M1, refer to the Service Manual for order and replacement 8. Replace the IP HANDLING BOARD, refer to the Service Manual for order and replacement 9. Replace the complete CASSETTE FIXATION UNIT, refer to the Service Manual for order and replacement
4153 | TRC_INITIALIZATION_FAILED | Error |  |  | initialization of TRC failed | Meaning: Start-up failed. Reason: Initialization of the TRC (TagReaderControl) failed. Cure: Check underlying error description.
4154 | CLEANING_UP_CHC_IHC_FAILED | Error |  |  | clean-up failed because IHC or CHC failed. See preceding error. | Meaning: Clean-up during start-up failed. Reason: Clean-up failed, because IHC (IP-HandlingControl) or CHC (CassetteHandlingControl) detected an error. Cure: Check underlying error description.
4155 | INIT_CHC_FAILED | Error |  |  | initialization of CHC failed | Meaning: Start-up failed. Reason: Initialization of the CHC (CassetteHandlingControl) failed. Cure: Check underlying error description.
4156 | INIT_IHC_FAILED | Error |  |  | initialization of IHC failed | Meaning: Start-up failed. Reason: Initialization of IHC (IP-HandlingControl) failed. Cure: Check underlying error description.
4157 | CONFIRM_WRITING_TAG_FAILED | Error | Error during writing on cassette chip after successful scan. Erase the image plate again. | Confirm | writing to RF-tag failed | Meaning: Writing data on RF-tag failed. Reason: RF-tag on cassette tray is defective or chipreader is defective or not connected. Cure: Try it again with another cassette; check cable between IP-handling board and chipreader; check if chipreader is mounted correctly; if still NOK, replace chipreader, then IP-handling board.
4158 | CONFIRM_ENABLE_TAGREADER_FOR_WRITING_FAILED | Error | Error during writing on cassette chip after successful scan. Erase the image plate again. | Confirm | enabling tag reader for writing to RF-tag failed | Meaning: Enabling the tag reader for writing data on RF-tag failed. Reason: RF-tag on cassette tray is defective or chipreader is defective or not connected. Cure: Try it again with another cassette; check cable between IP-handling board and chipreader; check if chipreader is mounted correctly; if still NOK, replace chipreader, then IP-handling board.
4159 | CONFIRM_ENABLE_TAGREADER_FAILED_ERASE_MODE | Error | Error during writing on cassette chip after successful erasure cycle. Erase the image plate again. | Confirm | enabling tag reader for writing to RF-tag failed during erasure cycle | Meaning: Enabling the tag reader for writing data on RF-tag failed during erasure cycle. Reason: RF-tag on cassette tray is defective or chipreader is defective or not connected. Cure: Try it again with another cassette; check cable between IP-handling board and chipreader; check if chipreader is mounted correctly; if still NOK, replace chipreader, then IP-handling board.
4160 | CONFIRM_WRITING_TAG_FAILED_ERASE_MODE | Error | Error during writing on cassette chip after successful erasure cycle. Erase the image plate again. | Confirm | writing to RF- tag failed during erasure cycle | Meaning: Writing data on RF-tag failed during erasure cycle. Reason: RF-tag on cassette tray is defective or chipreader is defective or not connected. Cure: Try it again with another cassette; check cable between IP-handling board and chipreader; check if chipreader is mounted correctly; if still NOK, replace chipreader, then IP-handling board.
4161 | WARNING_READ_TAGDATA_FAILED | Error |  |  | Reading of tag data failed | Meaning: Warning, reading data from RF-tag failed. Reason: During start-up we try to read the RF-tag, but if the cassette is open at start-up, we cannot read the tag data, because the reader is not close enough to the RF-tag. Cure: If error remains during normal cycles, try it again with another cassette; check cable between IP-handling board and chipreader; check if chipreader is mounted correctly; if still NOK, replace chipreader, then IP-handling board.
4162 | DESERIALIZATION_IPSTATUS_FAILED | Error |  |  | could not deserialize IPStatus | Meaning: Could not deserialize IPStatus data. Reason: Check underlying error. Cure: Check underlying error.
4163 | DESERIALIZATION_IPINFO_FAILED | Error |  |  | could not deserialize IPInfo | Meaning: Could not deserialize IPInfo data. Reason: Check underlying error. Cure: Check underlying error.
4164 | DESERIALIZATION_IPCCTRL_FAILED | Error |  |  | could not deserialize IPCtrl | Meaning: Could not deserialize IPCtrl data. Reason: Check underlying error. Cure: Check underlying error.
4165 | DESERIALIZATION_IPPSDATA_FAILED | Error |  |  | could not deserialize IPPSData | Meaning: Could not deserialize IPPSData. Reason: Check underlying error. Cure: Check underlying error.
4166 | IP_ERASED_ALREADY_DUE_TO_ERASE_TIMEOUT | Error |  |  | PC command 'erase' received, but IP was already erased automatically to prevent bowing of IP | Meaning: PC command 'erase' received, but IP was already erased automatically to prevent bowing of IP Reason: After end of scan, DigitizerControl waits 10 sec. for PC command 'erase'. After that time, it erases the IP and moves the IP automatically back into the cassette to prevent bowing of IP. Cure: None (no problem, digitizer is still working properly)
4167 | CAN_ADAPTER_ERROR | Error |  |  | fatal error in CAN adapter layer | Meaning: Fatal error in CAN adapter layer Reason: CAN adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help replace PMI board.
4168 | LOGGING_ADAPTER_ERROR | Error |  |  | fatal error in logging adapter layer | Meaning: Fatal error in logging adapter layer Reason: Logging adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help replace PMI board.
4169 | FW_ADAPTER_ERROR | Error |  |  | fatal error in FW adapter layer | Meaning: Fatal error in FireWire (FW) adapter layer Reason: FW adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and FW connection; update SW; if this doesn't help replace PMI board.
4170 | CAN_VCC_FAILURE | Error |  |  | CAN Vcc failure | Meaning: CAN voltage error; Reason: PMI power supply fuse F23 broken, cable power supply to PMI broken or PMI power supply defect; Cure: Replace F23 [1A T], replace cable power supply to PMI, replace PMI power supply
4171 | PMM_PLUS20V_FAILURE | Error |  |  | PMM +20V failure | Meaning: The check of the +20 V supply voltage for the Photomultiplier failed. The Digitizer is out of operation and requires service support. Reason: Check of the +20V supply voltage failed due to: - Defective flat cable between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD, broken wires or bad contacts - Defective PMI POWER DISTRIBUTION BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD 3. Exchange the PMI POWER DISTRIBUTION BOARD
4172 | PMM_MINUS10V_FAILURE | Error |  |  | PMM -10V failure | Meaning: The check of the -10 V supply voltage for the Photomultiplier failed. The Digitizer is out of operation and requires service support. Reason: Check of the -10V supply voltage failed due to: - Defective flat cable between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD, broken wires or bad contacts - Defective PMI POWER DISTRIBUTION BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD 3. Exchange the PMI POWER DISTRIBUTION BOARD
4173 | ANALOG_VCC_FAILURE | Error |  |  | Analog Vcc failure | Meaning: The check of the +8 V supply voltage for A/D conversion on the PMI Board failed. The Digitizer is out of operation and requires service support. Reason: Check of the +8V supply voltage failed due to: - Defective flat cable between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD, broken wires or bad contacts - Defective PMI POWER DISTRIBUTION BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD 3. Exchange the PMI POWER DISTRIBUTION BOARD
4174 | OPTICS_PLUS12V_FAILURE | Error |  |  | optics +12V failure | Meaning: The check of the +12 V supply voltage for the optic module failed. The Digitizer is out of operation and requires service support. Reason: Check of the +12V supply voltage failed due to: - Defective flat cable between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD, broken wires or bad contacts - Defective PMI POWER DISTRIBUTION BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD 3. Exchange the PMI POWER DISTRIBUTION BOARD
4175 | OPTICS_MINUS12V_FAILURE | Error |  |  | optics -12V failure | Meaning: The check of the -12 V supply voltage for the optic module failed. The Digitizer is out of operation and requires service support. Reason: Check of the -12V supply voltage failed due to: - Defective flat cable between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD, broken wires or bad contacts - Defective PMI POWER DISTRIBUTION BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between the PMI POWER DISTRIBUTION BOARD and the PMI BOARD 3. Exchange the PMI POWER DISTRIBUTION BOARD
4176 | DISPLAY_VCC_FAILURE | Error |  |  | Vcc failure for display unit | Meaning: Display voltage error; Reason: PMI power supply fuse F20 broken, cable power supply to PMI broken or PMI power supply defect; Cure: Replace F20 [2.5A T], replace cable power supply to PMI, replace PMI power supply
4177 | WRONG_RAMP_PIXEL_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4178 | INIT_BOARD_FAILED | Error |  |  | initBoard() failed | Meaning: Initialization of PMI board failed. Reason: Check underlying error. Cure: Check underlying error.
4179 | CHC_COMMUNICATION_FAILED | Error |  |  | no communication possible with CHC on IP-Handling board | Meaning: Communication error between PMI board and IP-handling board. Cure: Check CAN connection
4180 | FATAL_IHC_NO_RECEIVER_FOUND_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4181 | FATAL_IHC_UNKNOWN_CHANNEL_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4182 | FATAL_IHC_ERR_READING_ACKNOWLEDGE_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4183 | FATAL_IHC_CAN_ERROR_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4184 | FATAL_CHC_NO_RECEIVER_FOUND_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4185 | FATAL_CHC_UNKNOWN_CHANNEL_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4186 | FATAL_CHC_ERR_READING_ACKNOWLEDGE_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4187 | FATAL_CHC_CAN_ERROR_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4188 | FATAL_TRC_NO_RECEIVER_FOUND_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4189 | FATAL_TRC_UNKNOWN_CHANNEL_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4190 | FATAL_TRC_ERR_READING_ACKNOWLEDGE_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4191 | FATAL_TRC_CAN_ERROR_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4192 | FATAL_UIC_NO_RECEIVER_FOUND_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4193 | FATAL_UIC_UNKNOWN_CHANNEL_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4194 | FATAL_UIC_ERR_READING_ACKNOWLEDGE_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4195 | FATAL_UIC_CAN_ERROR_OBSOLETE | Error |  |  | OBSOLETE | OBSOLETE
4196 | FEEDER_DISABLED_REJECT_CASSETTE | Error | Cassette rejected because processing station is not ready anymore. Remove cassette. | Confirm | cassette is rejected because feeder was disabled during clamping of cassette | Meaning: Cassette is rejected because feeder was disabled during clamping of cassette. Reason: Processing station is not ready anymore to acquire images and therefore disabled the feeder of the digitizer. Cure: No problem; remove cassette, reboot PC
4197 | FATAL_REJECT_CASSETTE | Error |  |  | rejectCassette failed | Meaning: Rejecting the cassette failed with a fatal error. Reason: Check underlying error. Cure: Check underlying error.
4198 | DIGITIZER_START_UP | Info | Digitizer start-up. |  | digitizer is starting-up | Meaning: Digitizer is starting-up or not yet enabled to process inserted cassettes after start-up. Reason: Digitizer is starting-up or not yet enabled by the processing station to take cassettes. Cure: Check if digitizer is perfoming the start-up (you can motor movements). If you hear movements, wait until they are finished. If you hear nothing, restart the digitizer; if this doesn't help or if message doesn't disappear a few minutes after end of digitizer's start-up, reboot the PC.
4199 | TIMEOUT_SCAN_END_ICC | Error | Problem during scan. The image may not have the required quality. |  | timeout for 'scan' response from ICC after slowscan stopped | Meaning: The scanning unit didn't finish the scan process although the slowscan motor has already stopped. Reason: ICC (ImageCaptureControl) is still waiting for end of the scan, i.e. the required number of pixels have not been stored in the SDRAM of the FPGA. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI board and PMT/lens assembly; if OK, replace PMI; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly.
4200 | OPCODE_NOT_SUPPORTED | Error |  |  | operation code not supported | Meaning: The operation code given by the NX is not supported. (For example, FLFS has a special operation code.) Reason: Digitizer rejects the required operation mode, because it doesn't support this mode. Cure: Check if digitizer is configured correctly (i.e. as CR30-X) on the NX; if OK, inform R&D team.
4201 | CONFIRM_CLEAN_UP_CASS_CLOSED_ENABLE_TAGREADER_FAILED | Error | Cassette chip not readable or cassette tray is missing. Remove cassette, if present. If no cassette is present in the input slot, please contact service to fix the problem. | Confirm | Enabling of tag reader failed during clean-up after closing the cassette. | Meaning: Enabling of tag reader failed during start-up after closing the cassette. Reason: RF-tag not found. Cure: If a cassette is present in the input slot during start-up, check if cassette contains a tray with a RF-tag. If no cassette is present in the input slot during start-up, check if cassette detection sensor (edge sensor, microswitch on cassette holder) is connected and working properly.
4202 | CONFIRM_IP_WITHOUT_CASSETTE_FOUND_AT_START_UP | Error | Start-up: An image plate was found in the digitizer, but no cassette. The image plate will be moved to the cassette slot as soon as you confirm. Please confirm and withdraw the image plate. | Confirm | Start-up: an IP was found in the Digitizer, but no cassette. The IP will be transported to the cassette slot as soon as you confirm. Please confirm and withdraw the IP. | Meaning: An IP was found in the Digitizer, but no cassette during start-up. Reason: IP-jam occured before or manipulation by hand. Cure:The IP will be transported to the cassette slot as soon as you confirm. Please confirm and withdraw the IP.
4203 | FATAL_CMD_CLEAN_UP | Error |  |  | PC command 'cleanUp' stopped with FATAL response | Meaning: Fatal digitizer error occured during PC command 'cleanUp' was executed. Reason: Check underlying error. Cure: Check underlying error.
4204 | CLEAN_UP_IN_ERASURE_MODE | Error |  |  | digitizer received PC command 'cleanUp' in erasure mode; this is not allowed (SW error) | Meaning: Digitizer received PC command 'cleanUp' in erasure mode; this is not allowed. Reason: SW error. Cure: Install new software; inform R&D team.
4205 | CONFIRM_CLEAN_UP_BEFORE_SCANNING | Error | Image plate during self-test in the digitizer detected. The image plate was moved back into the cassette. Please check whether the image is present at the PC and re-scan the cassette, if necessary. If the image is present, the cassette was not erased possibly. Erase the cassette again in any case. | Confirm | digitizer received PC command 'cleanUp' before scanning started; the IP was not scanned and is transported back into the cassette | Meaning: Digitizer received PC command 'cleanUp' before scanning started; the IP was not scanned and is transported back into the cassette. Reason: PC connected to digitizer while digitizer was performing a cycle. Cure: Confirm message; remove cassette; insert cassette again to scan the image plate.
4206 | CLEAR_IMAGE_TX_ACTIVE_DURING_SCANNING | Error |  |  | digitizer received PC command 'imageTxActive(FALSE)' during scanning; this is not allowed (SW error) | Meaning: Digitizer received PC command 'imageTxActive(FALSE)' during scanning; this is not allowed. Reason: SW error. Cure: Install new software; inform R&D team.
4207 | INIT_CHC_OPENER_FAILED | Error |  |  | initialization of CHC's opener drive failed | Meaning: Initialization of CHC's (CassetteHandlingControl) opener drive failed. Reason: Cassette slot is blocked by something. Cure: Remove cassette or any other object out of input slot and confirm message. Then the opener drive will be tested again.
4208 | CONFIRM_REMOVE_CASS_FROM_SLOT | Error | Cassette slot is blocked by something. Check cassette slot and remove object. | Confirm | initialization of CHC's opener drive failed; remove cassette out of input slot | Meaning: Initialization of CHC's (CassetteHandlingControl) opener drive failed. Reason: Cassette slot is blocked by something. Cure: Remove cassette or any other object out of input slot and confirm message. Then the opener drive will be tested again and a dedicated error message is displayed if initialization fails again.
4209 | FATAL_ERROR_ASAP_DISPLAYS_ERROR | Error |  |  | fatal error, going into error state; ASAP will display error | Meaning: A fatal digitizer error occured which will be displayed by PC software component ASAP and not by the digitizer itself. Reason: A fatal digitizer error occured while a PC command was executed; therefore the fatal error code is sent in the response for the PC command and the PC (and not the digitizer) must display the corresponding error message. Cure: Check displayed error message.
4210 | CASSETTE_DETECTION_FAILURE | Error |  |  | cassette is still detected in the input slot at start-up after user confirmed that he removed the cassette (checked 5 s after cassette is released) | Meaning: Failure of the cassette detection or user didn't remove cassette without a tray from the input slot at start-up. Reason: Cassette is still detected in the input slot at start-up after user confirmed that he removed the cassette (checked 5 s after cassette is released). Cure: Check that cassette has been removed from the input slot; if no cassette was in the input slot, check if microswitch on cassette holder is connected and working properly
4211 | AUTO_CLEAN_UP_INIT_CHC_FAILED | Error |  |  | automatic clean-up failed because CHC initialization failed. See preceding error. | Meaning: Automatic clean-up after IP handling problem failed. Reason: Clean-up failed, because CHC (CassetteHandlingControl) detected an error. Cure: Check underlying error description.
4212 | AUTO_CLEAN_UP_INIT_IHC_FAILED | Error |  |  | automatic clean-up failed because IHC initialization failed. See preceding error. | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, because IHC (IP-HandlingControl) detected an error. Cure: Check underlying error description.
4213 | AUTO_CLEAN_UP_NO_IP_FOUND | Error | Open digitizer to check for image plate jam or check for an empty cassette and restart the digitizer. |  | automatic clean-up failed because IP was not found. | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, because IHC (IP-HandlingControl) did not find an IP. Cure: Open digitizer to check for image plate jam or check for an empty cassette and restart the digitizer.
4214 | AUTO_CLEAN_UP_NO_CASS_FOUND | Error |  |  | automatic clean-up failed because cassette was not found. | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, because CHC (CassetteHandlingControl) did not find the cassette, which must be present in this situation. Cure: Check if cassette is clamped correctly; restart digitizer to execute selftest again.
4215 | AUTO_CLEAN_UP_IHC_FAILURE | Error |  |  | automatic clean-up failed because IHC error occured. See preceding error. | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, because IHC (IP-HandlingControl) detected an error. Cure: Check underlying error description.
4216 | AUTO_CLEAN_UP_CHC_SELFTEST_FAILED | Error |  |  | automatic clean-up failed because selftest of CHC failed | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, becasue CHC (CassetteHandlingControl) detected an error during selftest. Cure: Check underlying error description.
4217 | AUTO_CLEAN_UP_IHC_SELFTEST_FAILED | Error |  |  | automatic clean-up failed because selftest of IHC failed | Meaning: Automatic clean-up after handling problem failed. Reason: Clean-up failed, because IHC (IP-HandlingControl) detected an error during selftest. Cure: Check underlying error description.
4218 | FEED_IP_AFTER_AUTO_CLEAN_UP_FAILED | Error |  |  | feeding the IP failed again after automatic clean-up | Meaning: Feeding the IP to the scanner failed again after automatic clean-up. Reason: IHC (IP-HandlingControl) detected an error during feeding the IP, automatic clean-up did not help. Cure: Check underlying error description.
4219 | ICC_DID_NOT_DETECT_BOS | Error |  |  | 'scan_IM_BOS' intermediate response from ICC still missing although slowscan has finished the scan already. | Meaning: The scanning unit didn't start the scan process although the slowscan motor has finished the scan already. Reason: ICC (ImageCaptureControl) is still waiting for begin of the scan, i.e. no pixels have been stored yet in the SDRAM of the FPGA. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI board and PMT/lens assembly; if OK, replace PMI; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly.
4220 | IPSTATUS_NOT_EXPOSED | Error |  |  | IP cannot be scanned because IP status is wrong. | Meaning: The IP cannot be scanned because the IP status is wrong. Reason: Digitizer rejects the scan because the NX did not set the IP status to 'exposed'. Normally the IP status is set to 'exposed' by pressing the 'ID' button on the NX. This is a software error. Cure: Check if digitizer is configured correctly (i.e. as CR30-X) on the NX; check NX software installation; if OK, inform R&D team.
4221 | PMM_5V | Error |  |  | Check PMM 5 failed | Reason: Fuse for PMM 5V on PMI board or power supply defective; Cure: Replace PMM 5V fuse
4222 | CONFIRM_CASSETTE_COVER_BLOCKED | Error | Digitizer cannot open the cassette. The cassette may be damaged. Press confirm to release the cassette. Use another cassette. Call your local service. | Confirm | cassette cover could not have been opened because cover is blocked | Meaning: Digitizer cannot open the cassette because the cassette cover is blocked. Reason: Cassette is damaged, cassette mechanics is defective. Cure: Confirm message; remove cassette; check if cassette cover can be opened by hand; if not, replace cassette
4223 | SCAN_WARNING_SLOWSCAN_SPEEDFACTOR | Error | Image geometry may be out of tolerance. Possible bad image quality. Please contact service. |  | Slowscan finished scan with warning that default slowscan speed factor was used for scanning | Meaning: The scan finished with a warning. Reason: Slowscan reported a warning that the default slowscan speed factor was used for scanning, because the slowscan factor which is stored on the slowscan's ELabel is out of the valid range. Therefore the image geometry may be out of tolerance. Cure: Replace slowscan board with a board which contains the correct speed factor for this device serial number. Ask service headquarter how to do this.
4224 | CONFIRM_AUTO_RETURN_RECONNECTED | Error | Image plate during self-test in the digitizer detected. The image plate was moved back into the cassette. Please check whether the image is present at the PC and re-scan the cassette, if necessary. If the image is present, the cassette was not erased possibly. Erase the cassette again in any case. | Confirm | digitizer received PC command 'cleanUp' after PC was disconnected before scan started; the IP was not scanned and was transported back into the cassette automatically | Meaning: Digitizer received PC command 'cleanUp' after or while IP was returned automatically into the cassette (because PC was disconnected before scan started); the IP was not scanned! Reason: PC disconnected and connected again while digitizer was performing a scan cycle. Cure: Confirm message; remove cassette; insert cassette again to scan the image plate.
4225 | ERASE_MODE_ERASE_IP_FAILED | Error |  |  | erasing the IP failed in the erase mode | Meaning: In the erase mode the erasing of the IP failed. The displayed specific error code of IP-HandlingControl explains the meaning more exactly. Reason: The displayed specific error code of IP-HandlingControl explains the reason. Cure: The displayed specific error code of IP-HandlingControl explains the cure.
4226 | SCAN_MODE_ERASE_IP_FAILED | Error |  |  | erasing the IP failed in the scan mode | Meaning: In the scan mode the erasing of the IP failed. The displayed specific error code of IP-HandlingControl explains the meaning more exactly. Reason: The displayed specific error code of IP-HandlingControl explains the reason. Cure: The displayed specific error code of IP-HandlingControl explains the cure.
4614 | ErrorAttributeIPTypeNotProduced | Error |  |  | ErrorAttribute IPType not produced.
4615 | ErrorAttributeIPSpeedclassNotProduced | Error |  |  | ErrorAttribute IPSpeedclass not produced.
4616 | ICProducing | Error |  |  | InfoCounter Producing failed.
4617 | ErrorProcessing | Error |  |  | ErrorProcessing failed.
4618 | ResponseProcessing | Error |  |  | ResponseProcessing failed.
4619 | Produce | Error |  |  | Produce failed.
4620 | Request | Error |  |  | Request failed.
4621 | RequestList | Error |  |  | RequestList failed.
4622 | RequestCounter | Error |  |  | RequestCounter failed.
4623 | ResetCounter | Error |  |  | ResetCounter failed.
4624 | RemoveEntry | Error |  |  | RemoveEntry failed.
4625 | EmptyList | Error |  |  | EmptyList failed.
4626 | RequestString | Error |  |  | RequestString failed.
4627 | GetIPData | Error |  |  | GetIPData failed.
4628 | ReadingIPData | Error |  |  | ReadingIPData failed.
4629 | IPDataExtracting | Error |  |  | IPDataExtracting failed.
4630 | NoAccessToPcToDic | Error |  |  | No valid pointer to access PcToDic.
4631 | ErrorsPer1000Cycles | Error |  |  | ErrorsPer1000Cycles failed.
4632 | CycleCountFromChassisELabel | Error |  |  | CycleCount could not be read from chassis eLabel.
4633 | AverageCyclesPerDay | Error |  |  | Calculation of Average Cycles Per Day failed.
4634 | CyclesBetweenFailures | Error |  |  | Calculation of CyclesBetweenFailures failed.
4635 | OperationHoursCounter | Error |  |  | Setting OperationHoursCounter failed.
4636 | CalculationOfMaintenanceRequired | Error |  |  | Calculation of MaintenanceRequired failed.
4637 | SerialNumberRegistration | Error | The digitizer is not registered correctly. Serial number mismatch. Please contact service to fix the problem. |  | Serial number not registered. | Meaning: Product registration inconsistent. Reason: Serial number mismatch. Cure: Register serial number via ServiceClient.
4638 | ConnectToServer | Error |  |  | Connect to Server failed.
4639 | NoValidInstallationDate | Error |  |  | No Valid Installation Date found in Infocounter.
4640 | NoValidLastMaintenance | Error |  |  | No Valid Date for last maintenance found in Infocounter.
4641 | MaintenanceRequired | Warning | Maintenance interval expired. Please contact service. | Confirm | Maintenance Interval is expired. Please contact service as soon as possible.
4642 | NodeIDsChanged | Error | Exchanged component detected. Please confirm exchange via ServiceClient. |  | A new module serial number has been detected. Input needed via the service client. | Meaning: At least one module serialnumber has been changed since last startup. Reason: The service technician has exchanged a complete module or a board. Cure: Execute 'Configure -> Exchanged Components' at Service Client.
4643 | CallGSO | Error |  |  | Parameter set of module is corrupted. | Meaning: Parameter set of module is corrupted. Reason: The structure of the eLabel on the module is corrupted. Cure: Check previous error messages to determine which module is corrupted. Call your local Agfa service support in order to get parameter set for corrupted module.
4644 | WrongPMIModuleConfiguration | Error |  |  | Wrong PMI module configuration. | Meaning: Wrong PMI module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace PMI module.
4645 | WrongOpticsModuleConfiguration | Error |  |  | Wrong Optics module configuration. | Meaning: Wrong Optics module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace Optics module.
4646 | WrongPMTModuleConfiguration | Error |  |  | Wrong PMT module configuration. | Meaning: Wrong PMT module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace PMT module.
4647 | WrongHandlingModuleConfiguration | Error |  |  | Wrong Handling module configuration. | Meaning: Wrong Handling module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace Handling module.
4648 | WrongSlowScanModuleConfiguration | Error |  |  | Wrong SlowScan module configuration. | Meaning: Wrong Optics module configuration. Reason: Module serialnumber is 9999 and not allowed for a module. Only sparepart boards have a module serialnumber 9999. Cure: Replace SlowScan module.
4649 | OrderRepairToolForPMI | Error |  |  | Parameter set of PMI board is corrupted. | Meaning: Parameter set of PMI board is corrupted. Reason: The structure of the eLabel on the PMI board is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted PMI board.
4650 | OrderRepairToolForOptics | Error |  |  | Parameter set of optic module is corrupted. | Meaning: Parameter set of optic module is corrupted. Reason: The structure of the eLabel on the optic module is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted optic module.
4651 | OrderRepairToolForPMT | Error |  |  | Parameter set of PMT/light collector module is corrupted. | Meaning: Parameter set of PMT/light collector module is corrupted. Reason: The structure of the eLabel on PMT is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted PMT/light collector module.
4652 | OrderRepairToolForHandling | Error |  |  | Parameter set of handling board is corrupted. | Meaning: Parameter set of handling board is corrupted. Reason: The structure of the eLabel on the handling board is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted handling board.
4653 | OrderRepairToolForSlowScan | Error |  |  | Parameter set of slowscan board is corrupted. | Meaning: Parameter set of slowscan board is corrupted. Reason: The structure of the eLabel on the slowscan board is corrupted. Cure: Call your local Agfa service support in order to get parameter set for corrupted slowscan board.
5120 | INITIALIZE_MODULE | Error |  |  | Testlog sent during initialisation phase
5121 | LOWER_PRE_SCAN_ROLLERS | Error |  |  | The SlowScanLift lowers the pre scan rollers. | Meaning: Reason: Cure:
5122 | WRONG_STATE | Error |  |  | The received event cannot be handled in the actual state of CaSSC. | Meaning: The received event cannot be handled in the actual state of CaSSC. The event and the state can be seen in the event logging. Reason: An event was sent at the wrong time to the CaSSC. Cure: Verify that the event is sent to the correct state of the CaSSC state table.
5123 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE_SLOW_SCAN | Error |  |  | The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScan. | Meaning: The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScan. Reason: A wrong state or has been set. Cure: Check the corresponding SW.
5124 | MOTOR_MOVEMENT_FAILED | Error |  |  | A movement of the SlowScan or the SlowScanLift was not accomplished correctly. The specific movement is coded in the previous error. | Meaning: A movement of the SlowScan or the SlowScanLift was not accomplished correctly. The specific movement is coded in the previous error. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
5125 | HOME_POS_LIFT_FAILED | Error |  |  | The home positioning of the SlowScanLift procedure was not accomplished correctly. | Meaning: The home positioning of the SlowScanLift procedure was not accomplished correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control and the light barrier.
5126 | LIFT_PRE_SCAN_ROLLERS_FAILED | Error |  |  | The pre-scan rollers were not lifted correctly. | Meaning: The pre-scan rollers were not lifted correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control
5127 | PRE_LIFT_PRE_SCAN_ROLLERS_FAILED | Error |  |  | The pre-scan rollers were not pre-lifted correctly. | Meaning: The pre-scan rollers were not pre-lifted correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control
5128 | LOWER_PRE_SCAN_ROLLERS_FAILED | Error |  |  | The pre-scan rollers were not lowered correctly. | Meaning: The pre-scan rollers were not lowered correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control
5129 | LOWER_PRE_SCAN_ROLLERS_FOR_ALIGNMENT_FAILED | Error |  |  | The pre-scan rollers were not lowered correctly for the alignment of the IP. | Meaning: The pre-scan rollers were not lowered correctly for the alignment of the IP. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5130 | LIFT_POST_SCAN_ROLLERS_FAILED | Error |  |  | The post-scan rollers were not lifted correctly. | Meaning: The post-scan rollers were not lifted correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5131 | LOWER_POST_SCAN_ROLLERS_FAILED | Error |  |  | The post-scan rollers were not lowered correctly. | Meaning: The post-scan rollers were not lowered correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5132 | GO_STEPS_FAILED | Error |  |  | The selected motor was not moved a defined number of steps correctly. | Meaning: The selected motor was not moved a defined number of steps correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5133 | IP_NOT_AT_PRE_SCAN_ROLLERS | Error |  |  | The IP was not transported to the pre-scan rollers correctly. | Meaning: The IP was not transported to the pre-scan rollers correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5134 | IP_NOT_AT_POST_SCAN_ROLLERS | Error |  |  | The IP was not transported to the post-scan rollers correctly. | Meaning: The IP was not transported to the post-scan rollers correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5135 | REAR_PART_OF_IP_NOT_AT_POST_SCAN_ROLLERS | Error |  |  | The rear part of the IP did not reach the post-scan rollers correctly. | Meaning: The rear part of the IP did not reach the post-scan rollers correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5136 | IP_NOT_AT_EOS | Error |  |  | The IP was not transported to the end of scan correctly. | Meaning: The IP was not transported to the end of scan correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5137 | IP_NOT_AT_BOE | Error |  |  | The IP was not transported to the begin of erasure correctly. | Meaning: The IP was not transported to the begin of erasure correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
5138 | WRONG_COMMAND | Error |  |  | The received command is unknown and cannot be handled by the CaSSC. | Meaning: The received command is unknown and cannot be handled by the CaSSC. Reason: A wrong event was sent to the CaSSC. Cure: Verify that the corrrect event is sent to the CaSSC.
5139 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE_SLOW_SCAN_LIFT | Error |  |  | The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScanLift. | Meaning: The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScanLift. Reason: A wrong state or has been set. Cure: Check the corresponding SW.
5140 | CALL_BACK_POINTER_ZERO | Error |  |  | The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. | Meaning: The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. Reason: -- Cure: Check the (initialization of) pointers.
5637 | LOCKER_ABS_POS_FAILED | Error |  |  | Locker Absolute Pos Movement terminated with failure.Obsolete | Meaning: Function disturbed. Homepos/Absolutepos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
5638 | LOCKER_HOME_POS_MAX_STEP_LIMIT | Error |  |  | Locker Home Pos Movement terminated with failure. | Meaning: Function disturbed. Homepos/Absolutepos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
5639 | LOCKER_LOCK_POS_MAX_STEP_LIMIT | Error |  |  | Locker Lock Pos Movement terminated with failure. | Meaning: Function disturbed. Lockpos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
5640 | OPENER_ABS_POS_FAILED | Error |  |  | Opener Absolute Pos Movement terminated with failure. | Meaning: Function disturbed. Homepos/Absolutepos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
5641 | OPENER_HOME_POS_MAX_STEP_LIMIT | Error |  |  | Opener Abs Pos Movement terminated with failure. | Meaning: Function disturbed. Homepos/Absolutepos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
5642 | OPENER_OPEN_POS_MAX_STEP_LIMIT | Error |  |  | Opener Open Pos Movement terminated with failure. | Meaning: Function disturbed. Openpos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
5643 | CASS_DETECTION_LOOK_FOR_CASS_DETECTION_FAILED | Error |  |  | Procedure lookForCassDetectionStatic terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Cassette Edge Sensor has an invalid value. Cure: Retry Cassette Detection.
5644 | FIXER_SELFTEST_FAILED | Error |  |  | Procedure selftest of SM-Fixer terminated with failure. Obsolete | Meaning: Retry possible, function disturbed. Reason: The Cassette Handling Motor SM-Fixer terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
5645 | INIT_STEPPERS_RETRY_MAX_LIMIT | Error |  |  | Init with Abs Pos Move, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. Check all 3 drives and corresponding mechanics: cassette opening, clamping and locking / unlocking
5646 | RETRY_MAX_LIMIT_REACHED | Error |  |  | Feeding Moves, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Feeding Cassette incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
5647 | CLOSE_COVER_RETRY_MAX_LIMIT | Error |  |  | Close Cover Move, Retries reach Max Limit. | Meaning: The cassette shutter could not be closed correctly. The CASSETTE OPENER MOTOR M3 did not reached the correct target position for cassette closed. The cassette is still clamped and could not be released. The Digitizer is out of operation and requires service support. Reason: The CASSETTE OPENER MOTOR M3 could not close the cassette shutter successfully due to: - Tray or Image Plate is still outside of the cassette and blocks the shutter - Image Plate is damaged or bended and blocks the shutter - Drawer unit is not properly adjusted and blocks the shutter - Shutter of the cassette is blocked of defect - Load detection for CASSETTE OPENER MOTOR M3 is imperfect, STALL CALIBRATION was not successful - Drive mechanism of CASSETTE OPENER MOTOR M3 is blocked or defect Cure: 1. Restart the Digitizer and check, if the problem persists 2. If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual 3. Check the tray and the Image Plate for any defects 4. Check the function of the cassette shutter and the cassette inside for any particle or removed fleece 5. Check the proper position of the DRAWER UNIT, refer to the Service Manual, adjustments 6. Check the cabling of the DRAWER MOTOR M4, refer to the Service Manual, troubleshooting guide 7. Check the drive mechanism of CASSETTE OPENER MOTOR M3 for proper function and smooth running 8. Perform a calibration of the CASSETTE OPENER MOTOR M3, refer to the Service Manual 9. Exchange the CASSETTE OPENER UNIT, refer to the Service Manual for order and replacement 10 Check the drive mechanism of CASSETTE FIXATION UNIT for proper function and smooth running
5648 | OPEN_COVER_RETRY_MAX_LIMIT | Error |  |  | Open Cover Move, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Open Cassette incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. Check mechanics of cassette handling module, opener.
5649 | FIXER_UNCLAMP_CASS_RETRY_MAX_LIMIT | Error |  |  | Unclamp Cassette Move, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Unclamp Cassette incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. Check mechanics of cassette clamping.
5650 | CLEAN_UP_FAILED | Error |  |  | Uncovered failure during cleanUp | Meaning: The Image Plate with the tray could no be returned to the cassette. The cassette is still clamped and could not be released. The Digitizer is out of operation and requires service support. Reason: The Digitizer could not return the Image Plate back to the cassette due to: - Image plate or tray is damaged or bended and blocks the shutter - Particles inside the cassette or removed cassette fleece - Shutter of the cassette is blocked or defect - Drive mechanism of the DRAWER UNIT is stiff or blocked - Load detection for DRAWER MOTOR M4 is imperfect, STALL CALIBRATION was not successfull - DRAWER UNIT is not properly adjusted and blocks the shutter - Drive mechanism in the CASSETTE FIXATION UNIT is stiff or blocked - Load detection for CASSETTE FIXING MOTOR M1 is imperfect, STALL CALIBRATION was not successfull - Drive mechanism of CASSETTE OPENER MOTOR M3 is blocked or defect - Load detection for CASSETTE OPENER MOTOR M3 is imperfect, STALL CALIBRATION was not successfull - Drive mechanism of CASSETTE UNLOCKING UNIT is blocked or defect - CASSETTE EDGE SENSOR ist not properly alligned or defect Cure: 1. Restart the Digitizer and check, if the problem persists 2. If the cassette is still clamped, remove the tray with Image Plate and the cassette manually, refer to the Service Manual 3. Check the tray and the Image Plate for any defects 4. Check the function of the cassette shutter and the cassette inside for any particle or removed fleece 5. Check the proper position of the DRAWER UNIT, refer to the Service Manual, adjustments 6. Check the cabling of the DRAWER MOTOR M4, refer to the Service Manual, troubleshooting guide 7. Check the drive mechanism of DRAWER UNIT for proper function and smooth running 8. Check the drive mechanism of CASSETTE FIXATION UNIT for proper function and smooth running 9. Check the drive mechanism of CASSETTE OPENER MOTOR M3 for proper function and smooth running 10. Check the drive mechanism of CASSETTE UNLOCKING UNIT for proper function and smooth running 11. Check the CASSETTE EDGE SENSOR for proper function 12. Perform a calibration of the DRAWER MOTOR M4, the CASSETTE OPENER MOTOR M3, the ROLLER MOTOR M5 and the CASSETTE FIXING MOTOR M1, refer to the Service Manual
5651 | READY_FOR_INIT_MODULE_SM_MOVE_FAILED | Error |  |  | Uncovered failure during cleanUp, readyForInit actions | Meaning: Function disturbed. Homepos/Absolutepos not found. Clean Up incomplete. Reason: Movent Failed, steps limit reached, or mechanical stop not found. Cure: Restart Digitizer. Check IP transport path for any object.
5652 | OPENER_ABS_POS_INIT_MODULE_RETRY_MAX_LIMIT | Error |  |  | OpenerAbsPos Move in initModule, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Initialisation of Stepper Motor Opener in Cassette Handling incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service. 1. Check Opener Mechanics for obstacles, Check the driving elements for getting stuck or blocked, Ensure easy smooth movement. 2. Check Repair Part 'TOOTHED BELT FOR CASSETTE UNIT' Related Repair Parts are also 'CASSETTE OPENER' and 'GS CR30 IP-HANDLING'
5653 | LOCKER_LOCK_IP_RETRY_MAX_LIMIT | Error |  |  | Lock IP Move failed repeatedly, Retries reach Max Limit. | Meaning: Function disturbed. AbsolutePos not reached. Lock IP incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found or way blocked. Cure: Restart Digitizer. If error persists please call Service.
5654 | LOCKER_UNLOCK_IP_RETRY_MAX_LIMIT | Error |  |  | UnLock IP Move failed repeatedly, Retries reach Max Limit. | Meaning: Function disturbed. AbsolutePos not reached. UnLock IP incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found or way blocked. Cure: Restart Digitizer. If error persists please call Service.
5655 | CHECK_FIXER_RETRY_MAX_LIMIT | Error |  |  | Fixer Selftest, Retries reach Max Limit. | Meaning: Function disturbed. Selftest at motor Fixer for clamp cassette incomplete. Reason: Maximum Retry limit reached. Selftest Fixer failed. Cure: Restart Digitizer. If error persists please call Service. 1. Check cables to motor Fixer. 2. Change Repair Part 'CASSETTE HOLDER'. 3. Change Repair Part 'CASSETTE HOLDER' together with 'GS CR30 IP-HANDLING' (mounted on the 'CASSETTE HOLDER'). Refer to the procedure for the exchange in the service documentation.
5656 | CHECK_OPENER_RETRY_MAX_LIMIT | Error |  |  | Opener Selftest, Retries reach Max Limit. | Meaning: Function disturbed. Selftest at motor Opener for open cassette cover incomplete. Reason: Maximum Retry limit reached. Selftest Opener failed. Cure: Restart Digitizer. If error persists please call Service. 1. Check cables to motor Opener. 2. Change Repair Part 'CASSETTE OPENER'. 3. Change Repair Part 'CASSETTE OPENER' and 'GS CR30 IP-HANDLING'. Refer to the procedure for the exchange in the service documentation.
5657 | CHECK_LOCKER_RETRY_MAX_LIMIT | Error |  |  | Locker Selftest, Retries reach Max Limit. | Meaning: Function disturbed. Selftest at motor Locker for locking the IP in cassette incomplete. Reason: Maximum Retry limit reached. Selftest Locker failed. Cure: Restart Digitizer. If error persists please call Service. 1. Check cables to motor Locker. 2. Change Repair Part 'CASSETTE RELEASE'. 3. Change Repair Part 'CASSETTE RELEASE' and 'GS CR30 IP-HANDLING'. Refer to the procedure for the exchange in the service documentation.
5658 | SELFTEST_RETRY_INVALID_PAR | Error |  |  | Selftest, Retries start with invalid parameter | Meaning: Function disturbed. Selftest incomplete. Reason: Retry at Selftest Procedure starts with invalid parameter. Cure: Restart Digitizer. If error persists please call Service.
5659 | INIT_MODULE_ABS_POS_RETRY_MAX_LIMIT | Error |  |  | AbsPos Move in initModule, Retries reach Max Limit. | Meaning: Function disturbed. Homepos/Absolutepos not found. Initialisation incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Unspecific failure for motor Fixer, Locker, Opener. Cure: Restart Digitizer. If error persists please call Service. 1. Check cables of motor Fixer, Locker, Opener.
5660 | LOCKER_INIT_MODULE_ABS_POS_RETRY_MAX_LIMIT | Error |  |  | Locker AbsPos Move in initModule, Retries reach Max Limit. | Meaning: Function disturbed. Locker Homepos/Absolutepos not found. Initialisation incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service. 1. Check Locker Mechanics for obstacles, Check the driving elements for getting stuck or blocked, Ensure easy smooth movement. Related Repair Parts are 'CASSETTE RELEASE' and 'GS CR30 IP-HANDLING'
5661 | FIXER_INIT_MODULE_ABS_POS_RETRY_MAX_LIMIT | Error |  |  | Fixer AbsPos Move in initModule, Retries reach Max Limit. | Meaning: Function disturbed. Fixer Homepos/Absolutepos not found. Initialisation incomplete. Reason: Maximum Retry limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service. 1. Check Fixer Mechanics for obstacles, Check the driving elements for getting stuck or blocked, Ensure easy smooth movement. Related Repair Parts are 'CASSETTE HOLDER' and 'GS CR30 IP-HANDLING'
5732 | WARN_CLAMP_CASS_INVALID_WIDTH | Warning | Unknown cassette format detected (cassette inserted correctly?). Remove cassette and insert it again. |  | Warning: Clamp Cassette Movement detected invalid Cassette Width | Meaning: Retry possible, function disturbed. Reason: The clamped Cassette has a invalid width. Cure: Re-clamp Cassette. If error persists please calibrate mechanical Cassette Unit.
5733 | WARN_LOCKER_HOME_MOVE_STOP_UNEXPECTED | Error |  |  | Warning: Locker home movement stopped early. Parameter is steps moved. | Meaning: Retry possible, function may be disturbed. Reason: The Locker home movement stopped early as stall reacted in position range before stopper. Cure: Restart Digitizer. If error persists please call Service.
5734 | WARN_SM_BIF_ERR_SIGNAL | Error |  |  | Warning: CHC-SM-BIF returns error signal. | Meaning: Function disturbed. Reason: Stepper Motor Interface Class sends Signal for Error Cure: Restart Digitizer. If error persists please call Service.
6144 | INITIALIZE_MODULE | Error |  |  | Testlog sent during initialisation phase.
6145 | WRONG_STATE | Error |  |  | The received event cannot be handled in the actual state of CaIPH. | Meaning: The received event cannot be handled in the actual state of CaSSC. The event and the state can be seen in the event logging. Reason: An event was sent at the wrong time to the CaIHC. Cure: Verify that the event is sent to the correct state of the CaIHC state table.
6146 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE | Error |  |  | The received event cannot be handled in the actual state of a low level state machine (in a passive class). | Meaning: The received event cannot be handled in the actual state of a low level state machine (in a passive class). Reason: A wrong state has been set. Cure: Check the corresponding SW.
6147 | MOVE_DRAWER_OUT_OF_CASSETTE_FAILED | Error |  |  | The drawer was not moved out of the cassette correctly. | Meaning: The drawer was not moved out of the cassette correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6148 | MOVE_DRAWER_OUT_OF_CASSETTE_FOR_INIT_FAILED | Error |  |  | The drawer was not moved out of the cassette correctly during the initialization procedure. | Meaning: The drawer was not moved out of the cassette correctly during the initialization procedure. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6149 | MOVE_DRAWER_INTO_CASSETTE_FAILED | Error |  |  | The drawer was not moved into the cassette correctly. | Meaning: The drawer was not moved into the cassette correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6150 | MOVE_DRAWER_INTO_CASSETTE_FOR_INIT_FAILED | Error |  |  | The drawer was not moved into the cassette correctly during the initialization procedure. | Meaning: The drawer was not moved into the cassette correctly during the initialization/calibration procedure Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6151 | PREPARE_MOVING_IP_TO_SCANNER_FAILED | Error |  |  | The preparation for moving the IP to the scanner was not accomplished correctly. | Meaning: The preparation for moving the IP to the scanner was not accomplished correctly. Reason: The movement or parts of it were not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6152 | MOVE_IP_TO_SCANNER_FAILED | Error |  |  | Moving the IP to the scanner was not accomplished correctly. | Meaning: Moving the IP to the scanner was not accomplished correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6153 | MOVE_IP_TO_ALIGNMENT_POS_FAILED | Error |  |  | Moving the IP to the alignment position was not accomplished correctly. | Meaning: Moving the IP to the alignment position was not accomplished correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6154 | MOVE_IP_TO_TRAY_FAILED | Error |  |  | Moving the IP to the tray was not accomplished correctly. | Meaning: Moving the IP to the tray was not accomplished correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6155 | GO_STEPS_FAILED | Error |  |  | The selected motor was not moved a defined number of steps correctly. | Meaning: The selected motor was not moved a defined number of steps correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6156 | CLAMP_IP_FAILED | Error |  |  | The IP was not clamped correctly. | Meaning: The IP was not clamped correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6157 | FREE_IP_FAILED | Error |  |  | The IP was not unclamped correctly. | Meaning: The IP was not unclamped correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6158 | MOVE_TO_ALIGNMENT_POS_FAILED | Error |  |  | The IP was not moved to the alignment position correctly. | Meaning: The IP was not moved to the alignment position correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6159 | MOTOR_MOVEMENT_FAILED | Error |  |  | A movement of the Drawer, the Roller or the RollerClamper was not accomplished correctly. The specific movement is coded in the previous error. | Meaning: A movement of the Drawer, the Roller or the RollerClamper was not accomplished correctly. The specific movement is coded in the previous error. Reason: The movement was not controlled correctly. Cure: Check mechnics transport roller unclamping/clamping; check mechanics drawer movement; check cassette for damages
6160 | IP_OUT_OF_TRAY_FAILED | Error |  |  | The IP could not be taken out of the cassette/tray correctly. | Meaning: The IP could not be taken out of the cassette/tray correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6161 | RETURN_IP_TO_TRAY_FAILED | Error |  |  | The IP could not be returned to the tray correctly. | Meaning: The IP could not be returned to the tray correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6162 | RETURN_IP_FROM_SLOW_SCAN_UNIT_FAILED | Error | Image plate jam. Please restart the digitizer and erase the image plate. |  | The IP could not be returned from the SlowScanUnit correctly. | Meaning: The IP could not be returned from the SlowScanUnit correctly. Reason: The movement was not controlled correctly or there is a problem with the mechanics Cure: Check the corresponding movement control and mechanics.
6163 | MOVE_DRAWER_AFTER_HOME_POS_REACHED_FAILED | Error |  |  | The drawer was not moved correctly after having reached the home position. | Meaning: The drawer was not moved correctly after having reached the home position. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6164 | MOVE_ROLLER_CLAMPING_AFTER_HOME_POS_REACHED_FAILED | Error |  |  | The RollerClamping motor was not moved correctly after having reached the home position. | Meaning: The RollerClamping motor was not moved correctly after having reached the home position. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters).
6165 | CABLE_ERASURE_UNIT_NOT_CONNECTED | Error | Failure at erasure module. Please restart digitizer. If error remains, please contact service. |  | The data cable of the erasure unit is not connected. | Meaning: The data cable of the erasure unit is not connected. Reason: -- Cure: Check the cable connection. Replace the IP handling erasure cable and/or the erasure unit if necessary.
6166 | TEMPERATURE_OF_ERASURE_UNIT_TOO_HIGH | Error | The temperature of the erasure unit is too high. Please wait until the cooling unit reduces the temperature. |  | The temperature of the erasure unit is too high. | Meaning: The temperature of the erasure unit is too high. Reason: -- Cure: Wait until the cooling unit reduces the temperature.
6167 | WRONG_SIGNAL_OBTAINED_FROM_TRANSPORT_LIGHT_BARRIER | Error |  |  | The signal obtained from the light barrier is wrong (i.e. end of IP detected instead of the beginning of the IP). | Meaning: The signal obtained from the light barrier is wrong (i.e. end of IP detected instaed of the beginning of the IP). Reason: -- Cure: Check the corresponding hardware and software. Replace the reflective light sensor if necessary.
6168 | FEED_IP_FAILED | Error | Open digitizer to check for image plate jam or check for an empty cassette and restart the digitizer. |  | Feed the IP failed, i.e. take the IP out of the cassette and transport it to the scanner. | Meaning: Feed the IP failed, i.e. take the IP out of the cassette and transport it to the scanner. Reason: The movement was not controlled correctly. Cure: Check for presence of IP; check drawer movement, transport roller clamping, obstacles in IP transport path. Perform stall calibration
6169 | END_OF_IP_NOT_DETECTED | Error |  |  | The end of the IP was not detected in the transport light barrier while transporting the IP back into the cassette. | Meaning: The end of the IP was not detected in the transport light barrier while transporting the IP back into the cassette. Reason: -- Cure: Check the hardware and software for detecting the IP.
6170 | WRONG_STATE_LOW_LEVEL_STATE | Error |  |  | The received event cannot be handled in the actual state of the low level state machine (passive class). | The received event cannot be handled in the actual state of the low level state machine of the Drawer (passive class). Reason: A wrong state has been set. Cure: Check the corresponding SW.
6171 | LED_CIRCUIT_ERASURE_UNIT_NOT_OPERATIONAL | Error | Failure at erasure module. Please restart digitizer. If error remains, please contact service. If an image has been scanned, the image plate must be erased again after the repair of the digitizer. |  | One of the LED circuits of the erasure unit is not operational. | Meaning: One of the LED circuits of the erasure unit is not operational, i.e. the corresponding self test failed. The circuit number is given as parameter in the error message. Reason: -- Cure: Check the voltage supply connection. Check the LED circuit. Replace the erasure unit and or the IP handling erasure cable if necessary.
6172 | WRONG_SIGNAL_LED_CIRCUIT_ERASURE_UNIT | Error |  |  | A wrong signal was obtained from one of the LED circuits when the the erasure unit was disabled. | Meaning: A wrong signal was obtained from one of the LED circuits, i.e.the erasure unit was disabled and the corresponding self test failed. The circuit number is given as parameter in the error message. Reason: The erasure unit may not have been switched off. Cure: Check the LED circuit. Replace the IP handling erasure cable and/or the erasure unit if necessary.
6173 | WRONG_VALUE_STATE_ERASURE_UNIT | Error |  |  | The erasure unit state passed to a SW operation has a wrong value. | Meaning: The erasure unit state passed to a SW operation has a wrong value (i.e. neiter ('enabled' nor 'disabled'). Reason: -- Cure: Check the corresponding SW operations. Replace the erasure unit and or the IP Handling erasure cable if necessary.
6175 | ERASURE_UNIT_NOT_SWITCHED_ON_CORRECTLY | Error | Failure at erasure module. Please restart digitizer. If error remains, please contact service. If an image has been scanned, the image plate must be erased again after the repair of the digitizer. |  | The erasure unit was not switched on correctly. | Meaning: The erasure unit was not switched on correctly. Reason: -- Cure: Check the corresponding hardware and software. Replace the erasure unit and or the IP Handling erasure cable if necessary.
6176 | ERASURE_UNIT_NOT_SWITCHED_OFF_CORRECTLY | Error | Failure at erasure module. Please restart digitizer. If error remains, please contact service. |  | The erasure unit was not switched off correctly. | Meaning: The erasure unit was not switched off correctly. Reason: -- Cure: Check the corresponding hardware and software. Replace the erasure unit and or the IP Handling erasure cable if necessary.
6177 | SELF_TEST_FAILED | Error |  |  | The IP handling self test failed. | Meaning: The IP handling self test failed. Reason: See the more detailed error loggings. Cure: See the more detailed error loggings.
6178 | WRONG_COMMAND | Error |  |  | The received command is unknown and cannot be handled by the CaIHC. | Meaning: The received command is unknown and cannot be handled by the CaIHC. Reason: A wrong event was sent to the CaIHC. Cure: Verify that the corrrect event is sent to the CaIHC.
6179 | WRONG_LED_CIRCUIT_NUMBER | Error |  |  | The received number of the (erasure unit) LED circuit is not within the required range of values (1 - 6). | Meaning: The received number of the (erasure unit) LED circuit is not within the required range of values (1 - 6). Reason: -- Cure: Verify that the corrrect LED circuit number is sent by the SW.
6180 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE_DRAWER | Error |  |  | The received event cannot be handled in the actual state of the low level state machine of the passive class Drawer. | Meaning: The received event cannot be handled in the actual state of the low level state machine of the passive class Drawer. Reason: A wrong state or has been set. Cure: Check the corresponding SW.
6181 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE_ROLLER | Error |  |  | The received event cannot be handled in the actual state of the low level state machine of the passive class Roller. | Meaning: The received event cannot be handled in the actual state of the low level state machine of the passive class Roller. Reason: A wrong state or has been set. Cure: Check the corresponding SW.
6182 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE_ROLLER_CLAMPING | Error |  |  | The received event cannot be handled in the actual state of the low level state machine of the passive class RollerClamping. | Meaning: The received event cannot be handled in the actual state of the low level state machine of the passive class RollerClamping. Reason: A wrong state or has been set. Cure: Check the corresponding SW.
6183 | MOVE_DRAWER_INTO_CASSETTE_FOR_CALIBRATION_FAILED | Error |  |  | The drawer was not moved into the cassette correctly during the calibration procedure. | Meaning: The calibration of the DRAWER MOTOR M4 failed. The DRAWER UNIT can not be moved correctly into the cassette during the calibration procedure. The Digitizer is out of operation and requires service support. Reason: The calibration of the DRAWER MOTOR M4 failed due to: - The travel path of the DRAWER UNIT is obstructed (e.g. screw blocks the guide rail) - Drive mechanism of the DRAWER UNIT is stiff or blocked - The transport rollers are clamped - The cassette entry slot is closed - DRAWER UNIT is not properly adjusted - Defective cable, broken wires or bad contacts - Defective DRAWER MOTOR M4 - Electronic or electrical error of the IP HANDLING BOARD Cure: 1. Check, if the travel path of the DRAWER UNIT is not obstructed 2. Check the drive mechanism of DRAWER UNIT for proper function and smooth running 3. If the DRAWER UNIT was replaced: Check, that the DRAWER UNIT is placed right angular on both toothed tracks 4. Check the complete drive mechanism of MOTOR ROLLER CLAMPING M6 for proper function and smooth running 5. Check the complete drive mechanism of MOTOR CASSETTE FIXING M1 for proper function and smooth running 6. Check the proper position of the DRAWER UNIT, refer to the Service Manual, adjustments 7. Check the cabling of the DRAWER MOTOR M4, refer to the Service Manual, troubleshooting guide 8. Perform a calibration of the DRAWER MOTOR M4 again 9. Exchange the DRAWER MOTOR M4 and the IP HANDLING BOARD, refer to the Service Manual for order and replacement
6184 | MOVE_DRAWER_OUT_OF_CASSETTE_FOR_CALIBRATION_FAILED | Error |  |  | The drawer was not moved out of the cassette correctly during the calibration procedure. | Meaning: The drawer was not moved out of the cassette correctly during the calibration procedure. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters) and mecjanics (this implies e.g. the Cassette Opener).
6185 | CALL_BACK_POINTER_ZERO | Error |  |  | The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. | Meaning: The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. Reason: -- Cure: Check the (initialization of) pointers.
6186 | STOP_ROLLER_FAILED | Error |  |  | The Roller did not stop correctly. | Meaning: The Roller did not stop correctly. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control.
6187 | WRONG_VALUE_IRUN_ADJUSTABLE_DRAWER | Error |  |  | The adjustable run current value of the Drawer read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable run current value of the Drawer read out of the eLabel is not within the required range. Therefore the default value is used. The wrong value is given as parameter. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Check the eLabel access (SW and HW).
6188 | WRONG_VALUE_IRUN_ADJUSTABLE_ROLLER_CLAMPING | Error |  |  | The adjustable run current value of the RollerClamping read out of the eLabel is not within the required range. Therefore the default value is used. | Meaning: The adjustable run current value of the RollerClamping read out of the eLabel is not within the required range. Therefore the default value is used. The wrong value is given as parameter. Reason: The value was not stored correctly in the eLabel or it was overwritten. Cure: Check the eLabel access (SW and HW).
6189 | WRONG_CHC_PROTVER | Error |  |  | Protocol version of CHC is wrong | Meaning: Incompatible SW components (protocol version of CassetteHandlingControl CHC is wrong). Reason: SW error. Cure: Install new SW version on all boards; inform R&D.
6190 | WRONG_SSC_PROTVER | Error |  |  | Protocol version of SSC is wrong | Meaning: Incompatible SW components (protocol version of SlowscanControl SSC is wrong). Reason: Different SW versions are installed on IP-Handling board and on Slowscan board. Cure: Install same SW version on all boards.
6191 | CAN_ADAPTER_ERROR | Error |  |  | fatal error in CAN adapter layer | Meaning: Fatal error in CAN communication layer. Reason: CAN error or SW Error Cure: Check CAN connections; if OK, install new SW and inform R&D.
6192 | LOGGING_ADAPTER_ERROR | Error |  |  | fatal error in logging adapter layer | Meaning: Fatal error in logger layer. Reason: CAN error or SW Error Cure: Check CAN connections; if OK, install new SW and inform R&D.
6193 | NUM_RETRIES_DRAWER_INTO_CASSETTE_EXCEEDED | Error |  |  | The number of retries for moving the Drawer into the cassette was exceeded. | Meaning: The number of retries for moving the Drawer into the cassette was exceeded. Reason: The number of steps required for a complete movement cannot be accomplished. Cure: Check the corresponding drawer mechanics. Check drawer height adjustment.
6194 | NUM_RETRIES_DRAWER_OUT_OF_CASSETTE_EXCEEDED | Error |  |  | The number of retries for moving the Drawer out of the cassette was exceeded. | Meaning: The number of retries for moving the Drawer out of the cassette was exceeded. Reason: The number of steps required for a complete movement cannot be accomplished. Cure: Check the corresponding drawer mecha nics.
6195 | NUM_RETRIES_CLAMP_IP_EXCEEDED | Error |  |  | The number of retries for clamping the IP (RollerClamper) was exceeded. | Meaning: The number of retries for clamping the IP was exceeded. Reason: The number of steps required for a complete movement cannot be accomplished. Cure: Check the corresponding clamping mecha nics.
6196 | NUM_RETRIES_UNCLAMP_IP_EXCEEDED | Error |  |  | The number of retries for unclamping the IP (RollerClamper) was exceeded. | Meaning: The number of retries for unclamping the IP was exceeded. Reason: The number of steps required for a complete movement cannot be accomplished. Cure: Check the corresponding clamping mecha nics.
6197 | CLAMP_IP_FOR_CALIBRATION_FAILED | Error |  |  | The transport rollers were not clamped correctly during the calibration procedure. | Meaning:The transport rollers were not clamped correctly during the calibration procedure. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters) and mechanics.
6198 | UNCLAMP_IP_FOR_CALIBRATION_FAILED | Error |  |  | The transport rollers were not unclamped correctly during the calibration procedure. | Meaning:The transport rollers were not unclamped correctly during the calibration procedure. Reason: The movement was not controlled correctly. Cure: Check the corresponding movement control (SW and stall parameters) and mechanics.
6199 | MOVE_IP_INTO_SLOW_SCAN_UNIT_FAILED | Error |  |  | The IP could not be moved into the SlowScanUnit correctly. | Meaning: The ROLLER IP SENSOR has not been released, while the image plate was already scanned. The cassette is still clamped and could not be released. The Digitizer is out of operation and requires service support. Reason: The ROLLER IP SENSOR has not been released after scan of the image plate due to: - The image plate is damaged or bended and can not enter the pre-scan rollers - The pre-scan rollers are not properly positioned in the SLOWSCAN or can not be clamped - The drive mechanism of the SLOWSCAN is stiff or blocked - Defective cable, broken wires or bad contacts of the ROLLER IP SENSOR - Defective ROLLER IP SENSOR - Electronic or electrical error of the IP HANDLING BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. If the cassette is still clamped, open the CASSETTE UNIT and check the position of the image plate in the Digitizer: 2a. If the image plate is in the detection range of ROLLER IP SENSOR: - Check if the image plate is damaged or bended - Check if the image plate is clamped between the pre-scan rollers - Check if the pre-scan rollers are properly positioned in the SLOWSCAN - Check the drive mechanism of the SLOWSCAN for proper function and smooth running - Replace the SLOWSCAN unit, refer to the Service Manual for order and replacement 2b. If the image plate is out of the detection range of ROLLER IP SENSOR: - Check the proper position of the ROLLER IP SENSOR - Check the cabling and connection between ROLLER IP SENSOR and IP HANDLING BOARD - Replace the ROLLER IP SENSOR, refer to the Service Manual for order and replacement - Replace the IP HANDLING BOARD, refer to the Service Manual for order and replacement
6200 | LIGHT_BARRIER_FOR_IP_DETECTION_NOT_CONNECTED | Error |  |  | The light barrier for detecting the IP is not connected correctly. | Meaning: The light barrier for detecting the IP is not connected correctly. Reason: -- Cure: Check the light barrier and the corresponding cabling. Replace the light barrier if necessary.
6201 | SLOW_SCAN_SPEED_NOT_SET_TO_ERASURE_SPEED_CORRECTLY | Error | Failure at erasure module. Please restart digitizer. If error remains, please contact service. If an image has been scanned, the image plate must be erased again after the repair of the digitizer. |  | The IP was not erased correctly. Please erase the IP again using the separate erasure cycle. | Meaning: The IP was not erased correctly because the speed of the SlowScanUnit was not set correctly to the erasure speed. Reason: Communication error or error in the corresponding software routine. Cure: Check the corresponding software routines.
6202 | IP_NOT_TRANSPORTED_THROUGH_SLOW_SCAN_UNIT_DURING_SCAN | Error |  |  | The IP was not transported correctly through the SlowScan unit during the scan. | Meaning: The ROLLER IP SENSOR has not been released, while the image plate was already scanned. The cassette is still clamped and could not be released. The Digitizer is out of operation and requires service support. Reason: The ROLLER IP SENSOR has not been released after scan of the image plate due to: - The image plate is damaged or bended and can not enter the pre-scan rollers - The pre-scan rollers are not properly positioned in the SLOWSCAN or can not be clamped - The drive mechanism of the SLOWSCAN is stiff or blocked - Defective cable, broken wires or bad contacts of the ROLLER IP SENSOR - Defective ROLLER IP SENSOR - Electronic or electrical error of the IP HANDLING BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. If the cassette is still clamped, open the CASSETTE UNIT and check the position of the image plate in the Digitizer: 2a. If the image plate is in the detection range of ROLLER IP SENSOR: - Check if the image plate is damaged or bended - Check if the image plate is clamped between the pre-scan rollers - Check if the pre-scan rollers are proper positioned in the SLOWSCAN - Check the drive mechanism of the SLOWSCAN for proper function and smooth running - Replace the SLOWSCAN unit, refer to the Service Manual for order and replacement 2b. If the image plate is out of the detection range of ROLLER IP SENSOR: - Check the proper position of the ROLLER IP SENSOR - Check the cabling and connection between ROLLER IP SENSOR and IP HANDLING BOARD - Replace the ROLLER IP SENSOR, refer to the Service Manual for order and replacement - Replace the IP HANDLING BOARD, refer to the Service Manual for order and replacement r function and smooth running - Replace the SLOWSCAN unit, refer to the Service Manual for order and replacement 2b. If the image plate is out of the detection range of ROLLER IP SENSOR: - Check the proper position of the ROLLER IP SENSOR - Check the cabling and connection between ROLLER IP SENSOR and IP HANDLING BOARD - Replace the ROLLER IP SENSOR, refer to the Service Manual for order and replacement - Replace the IP HANDLING BOARD, refer to the Service Manual for order and replacement
6203 | INITIALIZE_ROLLER_CLAMPING_FAILED | Error |  |  | The initialization of the RollerClamping Motor failed. | Meaning: The initialization of the MOTOR ROLLER CLAMPING M6 failed. The MOTOR ROLLER CLAMPING M6 did not detect the home position during the initialization. The Digitizer is out of operation and requires service support. Reason: The initialization of the MOTOR ROLLER CLAMPING M6 failed due to: - Toothed belt of the MOTOR ROLLER CLAMPING M6 slipped off or is defective - Drive mechanism of the MOTOR ROLLER CLAMPING M6 is blocked or defect - Detection of home position for MOTOR ROLLER CLAMPING M6 is imperfect, STALL CALIBRATION was not successfull - Defective cable, broken wires or bad contacts - Defective MOTOR ROLLER CLAMPING M6 - Electronic or electrical error of the IP HANDLING BOARD Cure: 1. Check the toothed belt of the MOTOR ROLLER CLAMPING M6 for proper position 2. Check the complete drive mechanism of MOTOR ROLLER CLAMPING M6 for proper function and smooth running 3. Restart the Digitizer and check, if the problem persists 4. Perform a calibration of the MOTOR ROLLER CLAMPING M6, refer to the Service Manual 5. Check the cabling and connection of the MOTOR ROLLER CLAMPING M6 6. Exchange the MOTOR ROLLER CLAMPING M6, refer to the Service Manual for order and replacement 7. Exchange the MOTOR ROLLER CLAMPING M6 and the IP HANDLING BOARD, refer to the Service Manual for order and replacement
6204 | INITIALIZE_DRAWER_FAILED | Error |  |  | The initialization of the RollerClamping Motor failed. | Meaning:The initialization of the Drawer Motor failed. Reason: Electronic or mechanical error of the Drawer module. Cure: Check if any obstacle blocks the motor movement; Retry, Restart, Call Service. Replace the SM Drawer if necessary; refer to the procedure for the exchange in the service documentation
6205 | NUM_STEPS_4_ABSPOS_ROLLER_CLAMPING_OUT_OF_RANGE | Error |  |  | The number of steps for determining the absolute position of the RollerClamping motor is out of range. | Meaning:The number of steps for determining the absolute position of the RollerClamping motor is out of range. Reason: Electronic or mechanical error of the RollerClamping module. Cure: Check if any obstacle blocks the motor movement; Retry, Restart, Call Service. Replace the SM RollerClamping if necessary; refer to the procedure for the exchange in the service documentation
6206 | NUM_STEPS_4_ABSPOS_DRAWER_OUT_OF_RANGE | Error |  |  | The number of steps for determining the absolute position of the Drawer motor is out of range. | Meaning:The number of steps for determining the absolute position of the Drawer motor is out of range. Reason: Electronic or mechanical error of the Drawer module. Cure: Check if any obstacle blocks the motor movement; Retry, Restart, Call Service. Replace the SM Drawer if necessary; refer to the procedure for the exchange in the service documentation
6207 | FEED_IP_OPEN_COVER_FAILED | Error |  |  | Feed the IP failed because the cover could not be opened. | Meaning: Feed the IP failed because the cover could not be opened. Reason: The cover could not be opened. Cure: Check the corresponding movement control (SW and stall parameters).
6208 | WRONG_UIC_PROTVER | Error |  |  | Protocol version of UIC is wrong | Meaning: Incompatible SW components (protocol version of UserInterface UIC is wrong). Reason: SW error. Cure: Install new SW version on all boards; inform R&D.
6656 | GeneralError | Error |  |  | General error in SSoFManager.
6657 | ApplicationConfigurationFailure | Error |  |  | Reading Application Configuration fails.
6658 | SSoFBridgeInitializationFailure | Error |  |  | Initialization of SSoFDigitizerBridge fails.
6659 | SSoFBridgeDeInitializationFailure | Error |  |  | Deinitialization of SSoFDigitizerBridge fails..
6660 | SSoFConfigurationFailure | Error |  |  | Configuration of SSoF fails.
6661 | SoftwareBusInitializationFailure | Error |  |  | Initialization of Softwarebus fails.
6662 | SSoFShutdownFailure | Error |  |  | Shutdown of SSoF fails.
6663 | SSoFDigitizerException | Error |  |  | Calling method at SSoFDigitizerBridge fails.
6664 | ConnectToServerException | Error |  |  | ConnectToServer fails.
6665 | DetachClientException | Error |  |  | DetachClient fails.
6666 | SynchronizationFailure | Error |  |  | Synchronization of EMS and DDM fails.
6667 | SyncThreadStoppingFailure | Error |  |  | Stopping SynchronizationThread fails.
6668 | SyncThreadStoppingException | Error |  |  | Stopping SynchronizationThread fails with exception.
6669 | SyncThreadStartingException | Error |  |  | Starting SynchronizationThread fails with exception.
7168 | CAS_LATENCY_NOT_SUPPORTED | Error |  |  | FPGA initialization failed because CAS_LATENCY is not supported | Meaning: FpgaAccess: FPGA initialization failed because CAS_LATENCY is not supported. Reason: SW error. Cure: Install new software; inform R&D team.
7169 | BURST_LENGTH_NOT_SUPPORTED | Error |  |  | FPGA initialization failed because BURST_LENGTH is not supported | Meaning: FpgaAccess: FPGA initialization failed because BURST_LENGTH is not supported. Reason: SW error. Cure: Install new software; inform R&D team.
7170 | SDRAM_CONTROLLER_START_FAILED | Error |  |  | FPGA initialization failed because SDRAM controller do not start | Meaning: FpgaAccess: FPGA initialization failed because SDRAM controller do not start. Reason: HW error. Cure: Replace PMI board.
7171 | TEST_FPGA_NOT_LOADED | Error |  |  | FPGA test failed FPGA is not loaded yet | Meaning: FpgaAccess: FPGA test failed FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case.
7172 | PATTGEN_FPGA_NOT_LOADED | Error |  |  | FPGA pattern generation failed because FPGA is not loaded yet | Meaning: FpgaAccess: FPGA pattern generation failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case.
7173 | PATTGEN_FLUSH_FIFO_UNDERRUN | Error |  |  | FPGA pattern generation failed because flush FIFO underrun occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO underrun occured. Reason: HW error. Cure: Replace PMI board.
7174 | PATTGEN_FLUSH_FIFO_OVERFLOW | Error |  |  | FPGA pattern generation failed because flush FIFO overflow occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO overflow occured. Reason: HW error. Cure: Replace PMI board.
7175 | PATTGEN_FLUSH_FIFO_DRAIN_NOK | Error |  |  | FPGA pattern generation failed because flush FIFO cannot be drained | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board.
7176 | FLUSH_FIFO_DRAIN_FAILED | Error |  |  | flush FIFO cannot be drained; HW problem, fatal error | Meaning: FpgaAccess: flush FIFO cannot be drained; HW problem, fatal error. Reason: HW error. Cure: Replace PMI board.
7177 | FPGA_TEST_FAILED | Error |  |  | FPGA test failed (one of the RAMs was not correctly accessible), fatal error | Meaning: FpgaAccess: FPGA test failed (one of the RAMs was not correctly accessible). Reason: HW error. Cure: Replace PMI board.
7178 | TIMEOUT_FILLING_SCAN_FIFO | Error |  |  | scan fifo write pump did not fill scan fifo with SDRAM data, fatal error | Meaning: FpgaAccess: scan fifo write pump did not fill scan fifo with SDRAM data. Reason: HW error. Cure: Replace PMI board.
7179 | UNKNOWN_FLUSH_FIFO_INPUT | Error |  |  | unknown input for flush fifo selected, fatal software error | Meaning: FpgaAccess: unknown input for flush fifo selected. Reason: SW error. Cure: Install new software; inform R&D team.
7180 | FILLSDRAM_FPGA_NOT_LOADED | Error |  |  | filling the SDRAM failed because FPGA is not loaded yet | Meaning: FpgaAccess: filling the SDRAM failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case.
7181 | FILLSDRAM_IMAGE_TOO_BIG | Error |  |  | filling the SDRAM failed because desired image is too big | Meaning: FpgaAccess: filling the SDRAM failed because desired image is too big. Reason: SW error. Cure: Install new software; inform R&D team.
7182 | PATTGEN_SDRAM_SIZE_UNKNOWN | Error |  |  | FPGA pattern generation failed because size of SDRAM is not determined yet (call initFpga first!) | Meaning: FpgaAccess: FPGA pattern generation failed because size of SDRAM is not determined yet. Reason: SW error (call initFpga first!). Cure: Install new software; inform R&D team.
7183 | PATTGEN_EXCEEDS_SDRAM | Error |  |  | FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM | Meaning: FpgaAccess: FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM. Reason: SW error. Cure: Install new software; inform R&D team.
7184 | LOAD_FPGA_CONFIGURATION | Error |  |  | Can't load FPGA configuration, fatal error | Meaning: FpgaAccess: Can't load FPGA configuration. Reason: Hardware defect or software mismatch. Cure: Upload PMI software and FPGA configuration; if still NOK, replace PMI board
7185 | LOAD_FPGA_HARDWARE | Error |  |  | Load FPGA: can't set FPGA into configuration mode, fatal error | Meaning: FpgaAccess: Load FPGA can't set FPGA into configuration mode. Reason: HW error. Cure: Replace PMI board.
7186 | STORED_FPGA_CONFIGURATION | Error |  |  | No valid FPGA version found on PMI | Meaning: FPGA version doesn't fit to hard- or software version. Reason: SW or HW error. Cure: Update software including FPGA configuration; if still NOK, replace PMI board
7187 | READ_PMI_CPU | Error |  |  | Can't read from FPGA CPU | Meaning: Can't read from FPGA CPU. Reason: Hardware error on PMI. Cure: Retry switch digitizer off/on; if still NOK, replace PMI board.
7188 | CAL_ERROR_FLAG | Error |  |  | PMI CPU Error flag is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Polygon or Laser or PMI board defective. Cure: Retry switch digitizer off/on, replace lens assembly CR30X, replace PMI board.
7189 | CAL_EOL | Error |  |  | EOL error counter is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Optical path is partly interrupted, calibration board is defective. Cure: Clean scan line with antistatic brush, check mounting of light guide, switch digitizer off/on, replace calibration plate
7190 | CAL_CAL_SENSORS | Error |  |  | Calibration sensor error counter is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Optical path is partly interrupted, calibration board is defective. Cure: Clean scan line with antistatic brush, switch digitizer off/on, replace calibration plate
7680 | IMAGE_SERVER_CREATION_FAILED | Error |  |  | ImageTransferControl cannot create image data server; fatal! | Meaning: ImageTransferControl cannot create image data server. Reason: SW error. Cure: Install new software.
7681 | INVALID_MAX_PACKET_SIZE | Error |  |  | ImageTransferControl received wrong FW packet size from PC; fatal! | Meaning: ImageTransferControl received wrong FW packet size from PC. Reason: SW error. Cure: Install new software; inform R&D team.
7682 | INVALID_FW_SPEED_CODE | Error |  |  | ImageTransferControl determined invalid FW speed; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&D team in any case.
7683 | BLOCK_TRANSFER_TIMEOUT_EXCEEDED | Error |  |  | ImageTransferControl: block transfer takes too long; either FW cable was disconnected during the block transfer or a fatal HW error occured! | Meaning: ImageTransferControl: block transfer takes too long. Reason: The transmission of an image data block takes too long. Cure: Check if FW cable is connected correctly; if still NOK, replace FW cable; if still NOK, replace PMI board; if still NOK, replace FW-board on PC.
7684 | INVALID_PC_NODE_ID | Error |  |  | ImageTransferControl: block transfer failed because PC's node ID is invalid; fatal error! | Meaning: ImageTransferControl: block transfer failed because PC's node ID is invalid. Reason: SW error. Cure: Install new software; inform R&D team.
7685 | SCAN_FIFO_OVERFLOW | Error |  |  | ImageTransferControl: scan fifo overflow during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo overflow during block transfer. Reason: HW error. Cure: Replace PMI board.
7686 | SCAN_FIFO_UNDERRUN | Error |  |  | ImageTransferControl: scan fifo underrun during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo underrun during block transfer. Reason: HW error. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI board and PMT/lens assembly; if OK, replace PMI board; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly.
7687 | FPGA_BUS_ERROR | Error |  |  | ImageTransferControl: FPGA bus error occured; fatal! | Meaning: ImageTransferControl: FPGA bus error occured. Reason: HW error. Cure: Replace PMI board.
7688 | INVALID_NR_OF_REQUESTED_BYTES | Error |  |  | ImageTransferControl: invalid number of requested bytes ( <=0); fatal SW bug! | Meaning: ImageTransferControl: invalid number of requested bytes ( <=0). Reason: SW error. Cure: Install new software; inform R&D team.
7689 | TIMEOUT_GET_PACKET_SIZE | Error |  |  | ImageTransferControl: timeout: cannot determine used FW packet size; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&D team in any case.
7690 | TIMEOUT_DATA_NOT_AVAILABLE | Error |  |  | ImageTransferControl: timeout: scanner didn't store enough data in SDRAM yet to start block transfer; wait, then try again! | Meaning: ImageTransferControl: timeout. Reason: Scanner didn't store enough data in SDRAM yet to start block transfer. Cure: No problem, PC software has to wait and try again.
7691 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED | Error |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump. Reason: HW error (DM (DataMover) pump or LLC (LinkLayerController) failed). Cure: Replace PMI board.
7692 | RECOVER_PACKET_FAILED | Error |  |  | ImageTransferControl: recovering the FW packet failed; fatal error! | Meaning: ImageTransferControl: recovering the FW packet failed. Reason: HW error. Cure: Replace PMI board.
7693 | TOO_MANY_PACKETS_PER_BLOCK | Error |  |  | ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block); request smaller block! | Meaning: ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block). Reason: SW error (PC must request smaller block!). Cure: Install new software; inform R&D team.
7694 | FLUSH_FIFO_DRAIN_FAILED | Error |  |  | flush FIFO cannot be drained; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board.
7695 | PIXELVIEWER_FLUSH_FIFO_OVERFLOW | Error |  |  | flush FIFO overflow occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO overflow occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board.
7696 | PIXELVIEWER_FLUSH_FIFO_UNDERRUN | Error |  |  | flush FIFO underrun occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO underrun occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board.
7697 | PIXELVIEWER_CRC1_AND_2_NOT_EQUAL | Error |  |  | after scanning in PixelViewer mode, the CRC1 and 2 are not equal; HW problem, fatal error! | Meaning: ImageTransferControl: after scanning in PixelViewer mode, the CRC1 and 2 are not equal. Reason: HW error. Cure: Replace PMI board.
7698 | PIXELVIEWER_START_NEXT_SCAN_FAILED | Error |  |  | starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running; HW failure (e.g. BOL missing) or SW synchronization problem, fatal error! | Meaning: ImageTransferControl: starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running. Reason: HW failure (e.g. BOL missing) or SW synchronization problem. Cure: Install new software; if still NOK, replace PMI board; inform R&D team in any case.
7699 | DM_ERROR_HANDLER_SETUP_FAILED | Error |  |  | registering the callback function for DM errors failed | Meaning: ImageTransferControl: registering the callback function for DM errors failed. Reason: SW error. Cure: Install new software; inform R&D team.
7700 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_START | Error |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer. Reason: HW error. Cure: Replace PMI board.
7701 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_OVFL | Error |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow. Reason: HW error. Cure: Replace PMI board.
7702 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_URUN | Error |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun. Reason: HW error. Cure: Replace PMI board.
7703 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_RETRANS | Error |  |  | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission. Reason: HW error. Cure: Replace PMI board.
7704 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER | Error |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost.
7705 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_START | Error |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. Actual transfer was not started yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost.
7706 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_PUMP | Error |  |  | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. DM pump was not enabled yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image transmission automatically again, image is not lost.
7707 | IMAGE_BUFFER_CREATION_FAILED | Error |  |  | ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. | Meaning: ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. Reason: not enough memory resources (probably software bug) Cure: Install new software.
7708 | MORE_DATA_REQUESTED_THAN_SCANNED | Error |  |  | ImageTransferUDPControl: more data requested than scanned | Meaning: ImageTransferUDPControl: more data requested than scanned. Reason: Index out of range Cure: Install new software; inform R&D team.
7709 | SERVER_DISCONNECTED_DURING_BLOCK_TRANSFER | Error |  |  | ImageTransferUDPControl: block transfer stopped because server was disconnected! | Meaning: ImageTransferUDPControl: block transfer stopped because server was disconnected! Reason: Cable was disconnected during transmission of image data to PC. Cure: Connect cable again; PC will request the image trasnmission automatically again, image is not lost.
7710 | NO_DATA_IN_DRAM_TIMEOUT | Error |  |  | ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. | Meaning: ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. Reason: Copying data into DRAM takes too long, or, no data anymore in FPGA. Cure: Install new software; inform R&D team.
7711 | DMA_NOT_PREPARED_TIMEOUT | Error |  |  | ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. | Meaning: ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. Reason: Timing problem Cure: Install new software; inform R&D team.
7712 | PATGEN_ENABLE_PROBLEM | Error |  |  | ImageTransferUDPControl: The Pattern Generator cannot be enabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be enabled. Reason: Caould be software or hardware problem Cure: Install new software; inform R&D team.
7713 | PATGEN_DISABLE_PROBLEM | Error |  |  | ImageTransferUDPControl: The Pattern Generator cannot be disabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be disabled. Reason: Caould be software or hardware problem Cure: Install new software; inform R&D team.
7714 | IMAGE_TRANSFER_CANCELLED | Error |  |  | ImageTransferUDPControl: the image transfer was cancelled | Meaning: ImageTransferUDPControl: the image transfer was cancelled
7715 | DM_ERROR | Error |  |  | ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error | Meaning: ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error. Reason: A DM error occurred during the transmission of an image block. Cure: The PC will automatically request the image block again (up to 10 times); if all repetitions fail an image transer error is displayed; in the latter case reboot the PC, the PC will start the image transmission automatically again, the image is not lost as long as the digitizer is not switched off.
8192 | LOST_MESSAGES | Error |  |  | Lost logmessages
10752 | INCIDENCE_OF_LIGHT | Error |  |  | Incidence of light --> adj.PMM/PMT failed ; fatal! | Meaning: Its to bright to adjust the Photo Multiplier. Reason: Housing of digitizer open. Cure: Close the digitizer housing.
10753 | PMM_ADJUST | Error |  |  | PMM/PMT adjust failed | Meaning: Photo multiplier malfunction. Reason: - Cure: Switch digitizer off/on and try a scan, if failed again replace PMT + LIGHT COLLECTOR.
10754 | GEO_CALIBRATION | Error |  |  | Geometric and jitter calibration failed | Meaning: Geometric and jitter calibration failed. Reason: Optic module, calibration board, sensor path misadjusted. Cure: Check proper seating of optic module, clean scan path with dust brush.
10755 | START_POLYGON | Error |  |  | Start polygon failed
10756 | ADC_ADJUST | Error |  |  | ADC adjustment out of spec | Meaning: ADC offset adjust out of range. Reason: ADC defective or PMM/PMT offset voltage out of range. Cure: Switch digitizer off/on, replace PMI board, replace PMT + LIGHT COLLECTOR.
10757 | START_SCAN | Error |  |  | Start scan failed | Meaning: Start scan failed, because scan is already running Reason: Software error Cure: Update software or inform RDE
10758 | POLYGON_LOCKED | Error |  |  | Polygon get not locked | Meaning: Polygon can't reach the requested speed. Reason: Polygon defective. Cure: Replace optic unit.
10759 | PREPARE_SCAN | Error |  |  | Prepare Scan failed
10760 | TOO_LESS_DATA | Error |  |  | Too less data stored in SDRAM --> SDRAM filled with RAMP | Meaning: Too less data stored in SDRAM --> SDRAM filled with stripes. Reason: Electrostatic discharge, error in begin or end of line signal (BOL,EOL). Cure: Check grounding of calibration board and grounding of front cover.
10761 | TEST_SCAN | Error |  |  | Test scan failed --> BOL or EOL out of range | Meaning: One or both of the outermost sensors of the CALIBRATION BOARD, which detects the begin (BOL) and the end (EOL) of the scan line, does not detect the expected amount of incident laser beam from the OPTIC MODULE. The Digitizer is out of operation and requires service support. Reason: One or both of the outermost sensors of the CALIBRATION BOARD does not detect the expected amount of incident laser beam because of: - The optical path of the laser beam is obstructed (e.g. dust cover, particles on the sensors, broken light collector, foreign body, ...) - Misalignment of the OPTIC MODULE - Defective of any of the following cables or connectors: - CALIBRATION BOARD - DISTRIBUTOR BOARD - DISTRIBUTOR BOARD - PMI BOARD - Defective DISTRIBUTOR BOARD - Defective CALIBRATION BOARD - Defective PMI-BOARD - Defective OPTIC MODULE Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check, that the optical path of the laser beam is not obstructed. If applicable remove the dust cover, refer to Service Bulletin No.37 3. Check the position of the laser beam with the new CALIBRATION BOARD or the LASER BEAM DETECTION TOOLKIT, refer to the Service Manual, adjustments and Service Bulletin No.28 4. Check the cabling of the following boards: - CALIBRATION BOARD - DISTRIBUTOR BOARD - DISTRIBUTOR BOARD - PMI BOARD 5. Exchange DISTRIBUTOR BOARD, refer to the Service Manual for order and replacement 6. Exchange the CALIBRATION BOARD, refer to the Service Manual for order and replacement 7. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement 8. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
10762 | PMM_REFCAL | Error |  |  | Calculation of REFCAL from IPF out of range | Meaning: The IP depended calculation of sensitivity factor results in a value out of range. Reason: Wrong PMM/PMT adjustment values stored in PMT eLabel. Cure: Replace PMT + LIGHT COLLECTOR.
10763 | PMM_AGING | Error |  |  | PMT sensitivity become critical --> reaches 94% = 750V | Meaning: The PMTs end is coming soon. Cure: Replace PMT + LIGHT COLLECTOR
10764 | PMM_FINE_ADJUST | Error |  |  | PMM/PMT fine adjust failed --> possibly bad image quality | Meaning: PMT sensitivity adjustment failed Reason: - PMI POWER DISTRIBUTION board defective - PMT defective or end of life Cure: 1. Replace PMI POWER DISTRIBUTION board 2. If error still persists replace PMT + LIGHT COLLECTOR
10765 | START_GEO_CALIBRATION | Error |  |  | Start geometric and jitter calibration failed
10766 | IPF_OUT_OF_RANGE | Error |  |  | IP sensitivity factor out of range set to best fit
10767 | TEST_SCAN_CAL | Error |  |  | Test scan failed: Not all CAL sensors detected --> defect or IP in digitizer | Meaning: Can't detect the expected amount of calibration signals. Reason: Optic path interrupted, IP in digitizer, broken light collector, cable PMI to calibration board defective, CAL-board defective Cure: Clean the scan line, check optic path, check cable PMI to calibration board, replace calibration board.
10768 | READ_PMM_ELABEL | Error |  |  | Can't read PMM settings from PMM eLabel
10769 | READY_TO_SCAN_CAL_SENSORS | Error |  |  | Calibration sensor defective or two IP's in scanner | Meaning: Can't detect the expected amount of calibration signals. Reason: CAL-board defective or optic path interrupted (IP or dust). Cure: Clean the scan line, replace calibration board.
10770 | PMM_BLUE_LED | Error |  |  | Photo Multiplier Module (PMM/PMT), blue LED or power supply board defective | Meaning: Can't detect PMT blue LED. Reason: - PMT blue LED defective or not connected - PMI POWER DISTRIBUTION board defective - PMT defective Cure: 1. Check connection and cable from blue LED to PMT 2. Replace PMI POWER DISTRIBUTION board 3. If error still persists replace PMT + LIGHT COLLECTOR
10771 | FPGA_BUS | Error |  |  | FPGA BUS error detected | Meaning: FPGA BUS error occurred. Reason: unknown Cure: Restart digitizer, if error occurred again replace PMI board.
10772 | AUTOEOL_ADDED | Error |  |  | FPGA has added an EOL at end of EOL valid | Meaning: No EOL detected, FPGA adds one at end of EOL valid window. Reason: polygon speed out of range. Cure: Restart digitizer, replace optic module.
10773 | PMM_OFFSET_VOLTAGE | Error |  |  | PMM/PMT offset voltage out of range | Meaning: PMT offset voltage adjustment failed Reason: - Incidence of light - PMI POWER DISTRIBUTION board defective - PMT defective (end of life) Cure: 1. Make sure all covers are closed 2. Replace PMI POWER DISTRIBUTION board 3. If error still persists replace PMT + LIGHT COLLECTOR
11264 | PC_TO_DIC_ADAPTER_RECEIVE_DEFAULT_NOT_POSSIBLE | Error |  |  | Adapter tried to receive a message using the 'receiveDefault' method, but it contains unsupported number of bytes
11265 | BASE_APP_SEND_FAILED_NO_CONNECTION | Error |  |  | RTPort_send failed because connection to receiver is not established
11266 | BASE_APP_SEND_FAILED_NO_MEMORY | Error |  |  | RTPort_send failed because no memory available (list of unused messageobjects exhausted)
11267 | BASE_APP_SEND_FAILED_INVALID_PRIO | Error |  |  | RTPort_send failed because message priority is invalid
11268 | BASE_APP_SEND_FAILED_OTHER_REASON | Error |  |  | RTPort_send failed for any other reason
11776 | SM_SELFTEST_FAILED | Error |  |  | Procedure selftest of SMx terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11777 | SM_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of SMx terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11778 | SM_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of SMx terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11779 | SM_COS_COIL_FAILURE | Error |  |  | Procedure selftest of SMx terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11780 | SM_DRIVER_FAILURE | Error |  |  | Procedure selftest of SMx terminated with DIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11781 | SM_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of SMx terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11782 | ADC_MEAS_TOT_FAILURE | Error |  |  | Selftest of SMx ended due to ADC measurement timeout failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service.
11876 | WARN_INVALID_PAR | Error |  |  | Warning: Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
11877 | WARN_INVALID_ACCESS_FUNC | Error |  |  | Warning: Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
11878 | WARN_INVALID_POINTER | Error |  |  | Warning: Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
11879 | WARN_LOW_VOLT_MEAS_POS | Error |  |  | Warning: Test pos current, measured voltage value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at driver voltage of an active pos current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11880 | WARN_LOW_VOLT_MEAS_NEG | Error |  |  | Warning: Test neg current, measured voltage value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at driver voltage of an active neg current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11881 | WARN_LOW_SIN_CURRENT_MEAS | Error |  |  | Warning: Test measurement sin value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at sin current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11882 | WARN_LOW_COS_CURRENT_MEAS | Error |  |  | Warning: Test measurement cos value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at cos current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11883 | WARN_LOW_SIN_NEG_CURRENT_MEAS | Error |  |  | Warning: Test measurement sin neg value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at sin negative current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
11884 | WARN_LOW_COS_NEG_CURRENT_MEAS | Error |  |  | Warning: Test measurement cos neg value has low reserve to limit | Meaning: Retry possible, function normal, risk for malfunction increased. Reason: The SM Test at cos negative current has a low reserve to the error limit. Cure: Restart, check electronic and motor.
12288 | FIXER_SELFTEST_FAILED | Error |  |  | Procedure selftest of FIXER terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor FIXER terminated its selftest with an errror condition Cure: Retry, Restart, check cabling, motor and IP handling board.
12289 | FIXER_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of FIXER terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor FIXER terminated its selftest with an error condition Cure: Retry, restart, check cabling, motor and IP handling board.
12290 | FIXER_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of FIXER terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor FIXER terminated its selftest with an error condition Cure: Retry, Restart, check cabling and motor. Call Service.
12291 | FIXER_COS_COIL_FAILURE | Error |  |  | Procedure selftest of FIXER terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor FIXER terminated its selftest with an error condition Cure: Check motor cabling and motor.
12292 | FIXER_DRIVER_FAILURE | Error |  |  | Procedure selftest of FIXER terminated with DRIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor FIXER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12293 | FIXER_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of FIXER terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor FIXER terminated its selftest with an error condition Cure: Retry, restart, check cabling, motor and IP handling board.
12294 | LOCKER_SELFTEST_FAILED | Error |  |  | Procedure selftest of LOCKER terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor LOCKER terminated its selftest with an errror condition Cure: Retry, Restart, check cabling, motor and IP handling board.
12295 | LOCKER_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of LOCKER terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12296 | LOCKER_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of LOCKER terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12297 | LOCKER_COS_COIL_FAILURE | Error |  |  | Procedure selftest of LOCKER terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12298 | LOCKER_DRIVER_FAILURE | Error |  |  | Procedure selftest of LOCKER terminated with DRIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12299 | LOCKER_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of LOCKER terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor LOCKER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12300 | OPENER_SELFTEST_FAILED | Error |  |  | Procedure selftest of OPENER terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its selftest with an error condition Cure: Retry, Restart, check cabling and motor. Call Service.
12301 | OPENER_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of OPENER terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12302 | OPENER_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of OPENER terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12303 | OPENER_COS_COIL_FAILURE | Error |  |  | Procedure selftest of OPENER terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12304 | OPENER_DRIVER_FAILURE | Error |  |  | Procedure selftest of OPENER terminated with DRIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12305 | OPENER_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of OPENER terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its selftest with an error condition Cure: Retry, Restart, Call Service.
12306 | OPENER_CLOSE_COVER_CAL_POS_FAILED | Error |  |  | Procedure closeCoverCal for calibration of OPENER terminated with detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its movement with an error condition Cure: Retry, Restart, Call Service.
12307 | OPENER_OPEN_COVER_CAL_POS_FAILED | Error |  |  | Procedure openCoverCal for calibration of OPENER terminated with detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor OPENER terminated its movement with an error condition Cure: Retry, Restart, Call Service.
12308 | OPENER_OPEN_POS_STEP_DEVIATION | Error |  |  | Opener Open Pos Movement terminated with failure. | Meaning: Function disturbed. Open Cover Pos not found. Reason: Moved steps have deviation above the accepted limit or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12309 | OPENER_CLOSE_POS_STEP_DEVIATION | Error |  |  | Opener Close Pos Movement terminated with failure. | Meaning: Function disturbed. Close Cover position not found. Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Cure: Restart Digitizer. If error persists please call Service.
12310 | OPENER_CLOSE_POS_FAILED | Error |  |  | Opener Close Pos Movement terminated with failure. | Meaning: Function disturbed. Openpos not found. Reason: Maximum step limit reached, or mechanical position blocked. Cure: Restart Digitizer. If error persists please call Service.
12311 | OPENER_OPEN_POS_MAX_STEP_LIMIT | Error |  |  | Opener Open Pos Movement terminated with failure. | Meaning: Function disturbed. Open Cover Pos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12312 | OPENER_OPEN_POS_FAILED | Error |  |  | Opener Open Pos Movement terminated with failure. | Meaning: Function disturbed. Open Cover Pos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12313 | FIXER_RE_CLAMP_POS_FAILED | Error |  |  | Fixer Re-Clamp Pos Movement terminated with failure. | Meaning: Function disturbed. Fixer Clamp Pos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12314 | FIXER_RE_UNCLAMP_POS_FAILED | Error |  |  | Fixer Re-UnClamp Pos Movement terminated with failure. | Meaning: Function disturbed. Fixer UnClamp Pos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12315 | FIXER_UNCLAMP_POS_STEP_DEVIATION | Error |  |  | Fixer UnClampPos Movement terminated with failure. | Meaning: Function disturbed. UnClamp position not found. Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Cure: Restart Digitizer. If error persists please call Service.
12316 | FIXER_UNCLAMP_POS_MAX_STEP_LIMIT | Error |  |  | Fixer Home UnClamp Movement stopped due to maximum step limit | Meaning: Function disturbed. Unclamp-Pos not found. Reason: Maximum step limit reached. Cure: Restart Digitizer. If error persists please call Service.
12317 | FIXER_CLAMP_CASS_FAILED | Error |  |  | Fixer Clamp Cassette Movement terminated with failure. | Meaning: Function disturbed. Fixer Clamp Pos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12318 | LOCKER_LOCK_IP_MOVE_STOP_UNEXPECTED | Error |  |  | Locker home movement stopped too early. | Meaning: Retry possible, function may be disturbed. Reason: The Locker home movement stopped early as stall reacted in position range before stopper. Cure: Restart Digitizer. If error persists please call Service.
12319 | LOCKER_LOCK_IP_POS_STEP_DEVIATION | Error |  |  | Locker Lock Ip Pos Movement terminated with failure. | Meaning: Function disturbed. LockIp position not found. Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Cure: Restart Digitizer. If error persists please call Service.
12320 | LOCKER_LOCK_IP_POS_FAILED | Error |  |  | Locker Lock Ip Pos Movement terminated with failure. | Meaning: Function disturbed. Lock Ip pos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12321 | LOCKER_UNLOCK_IP_POS_STEP_DEVIATION | Error |  |  | Locker UnLock Ip Pos Movement terminated with failure. | Meaning: Function disturbed. UnLockIp position not found. Reason: Moved steps have deviation above the accepted limit, or mechanical position blocked. Cure: Restart Digitizer. If error persists please call Service.
12322 | LOCKER_UNLOCK_IP_POS_FAILED | Error |  |  | Locker UnLock Ip Pos Movement terminated with failure. | Meaning: Function disturbed. UnLock Ip pos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12323 | LOCKER_UNLOCK_IP_MOVE_STOP_UNEXPECTED | Error |  |  | Locker UnLock Ip Pos Movement terminated with failure. | Meaning: Function disturbed. UnLock Ip pos not found. Reason: Stop before breakpoint step limit reached, or mechanical way blocked. Cure: Restart Digitizer. If error persists please call Service.
12324 | LOCKER_ABS_POS_FAILED | Error |  |  | Locker Absolute Pos Movement terminated with failure. | Meaning: Function disturbed. Homepos/Absolutepos not found. Reason: Maximum step limit reached, or mechanical stop not found. Cure: Restart Digitizer. If error persists please call Service.
12325 | LOCKER_CAL_LOCK_IP_MOVE_STOP_UNEXPECTED | Error |  |  | Locker calibration Lock Ip Pos movement stopped too early. | Meaning: Retry possible, function may be disturbed. Reason: The Locker calibration Lock Ip Pos movement stopped early as stall reacted in position range before target position reached. Cure: Restart Digitizer. If error persists please call Service.
12326 | LOCKER_CAL_LOCK_IP_MOVE_FAILED | Error |  |  | Locker calibration Lock Ip Pos movement terminated with failure. | Meaning: Retry possible, function may be disturbed. Reason: Maximum step limit reached, or mechanical stop at wrong position. Cure: Restart Digitizer. If error persists please call Service.
12327 | LOCKER_CAL_UNLOCK_IP_MOVE_FAILED | Error |  |  | Locker calibration UnLock Ip Pos movement terminated with failure. | Meaning: Retry possible, function may be disturbed. Reason: Maximum step limit reached, or mechanical problem Cure: Restart Digitizer. If error persists please call Service.
12328 | FIXER_CLAMP_CASS_CAL_MAX_STEP_LIMIT | Error |  |  | Clamp Cassette Movement stopped due to maximum step limit | Meaning: Function disturbed. Reason: Maximum step limit reached. Cure: Re-Start Calibration Process Cassette. Restart Digitizer. If error persists please call Service.
12329 | PREPARE_SM_SELFTEST_FAILED | Error |  |  | Preparation of the Selftest of the stepper motors failed. | Meaning: Function disturbed. Reason: SW-Function to set all motor currents failed Cure: Load new SW. Restart Digitizer. If error persists please call Service.
12388 | WARN_FIXER_INVALID_PAR | Error |  |  | Warning: FIXER Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The FIXER Test has a invalid par. Cure: Restart, Call Service
12389 | WARN_FIXER_INVALID_ACCESS_FUNC | Error |  |  | Warning: FIXER Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The FIXER Test has a invalid par. Cure: Restart, Call Service
12390 | WARN_FIXER_INVALID_POINTER | Error |  |  | Warning: FIXER Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The FIXER Test has a invalid par. Cure: Restart, Call Service
12391 | WARN_LOCKER_INVALID_PAR | Error |  |  | Warning: LOCKER Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The LOCKER Test has a invalid par. Cure: Restart, Call Service
12392 | WARN_LOCKER_INVALID_ACCESS_FUNC | Error |  |  | Warning: LOCKER Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The LOCKER Test has a invalid par. Cure: Restart, Call Service
12393 | WARN_LOCKER_INVALID_POINTER | Error |  |  | Warning: LOCKER Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The LOCKER Test has a invalid par. Cure: Restart, Call Service
12394 | WARN_OPENER_INVALID_PAR | Error |  |  | Warning: OPENER Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The OPENER Test has a invalid par. Cure: Restart, Call Service
12395 | WARN_OPENER_INVALID_ACCESS_FUNC | Error |  |  | Warning: OPENER Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The OPENER Test has a invalid par. Cure: Restart, Call Service
12396 | WARN_OPENER_INVALID_POINTER | Error |  |  | Warning: OPENER Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The OPENER Test has a invalid par. Cure: Restart, Call Service
12800 | CLAMPER_SELFTEST_FAILED | Error |  |  | Procedure selftest of Clamper terminated with failure. | Meaning: Self test of the MOTOR ROLLER CLAMPING M6 failed repeatedly. Electrical self test of stepper motor M6 for ROLLER CLAMPING terminated with an error. Repeat test cycles shows the same error. The Digitizer is out of operation and requires service support. Reason: Self test of the MOTOR ROLLER CLAMPING M6 failed repeatedly due to: - Defective cable, broken wires or bad contacts - Defective MOTOR ROLLER CLAMPING M6 - Electronic or electrical error of the IP HANDLING BOARD. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection of the MOTOR ROLLER CLAMPING M6 3. Exchange the MOTOR ROLLER CLAMPING M6, refer to the Service Manual for order and replacement 4. Exchange the MOTOR ROLLER CLAMPING M6 and the IP HANDLING BOARD, refer to the Service Manual for order and replacement 5. Install the most recent Digitizer software.
12801 | CLAMPER_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of Clamper terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller Clamping if necessary; refer to the procedure for the exchange in the service documentation.
12802 | CLAMPER_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of Clamper terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller Clamping if necessary; refer to the procedure for the exchange in the service documentation.
12803 | CLAMPER_COS_COIL_FAILURE | Error |  |  | Procedure selftest of Clamper terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller Clamping if necessary; refer to the procedure for the exchange in the service documentation.
12804 | CLAMPER_DRIVER_FAILURE | Error |  |  | Procedure selftest of Clamper terminated with DIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller Clamping if necessary; refer to the procedure for the exchange in the service documentation.
12805 | CLAMPER_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of Clamper terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller Clamping if necessary; refer to the procedure for the exchange in the service documentation.
12806 | DRAWER_SELFTEST_FAILED | Error |  |  | Procedure selftest of Drawer terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Drawer if necessary; refer to the procedure for the exchange in the service documentation.
12807 | DRAWER_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of Drawer terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Drawer if necessary; refer to the procedure for the exchange in the service documentation.
12808 | DRAWER_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of Drawer terminated with SIN COIL failure. | Meaning: Self test of the DRAWER MOTOR M4 failed repeatedly. Electrical self test of stepper motor M4 of the DRAWER UNIT terminated with an error. Repeat test cycles shows the same error. The Digitizer is out of operation and requires service support. Reason: Self test of the DRAWER MOTOR M4 failed repeatedly due to: - Defective cable, broken wires or bad contacts - Defective DRAWER MOTOR M4 - Electronic or electrical error of the IP HANDLING BOARD Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection of the DRAWER MOTOR M4 3. Exchange the DRAWER MOTOR M4, refer to the Service Manual for order and replacement 4. Exchange the DRAWER MOTOR M4 and the IP HANDLING BOARD, refer to the Service Manual for order and replacement 5. Install the most recent Digitizer software.
12809 | DRAWER_COS_COIL_FAILURE | Error |  |  | Procedure selftest of Drawer terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Drawer if necessary; refer to the procedure for the exchange in the service documentation.
12810 | DRAWER_DRIVER_FAILURE | Error |  |  | Procedure selftest of Drawer terminated with DIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Drawer if necessary; refer to the procedure for the exchange in the service documentation.
12811 | DRAWER_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of Drawer terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Drawer if necessary; refer to the procedure for the exchange in the service documentation.
12812 | ROLLER_SELFTEST_FAILED | Error |  |  | Procedure selftest of Roller terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller if necessary; refer to the procedure for the exchange in the service documentation.
12813 | ROLLER_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of Roller terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller if necessary; refer to the procedure for the exchange in the service documentation.
12814 | ROLLER_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of Roller terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller if necessary; refer to the procedure for the exchange in the service documentation.
12815 | ROLLER_COS_COIL_FAILURE | Error |  |  | Procedure selftest of Roller terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller if necessary; refer to the procedure for the exchange in the service documentation.
12816 | ROLLER_DRIVER_FAILURE | Error |  |  | Procedure selftest of Roller terminated with DIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller if necessary; refer to the procedure for the exchange in the service documentation.
12817 | ROLLER_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of Roller terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Roller if necessary; refer to the procedure for the exchange in the service documentation.
12900 | WARN_CLAMPER_INVALID_PAR | Error |  |  | Warning: Clamper Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12901 | WARN_CLAMPER_INVALID_ACCESS_FUNC | Error |  |  | Warning: Clamper Test Clamper used invalid access par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12902 | WARN_CLAMPER_INVALID_POINTER | Error |  |  | Warning: Clamper Test Clamper used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12903 | WARN_DRAWER_INVALID_PAR | Error |  |  | Warning: Drawer Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12904 | WARN_DRAWER_INVALID_ACCESS_FUNC | Error |  |  | Warning: Drawer Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12905 | WARN_DRAWER_INVALID_POINTER | Error |  |  | Warning: Drawer Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12906 | WARN_ROLLER_INVALID_PAR | Error |  |  | Warning: Roller Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12907 | WARN_ROLLER_INVALID_ACCESS_FUNC | Error |  |  | Warning: Roller Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12908 | WARN_ROLLER_INVALID_POINTER | Error |  |  | Warning: Roller Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
12909 | PREPARE_SM_SELFTEST_FAILED | Error |  |  | The preparation of the self test of the stepper motors failed. | Meaning: Function disturbed. Reason: SW-Function to set all motor currents failed. Cure: Load new SW. Restart Digitizer. If error persists please call Service.
13312 | SLOW_SCAN_SELFTEST_FAILED | Error |  |  | Procedure selftest of Slow_Scan terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan if necessary; refer to the procedure for the exchange in the service documentation.
13313 | SLOW_SCAN_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of Slow_Scan terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan if necessary; refer to the procedure for the exchange in the service documentation.
13314 | SLOW_SCAN_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of Slow_Scan terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan if necessary; refer to the procedure for the exchange in the service documentation.
13315 | SLOW_SCAN_COS_COIL_FAILURE | Error |  |  | Procedure selftest of Slow_Scan terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan if necessary; refer to the procedure for the exchange in the service documentation.
13316 | SLOW_SCAN_DRIVER_FAILURE | Error |  |  | Procedure selftest of Slow_Scan terminated with DIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan if necessary; refer to the procedure for the exchange in the service documentation.
13317 | SLOW_SCAN_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of Slow_Scan terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan if necessary; refer to the procedure for the exchange in the service documentation.
13318 | SLOW_SCAN_LIFT_SELFTEST_FAILED | Error |  |  | Procedure selftest of Slow_Scan_Lift terminated with failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan Lift if necessary; refer to the procedure for the exchange in the service documentation.
13319 | SLOW_SCAN_LIFT_UNPLUGGED_FAILURE | Error |  |  | Procedure selftest of Slow_Scan_Lift terminated with UNPLUGGED failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan Lift if necessary; refer to the procedure for the exchange in the service documentation.
13320 | SLOW_SCAN_LIFT_SIN_COIL_FAILURE | Error |  |  | Procedure selftest of Slow_Scan_Lift terminated with SIN COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan Lift if necessary.
13321 | SLOW_SCAN_LIFT_COS_COIL_FAILURE | Error |  |  | Procedure selftest of Slow_Scan_Lift terminated with COS COIL failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan Lift if necessary; refer to the procedure for the exchange in the service documentation.
13322 | SLOW_SCAN_LIFT_DRIVER_FAILURE | Error |  |  | Procedure selftest of Slow_Scan_Lift terminated with DIVER failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan Lift if necessary; refer to the procedure for the exchange in the service documentation.
13323 | SLOW_SCAN_LIFT_GROUND_CURRENT_FAILURE | Error |  |  | Procedure selftest of Slow_Scan_Lift terminated with overcurrent to Ground detected failure. | Meaning: Retry possible, function disturbed. Reason: The Handling Motor SM terminated its selftest with an errror condition Cure: Retry, Restart, Call Service. Replace the SM Slow Scan Lift if necessary; refer to the procedure for the exchange in the service documentation.
13412 | WARN_SLOW_SCAN_INVALID_PAR | Error |  |  | Warning: Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
13413 | WARN_SLOW_SCAN_INVALID_ACCESS_FUNC | Error |  |  | Warning: Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
13414 | WARN_SLOW_SCAN_INVALID_POINTER | Error |  |  | Warning: Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
13415 | WARN_SLOW_SCAN_LIFT_INVALID_PAR | Error |  |  | Warning: Test detected invalid par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
13416 | WARN_SLOW_SCAN_LIFT_INVALID_ACCESS_FUNC | Error |  |  | Warning: Test used invalid access par | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
13417 | WARN_SLOW_SCAN_LIFT_INVALID_POINTER | Error |  |  | Warning: Test used invalid pointer | Meaning: Retry possible, function disturbed. Reason: The SM Test has a invalid par. Cure: Restart, Call Service
13418 | PREPARE_SM_SELFTEST_FAILED | Error |  |  | The preparation of the self test of the stepper motors failed. | Meaning: Function disturbed. Reason: SW-Function to set all motor currents failed. Cure: Load new SW. Restart Digitizer.
13824 | INIT_MODULE_FAILED_PLACEHOLDER | Error |  |  | Testlog sent during initialisation phase | Meaning: Serious error, function disturbed. Reason: The needed elements could not be intialised to valid values or function Cure: Restart Digitizer. If error persists please call Service.
13924 | WARN_CASS_INVALID_PLACEHOLDER | Error |  |  | Warning: Clamp Cassette Movement detected invalid Cassette Width | Meaning: Retry possible, function disturbed. Reason: The clamped Cassette has a invalid width. Cure: Re-clamp Cassette. If error persists please calibrate mechanical Cassette Unit.
13925 | FIXER_CALIBRATION_FAILED | Error |  |  | The calibration of the Fixer failed. | Meaning: The calibration of the Fixer failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
13926 | FIXER_CALIBRATION_ABORTED | Error |  |  | The calibration procedure of the Fixer was aborted. | Meaning: The calibration procedure of the Fixer was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
13927 | FIXER_CALIBRATION_INVALID_VALUE | Error |  |  | The determined optimal run current value of the Fixer was invalid. | Meaning: The determined optimal run current value of the Fixer was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
13928 | LOCKER_CALIBRATION_FAILED | Error |  |  | The calibration of the Locker failed. | Meaning: The calibration of the Locker failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
13929 | LOCKER_CALIBRATION_ABORTED | Error |  |  | The calibration procedure of the Locker was aborted. | Meaning: The calibration procedure of the Locker was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
13930 | LOCKER_CALIBRATION_INVALID_VALUE | Error |  |  | The determined optimal run current value of the Locker was invalid. | Meaning: The determined optimal run current value of the Locker was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
13931 | OPENER_CALIBRATION_FAILED | Error |  |  | The calibration of the Opener failed. | Meaning: The calibration of the Opener failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
13932 | OPENER_CALIBRATION_ABORTED | Error |  |  | The calibration procedure of the Opener was aborted. | Meaning: The calibration procedure of the Opener was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
13933 | OPENER_CALIBRATION_INVALID_VALUE | Error |  |  | The determined optimal run current value of the Opener was invalid. | Meaning: The determined optimal run current value of the Opener was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14436 | DRAWER_CALIBRATION_FAILED | Error |  |  | The calibration of the Drawer failed. | Meaning: The calibration of the Drawer failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14437 | DRAWER_CALIBRATION_ABORTED | Error |  |  | The calibration procedure of the Drawer was aborted. | Meaning: The calibration procedure of the Drawer was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14438 | DRAWER_CALIBRATION_INVALID_VALUE | Error |  |  | The determined optimal run current value of the Drawer was invalid. | Meaning: The determined optimal run current value of the Drawer was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14439 | ROLLER_CLAMPING_CALIBRATION_FAILED | Error |  |  | The calibration of the RollerClamping failed. | Meaning: The calibration of the RollerClamping failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding mechanics and electronics.
14440 | ROLLER_CLAMPING_CALIBRATION_ABORTED | Error |  |  | The calibration procedure of the RollerClamping was aborted. | Meaning: The calibration procedure of the RollerClamping was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
14441 | ROLLER_CLAMPING_CALIBRATION_INVALID_VALUE | Error |  |  | The determined optimal run current value of the RollerClamping was invalid. | Meaning: The determined optimal run current value of the RollerClamping was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding mechanics or electronics. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
14848 | READ_OPTICS_TYPE | Error |  |  | Can't read optics type from optics eLabel -> using default type | Meaning: Default value for optic type are used. Reason: Unreadable entry in 'optic type' of optic module elabel. Cure: Replace optic module.
14849 | READ_POLYGON_TYPE | Error |  |  | Can't read polygon type from optics eLabel -> using default type | Meaning: Default values for polygon type are used. Reason: Unreadable entry in polygon type of optic module elabel. Cure: Replace optic module.
14850 | READ_OPERATION_MODE | Error |  |  | Can't read operation mode from PMI eLabel -> setting to normal mode | Meaning: 'Normal mode' is used as operation mode. Reason: Unreadable entry in 'operation mode' of PMI elabel. Cure: Replace PMI board.
14851 | READ_POLYGON_SPEED_OFFSET | Error |  |  | Can't read polygon speed offset from optics eLabel -> using offset 0
14852 | INIT_PMM | Error |  |  | Initializing I2C hardware for PMM failed | Meaning: Communication error PMI to PMM. Reason: PMI and PMM connection not OK or PMI or PMM defective. Cure: 1) Check PMM cabling. 2) Replace PMM cabling. 3) Replace PMI board. 4) Replace PMM.
14853 | CAL_BOARD_CONNECTOR | Error |  |  | Calibration board not detected | Meaning: Calibration board not detected. Reason: Cable broken or calibration board defective. Cure: 1) Check calibration board cabling. 2) Replace calibration board, if error persists.
14854 | OPTICS_BOARD_CONNECTOR | Error |  |  | Optics board not detected | Meaning: Optic board not detected. Reason: Cable broken/not connected or optic board defective. Cure: 1) Check optic module cabling. 2) Replace optic module, if error persists.
14855 | LASER_DIODE | Error |  |  | Check laser diode failed | Meaning: Check laser diode failed. Reason: Laser diode defective. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
14856 | OPTICS_TYPE | Error |  |  | Unknown optics type detected | Meaning: Unknown optics type detected Reason: Unknown optics type in elabel. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
14857 | INIT_ELABEL | Error |  |  | Error initializing optics eLabel | Meaning: Initialization of OPTIC MODULE e-label (module specific parameter storage) failed. The Digitizer is out of operation and requires service support. Reason: OPTIC MODULE e-label can't be initialized due to: - OPTIC MODULE not connected - Defective of the PMI-BOARD - OPTIC MODULE cable or connectors - Corrupted data structure on the e-label - Defective OPTIC MODULE Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between PMI-BOARD and OPTIC MODULE 3. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement 4. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
14858 | PMM_ELABEL | Error |  |  | Error getting data from PMM eLabel | Meaning: Can't read data from PMM eLabel. Reason: PMM eLabel defective or PMM not connected. Cure: 1) Check PMM cabling. 2) Replace PMM cabling. 3) Replace PMI board. 4) Replace PMM.
14859 | PMM_BOARD_CONNECTOR | Error |  |  | PMM board not detected | Meaning: PMM board not detected Reason: PMM is not connected to PMI board. Cure: 1) Check PMM cabling. 2) Replace PMM cabling, if error persists. 3) Replace PMM, if error persists.
14860 | START_POLYGON | Error |  |  | Start or Lock polygon failed | Meaning: Polygon does not reach its speed. Reason: Power supply or polygon defective. Cure: Check power supply, replace optics module.
14861 | UPDATE_PMM | Error |  |  | Error: PMM firmware update failed - Switch digitizer Off/On for retry | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists.
14862 | UPDATE_PMM_FAILED | Error |  |  | Error: PMM firmware update failed | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists.
14863 | PMMADJ_LPV_TOO_HIGH | Error |  |  | Prepare PMM adjust: LPV is to high
14864 | PMMADJ_LED_BRIGHTNESS | Error |  |  | Prepare PMM adjust: Error LED not bright enough
14865 | PMMADJ_SEL_SENS | Error |  |  | Prepare PMM adjust: Error selecting sensitivity
14866 | PMMADJ_INCIDENCE_OF_LIGHT | Error |  |  | Prepare PMM adjust: Error because incidence of light
14867 | PMMADJ_WRITE_ROV | Error |  |  | Analyse PMM adjust: Error writing new REFCAL to PMM
14868 | PMMADJ_PARAMETER | Error |  |  | Analyse PMM adjust: Error parameter real dose or target index out of range
14869 | BOL_TIMEOUT | Error |  |  | Lost BOL signal | Meaning: ESD or optics problem. Reason: electric discharge or optics module defective. Cure: if permanent check optics module.
14870 | PMI_ELABEL | Error |  |  | Error getting data from PMI eLabel | Meaning: Error getting data from PMI eLabel. Reason: PMI eLabel defective. Cure: Replace PMI board.
14871 | DIST_BOL_EOL_OUT_OF_RANGE | Error |  |  | Distance BOL to EOL out of range | Meaning: Distance BOL to EOL out of range. Reason: Polygon speed out of range, Optics module out of range. Cure: Check polygon speed (image width), replace optics module.
14872 | PMI_HW_VER | Error |  |  | Software needs PMI-Board hardware version 1 or later | Meaning: PMI software does not fit to PMI hardware. Reason: Software update to an incompatible PMI board. Cure: Replace PMI board to index 1 or higher.
14873 | PMI_FPGA_VER | Error |  |  | FPGA version does not fit to software version | Meaning: PMI software does not fit to PMI FPGA hardware. Reason: Software update failed or not completed. Cure: Update PMI software.
14874 | OPTICS_ELABEL_CORRUPTED | Error |  |  | Optics eLabel data corrupted | Meaning: Optics eLabel can't be read, need to be restored from PC. Reason: Cable from PMI to optics module too long or unknown. Cure: Restore optics eLabel from PC.
14875 | LD_AGING | Error |  |  | Laser diode is aging | Meaning: Laser diode is supplied with additional electrical power (1.5 of the standard value) to compensate aging for a while. Reason: Laser diode is short before end of life.
15360 | DIC_TO_IHC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while DIC is sending to IHC | Meaning: No receiver found after sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (IHC) on IP-Handling board. Reason: Software error, input queue of IHC may be full. Cure: Install new software; inform R&D team.
15361 | DIC_TO_IHC_UNKNOWN_CHANNEL | Info |  |  | unknown CAN channel for sending from DIC to IHC | Meaning: The CAN channel number for sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (IHC) on IP-Handling board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15362 | DIC_TO_IHC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | DIC could not read acknowledge from IHC | Meaning: Could not read acknowledge after sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (IHC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on IP-Handling board; if connections are OK, install new software and inform R&D team.
15363 | DIC_TO_IHC_CAN_ERROR | Fatal error |  |  | sending from DIC to IHC failed with CAN fast channel error | Meaning: CAN error during sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (IHC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15364 | DIC_TO_CHC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while DIC is sending to CHC | Meaning: No receiver found after sending message from DigitizerControl (DIC) on PMI-board to CassetteHandlingControl (CHC) on IP-Handling board. Reason: Software error, input queue of CHC may be full. Cure: Install new software; inform R&D team.
15365 | DIC_TO_CHC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from DIC to CHC | Meaning: The CAN channel number for sending message from DigitizerControl (DIC) on PMI-board to CassetteHandlingControl (CHC) on IP-Handling board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15366 | DIC_TO_CHC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | DIC could not read acknowledge from CHC | Meaning: Could not read acknowledge after sending message from DigitizerControl (DIC) on PMI-board to CassetteHandlingControl (CHC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on IP-Handling board; if connections are OK, install new software and inform R&D team.
15367 | DIC_TO_CHC_CAN_ERROR | Fatal error |  |  | sending from DIC to CHC failed with CAN fast channel error | Meaning: CAN error during sending message from DigitizerControl (DIC) on PMI-board to CassetteHandlingControl (CHC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15368 | DIC_TO_TRC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while DIC is sending to TRC | Meaning: No receiver found after sending message from DigitizerControl (DIC) on PMI-board to TagReaderControl (TRC) on IP-Handling board. Reason: Software error, input queue of TRC may be full. Cure: Install new software; inform R&D team.
15369 | DIC_TO_TRC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from DIC to TRC | Meaning: The CAN channel number for sending message from DigitizerControl (DIC) on PMI-board to TagReaderControl (TRC) on IP-Handling board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15370 | DIC_TO_TRC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | DIC could not read acknowledge from TRC | Meaning: Could not read acknowledge after sending message from DigitizerControl (DIC) on PMI-board to TagReaderControl (TRC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on IP-Handling board; if connections are OK, install new software and inform R&D team.
15371 | DIC_TO_TRC_CAN_ERROR | Fatal error |  |  | sending from DIC to TRC failed with CAN fast channel error | Meaning: CAN error during sending message from DigitizerControl (DIC) on PMI-board to TagReaderControl (TRC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15372 | UID_TO_UIC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while UID is sending to UIC | Meaning: No receiver found after sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on IP-Handling board. Reason: Software error, input queue of UIC may be full. Cure: Install new software; inform R&D team.
15373 | UID_TO_UIC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from UID to UIC | Meaning: The CAN channel number for sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on IP-Handling board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15374 | UID_TO_UIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | UID could not read acknowledge from UIC | Meaning: Could not read acknowledge after sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on IP-Handling board ; if connections are OK, install new software and inform R&D team.
15375 | UID_TO_UIC_CAN_ERROR | Fatal error |  |  | sending from UID to UIC failed with CAN fast channel error | Meaning: CAN error during sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15376 | SSC_TO_IHC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while SSC is sending to IHC | Meaning: No receiver found after sending message from SlowscanControl (SSC) on Slowscan board to IP-HandlingControl (IHC) on IP-Handling board. Reason: Software error, input queue of IHC may be full. Cure: Install new software; inform R&D team.
15377 | SSC_TO_IHC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from SSC to IHC | Meaning: The CAN channel number for sending message from SlowscanControl (SSC) on Slowscan board to IP-HandlingControl (IHC) on IP-Handling board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15378 | SSC_TO_IHC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | SSC could not read acknowledge from IHC | Meaning: Could not read acknowledge after sending message from SlowscanControl (SSC) on Slowscan board to IP-HandlingControl (IHC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on IP-Handling board; if connections are OK, install new software and inform R&D team.
15379 | SSC_TO_IHC_CAN_ERROR | Fatal error |  |  | sending from SSC to IHC failed with CAN fast channel error | Meaning: CAN error during sending message from SlowscanControl (SSC) on Slowscan board to IP-HandlingControl (IHC) on IP-Handling board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15380 | IHC_TO_DIC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while IHC is sending to DIC | Meaning: No receiver found after sending message from IP-HandlingControl (IHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: Software error, input queue of DIC may be full. Cure: Install new software; inform R&D team.
15381 | IHC_TO_DIC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from IHC to DIC | Meaning: The CAN channel number for sending message from IP-HandlingControl (IHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15382 | IHC_TO_DIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | IHC could not read acknowledge from DIC | Meaning: Could not read acknowledge after sending message from IP-HandlingControl (IHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15383 | IHC_TO_DIC_CAN_ERROR | Fatal error |  |  | sending from IHC to DIC failed with CAN fast channel error | Meaning: CAN error during sending message from IP-HandlingControl (IHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15384 | CHC_TO_DIC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while CHC is sending to DIC | Meaning: No receiver found after sending message from CassetteHandlingControl (CHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: Software error, input queue of DIC may be full. Cure: Install new software; inform R&D team.
15385 | CHC_TO_DIC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from CHC to DIC | Meaning: The CAN channel number for sending message from CassetteHandlingControl (CHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15386 | CHC_TO_DIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | CHC could not read acknowledge from DIC | Meaning: Could not read acknowledge after sending message from CassetteHandlingControl (CHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15387 | CHC_TO_DIC_CAN_ERROR | Fatal error |  |  | sending from CHC to DIC failed with CAN fast channel error | Meaning: CAN error during sending message from CassetteHandlingControl (CHC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15388 | TRC_TO_DIC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while TRC is sending to DIC | Meaning: No receiver found after sending message from TagReaderControl (TRC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: Software error, input queue of DIC may be full. Cure: Install new software; inform R&D team.
15389 | TRC_TO_DIC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from TRC to DIC | Meaning: The CAN channel number for sending message from TagReaderControl (TRC) on IP-Handling board to DigitizerControl (DIC) on PMI-board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15390 | TRC_TO_DIC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | TRC could not read acknowledge from DIC | Meaning: Could not read acknowledge after sending message from TagReaderControl (TRC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15391 | TRC_TO_DIC_CAN_ERROR | Fatal error |  |  | sending from TRC to DIC failed with CAN fast channel error | Meaning: CAN error during sending message from TagReaderControl (TRC) on IP-Handling board to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15392 | UIC_TO_UID_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while UIC is sending to UID | Meaning: No receiver found after sending message from UserInterfaceControl (UIC) on IP-Handling board to UserInterfacDispatcher(UID) on PMI-board. Reason: Software error, input queue of UID may be full. Cure: Install new software; inform R&D team.
15393 | UIC_TO_UID_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from UIC to UID | Meaning: The CAN channel number for sending message from UserInterfaceControl (UIC) on IP-Handling board to UserInterfacDispatcher(UID) on PMI-board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15394 | UIC_TO_UID_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | UIC could not read acknowledge from UID | Meaning: Could not read acknowledge after sending message from UserInterfaceControl (UIC) on IP-Handling board to UserInterfacDispatcher(UID) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15395 | UIC_TO_UID_CAN_ERROR | Fatal error |  |  | sending from UIC to UID failed with CAN fast channel error | Meaning: CAN error during sending message from UserInterfaceControl (UIC) on IP-Handling board to UserInterfacDispatcher(UID) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15396 | IHC_TO_SSC_NO_RECEIVER_FOUND | Fatal error |  |  | no suitable receiver found while IHC is sending to SSC | Meaning: No receiver found after sending message from IP-HandlingControl (IHC) on IP-Handling board to SlowscanControl(SSC) on Slowscan board. Reason: Software error, input queue of SSC may be full. Cure: Install new software; inform R&D team.
15397 | IHC_TO_SSC_UNKNOWN_CHANNEL | Fatal error |  |  | unknown CAN channel for sending from IHC to SCC | Meaning: The CAN channel number for sending message from IP-HandlingControl (IHC) on IP-Handling board to SlowscanControl(SSC) on Slowscan board is unknown. Reason: Software error. Cure: Install new software; inform R&D team.
15398 | IHC_TO_SSC_ERR_READING_ACKNOWLEDGE | Fatal error |  |  | IHC could not read acknowledge from SSC | Meaning: Could not read acknowledge after sending message from IP-HandlingControl (IHC) on IP-Handling board to SlowscanControl(SSC) on Slowscan board. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on Slowscan board; if connections are OK, install new software and inform R&D team.
15399 | IHC_TO_SSC_CAN_ERROR | Fatal error |  |  | sending from IHC to SSC failed with CAN fast channel error | Meaning: CAN error during sending message from IP-HandlingControl (IHC) on IP-Handling board to SlowscanControl(SSC) on Slowscan board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&D team.
15972 | MAXIMUM_NUMBER_OF_PLATEAUS_EXCEEDED | Error |  |  | The maximum number of plateaus was exceeded. | Meaning: The maximum number of plateaus was exceeded. Reason: The system (mechanics/stepper motor) is not stable enough for a calibration. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
15973 | NO_PLATEAU_FOUND_FOR_EVALUATION | Error |  |  | No plateau was found for determining the optimal run current level. | Meaning: No plateau was found for determining the optimal run current level. Reason: The system (mechanics/stepper motor) is not stable enough for a calibration. Cure: Check the corresponding mechanics and electronics and retry the calibration procedure.
15974 | WRONG_RUN_CURRENT_VALUE_SET | Error |  |  | The run current value set for the calibration cycle is wrong. | Meaning: The run current value set for the calibration cycle is wrong. The required and the set values are logged. Reason: -- Cure: Check the corresponding SW routine
15975 | FUNCTION_POINTER_ZERO | Error |  |  | A function pointer was zero. | Meaning: A function pointer was zero. Reason: -- Cure: Check the corresponding SW routine
15976 | MOTOR_SIM_CALIBRATION_FAILED | Error |  |  | The calibration of the Motor Simulation failed. | Meaning: The calibration of the Motor Simulation failed. Reason: The reason is given in the additional error code. Cure: Retry the calibration. If the problem persists check the corresponding Software or Settings.
15977 | MOTOR_SIM_CALIBRATION_ABORTED | Error |  |  | The calibration procedure of the Motor Simulation was aborted. | Meaning: The calibration procedure of the Motor Simulation was aborted. Reason: One of the calibration movements was not accomplished correctly. Cure: Restart the Digitizer and the calibration procedure. If error persists please call Service.
15978 | MOTOR_SIM_CALIBRATION_INVALID_VALUE | Error |  |  | The determined optimal run current value of the Motor Simulation was invalid. | Meaning: The determined optimal run current value of the Motor Simulation was invalid, i.e. the value is not within the required range. Reason: There is a problem either with the corresponding Software or Settings. Cure: Check the corresponding Software and Settings and retry the calibration procedure.
15979 | CB_FUNCTION_POINTER_ZERO | Error |  |  | A cb function pointer was zero. | Meaning: A cb function pointer was zero. Reason: -- Cure: Check the corresponding SW routine
15980 | WRONG_RUN_CURRENT_VALUE_STORED | Error |  |  | The run current value stored in the eLabel is wrong. | Meaning: The run current value stored in the eLabel is wrong. The required and the set values are logged. Reason: -- Cure: Check the corresponding SW routines and eventually the hardware.
16384 | GeneralError | Error |  |  | General error in ImageProcessing.
16385 | PXC_InsufficientCalWidth | Error | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Calibration line does not cover the complete width of the image to be corrected. Parts of the image will not be corrected. | Meaning: Only parts of the image will be corrected. The rest might have banding artefacts. Reason: Calibration line does not cover the complete width of the image to be corrected. Cure: Calibrate with the biggest used cassette.
16386 | PXC_CalcCorrLines | Error | Cassette format is not calibrated. Possible bad image quality. Please contact service to calibrate the digitizer. |  | Calculating the correction lines failed. Using constant correction line. | Meaning: The image is uncorrected. Reason: An exception occurred during preparation of the correction line. Might be software bug. Cure: Check the log file to find out more information about the error.
16409 | CIM_WriteImageMemoryException | Error |  |  | Exception at writing IP calibration image to memory for storing on ID tag.
16410 | CIM_WriteImageFileException | Error |  |  | Exception at writing IP calibration image to file.
16411 | CIM_WriteNoImage | Error |  |  | Writing IP calibration image failed because no image present. Software error.
16414 | CIM_RescaleNoImage | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | No image to rescale. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration or IP correction was disabled for this scan. Cure: Return the image plate to production for calibration.
16415 | CIM_RescaleIPCalDotsPerMeter | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | The IP calibration image (TIFF file) contains wrong values for DotsPerMeter. | Meaning: Parameter error in the image plate calibration image. Reason: This could be prototype plate from DX-S development. But usually, this problem should not occur. Cure: Delete the calibration image from the hard disk and return the image plate to production for calibration.
16416 | CIM_RescaleIPCalFailed | Error |  |  | Rescaling the given IP calibration image failed.
16417 | CIM_RescaleException | Error |  |  | Exception at rescaling IP calibration image.
16424 | CIM_ShiftAddCutNoCal | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | No IP calibration image available. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration. Cure: Return the image plate to production for calibration.
16425 | CIM_ShiftAddCutException | Error |  |  | Exception at shifting / extending / cutting IP calibration image.
16426 | CIM_ShiftAddCutTopExtra | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many lines on top of the calibration image had to be extrapolated. | Meaning: Too many lines on top of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
16427 | CIM_ShiftAddCutBottomExtra | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many lines at the bottom of the calibration image had to be extrapolated. | Meaning: Too many lines at the bottom of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
16428 | CIM_ShiftAddCutLeftExtra | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many columns at the left side of the calibration image had to be extrapolated. | Meaning: Too many columns at the left side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
16429 | CIM_ShiftAddCutRightExtra | Error | Image plate calibration not available. Possible reduced image quality. Please contact service. |  | Too many columns at the right side of the calibration image had to be extrapolated. | Meaning: Too many columns at the right side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
18432 | SRVC_SetError | Error |  |  | Setting an error failed.
18433 | SRVC_RemoveError | Error |  |  | Removing an error failed.
18435 | SRVC_StartCalibration | Info | Start shading calibration. | Confirm | Start shading calibration